## TR3 square Profile Parameters
ParameterGroup
Electrons profile analysis (current protocol: 'DataBook')
| Radiation device | Date | Time | Radiation type | SSD | Energy | Applicator | Scan type | Start position [De] | End position [De] | Field size inline | Field size crossline | Gantry angle | Collimator angle | Field detector | Penumbra margin | Scan direction | Scan mode | Flatness | Symmetry | FieldWidth | Penumbra | Center
| TR3 | 2018-09-05 | 8:39 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 7.02 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:40 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 7.13 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:41 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.9% | 100% | 7.35 cm | 1.39 cm - 1.39 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:41 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.6% | 100% | 7.44 cm | 1.85 cm - 1.85 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:42 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 6.50 cm | 6.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.3% | 100% | 7.53 cm | 2.31 cm - 2.31 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:44 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 7.01 cm | 0.40 cm - 0.40 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:45 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.5% | 100% | 7.17 cm | 0.78 cm - 0.78 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:46 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.6% | 100% | 7.30 cm | 1.46 cm - 1.46 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:46 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.9% | 100% | 7.34 cm | 1.73 cm - 1.73 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:47 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.2% | 100% | 7.40 cm | 2.20 cm - 2.20 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:49 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 7.00 cm | 0.44 cm - 0.44 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:50 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 7.06 cm | 0.53 cm - 0.53 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:50 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.3% | 100% | 7.16 cm | 1.01 cm - 1.01 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:51 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.1% | 100% | 7.23 cm | 1.59 cm - 1.59 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:52 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.2% | 100% | 7.24 cm | 1.83 cm - 1.83 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:55 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 7.01 cm | 0.58 cm - 0.58 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:56 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 7.04 cm | 0.68 cm - 0.68 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:56 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.2% | 100% | 7.10 cm | 1.07 cm - 1.07 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:57 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.4% | 100% | 7.13 cm | 1.36 cm - 1.36 cm | 0.00 cm
| TR3 | 2018-09-05 | 8:58 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.5% | 100% | 7.13 cm | 1.57 cm - 1.57 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:02 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 6.99 cm | 0.71 cm - 0.71 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:03 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.3% | 100% | 7.01 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:03 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.5% | 100% | 7.03 cm | 0.97 cm - 0.97 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:04 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.60 cm | 1.60 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.7% | 100% | 7.04 cm | 1.18 cm - 1.18 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:05 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.1% | 100% | 7.04 cm | 1.30 cm - 1.30 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:46 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 8.03 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:46 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 8.16 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:47 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.3% | 100% | 8.38 cm | 1.39 cm - 1.39 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:48 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.9% | 100% | 8.47 cm | 1.86 cm - 1.86 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:48 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 6.50 cm | 6.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.9% | 100% | 8.57 cm | 2.35 cm - 2.35 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:51 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 8.03 cm | 0.40 cm - 0.40 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:51 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.8% | 100% | 8.20 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:52 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.1% | 100% | 8.34 cm | 1.48 cm - 1.48 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:53 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.4% | 100% | 8.37 cm | 1.75 cm - 1.75 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:54 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.7% | 100% | 8.44 cm | 2.24 cm - 2.24 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:56 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 8.02 cm | 0.44 cm - 0.44 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:57 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 8.08 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:57 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.1% | 100% | 8.20 cm | 1.03 cm - 1.03 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:58 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.6% | 100% | 8.26 cm | 1.61 cm - 1.61 cm | 0.00 cm
| TR3 | 2018-09-05 | 9:59 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.4% | 100% | 8.29 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:03 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 8.03 cm | 0.58 cm - 0.58 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:03 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 8.07 cm | 0.69 cm - 0.69 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:04 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.8% | 100% | 8.14 cm | 1.09 cm - 1.09 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:05 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.7% | 100% | 8.16 cm | 1.38 cm - 1.38 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:06 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.5% | 100% | 8.17 cm | 1.57 cm - 1.57 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:08 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 8.01 cm | 0.71 cm - 0.71 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:09 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 8.02 cm | 0.80 cm - 0.80 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:09 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.2% | 100% | 8.05 cm | 0.99 cm - 0.99 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:10 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.60 cm | 1.60 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.2% | 100% | 8.06 cm | 1.19 cm - 1.19 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:11 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.3% | 100% | 8.06 cm | 1.30 cm - 1.30 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:43 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 6.08 cm | 0.37 cm - 0.37 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:44 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 6.18 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:45 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.5% | 100% | 6.37 cm | 1.38 cm - 1.38 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:45 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.2% | 100% | 6.46 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:46 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.7% | 100% | 6.58 cm | 2.29 cm - 2.29 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:48 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 6.06 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:49 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.9% | 100% | 6.20 cm | 0.78 cm - 0.78 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:50 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.6% | 100% | 6.32 cm | 1.46 cm - 1.46 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:50 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.7% | 100% | 6.36 cm | 1.72 cm - 1.72 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:51 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.7% | 100% | 6.44 cm | 2.18 cm - 2.18 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:53 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 6.06 cm | 0.42 cm - 0.42 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:54 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.3% | 100% | 6.10 cm | 0.52 cm - 0.52 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:54 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.5% | 100% | 6.20 cm | 1.03 cm - 1.03 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:55 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.2% | 100% | 6.25 cm | 1.60 cm - 1.60 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:56 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.2% | 100% | 6.27 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:59 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 6.07 cm | 0.56 cm - 0.56 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:00 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4% | 100% | 6.10 cm | 0.68 cm - 0.68 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:01 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.3% | 100% | 6.15 cm | 1.10 cm - 1.10 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:01 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.6% | 100% | 6.17 cm | 1.39 cm - 1.39 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:02 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.5% | 100% | 6.17 cm | 1.59 cm - 1.59 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:04 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5% | 100% | 6.07 cm | 0.68 cm - 0.68 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:05 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.1% | 100% | 6.08 cm | 0.78 cm - 0.78 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:06 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.6% | 100% | 6.10 cm | 0.99 cm - 0.99 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:06 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.60 cm | 1.60 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11% | 100% | 6.11 cm | 1.20 cm - 1.20 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:07 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.5% | 100% | 6.10 cm | 1.33 cm - 1.33 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:11 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 5.02 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:11 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 5.11 cm | 0.56 cm - 0.56 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:12 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14% | 100% | 5.29 cm | 1.37 cm - 1.37 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:13 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17% | 100% | 5.41 cm | 1.80 cm - 1.80 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:13 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.9% | 100% | 5.59 cm | 2.21 cm - 2.21 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:16 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 5.01 cm | 0.40 cm - 0.40 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:16 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.2% | 100% | 5.13 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:17 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.8% | 100% | 5.26 cm | 1.44 cm - 1.44 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:17 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.5% | 100% | 5.31 cm | 1.67 cm - 1.67 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:18 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19% | 100% | 5.46 cm | 2.10 cm - 2.10 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:20 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 5.00 cm | 0.44 cm - 0.44 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:21 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.5% | 100% | 5.05 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:21 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.8% | 100% | 5.13 cm | 1.03 cm - 1.03 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:22 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.2% | 100% | 5.22 cm | 1.57 cm - 1.57 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:22 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.9% | 100% | 5.26 cm | 1.79 cm - 1.79 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:25 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.4% | 100% | 5.02 cm | 0.59 cm - 0.59 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:25 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.3% | 100% | 5.04 cm | 0.70 cm - 0.70 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:26 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.9% | 100% | 5.09 cm | 1.10 cm - 1.10 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:27 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15% | 100% | 5.11 cm | 1.39 cm - 1.39 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:27 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.9% | 100% | 5.13 cm | 1.58 cm - 1.58 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:29 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.7% | 100% | 5.01 cm | 0.71 cm - 0.71 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:30 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.2% | 100% | 5.03 cm | 0.80 cm - 0.80 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:30 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11% | 100% | 5.05 cm | 0.99 cm - 0.99 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:31 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.60 cm | 1.60 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.5% | 100% | 5.06 cm | 1.19 cm - 1.19 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:32 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.8% | 100% | 5.05 cm | 1.31 cm - 1.31 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:35 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 4.61 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:36 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.8% | 100% | 4.68 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:36 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.9% | 100% | 4.86 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:37 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.7% | 100% | 4.98 cm | 1.76 cm - 1.76 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:37 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.3% | 100% | 5.20 cm | 2.16 cm - 2.16 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:40 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 4.59 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:40 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8% | 100% | 4.70 cm | 0.78 cm - 0.78 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:41 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.7% | 100% | 4.83 cm | 1.41 cm - 1.41 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:42 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.3% | 100% | 4.90 cm | 1.65 cm - 1.65 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:42 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.4% | 100% | 5.08 cm | 2.04 cm - 2.04 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:44 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2% | 100% | 4.59 cm | 0.43 cm - 0.43 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:45 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.9% | 100% | 4.62 cm | 0.53 cm - 0.53 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:46 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.7% | 100% | 4.71 cm | 1.01 cm - 1.01 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:46 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.9% | 100% | 4.80 cm | 1.54 cm - 1.54 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:47 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.4% | 100% | 4.86 cm | 1.75 cm - 1.75 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:49 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.9% | 100% | 4.60 cm | 0.58 cm - 0.58 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:50 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.9% | 100% | 4.62 cm | 0.69 cm - 0.69 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:50 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.7% | 100% | 4.67 cm | 1.08 cm - 1.08 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:51 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.9% | 100% | 4.70 cm | 1.36 cm - 1.36 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:51 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.7% | 100% | 4.71 cm | 1.56 cm - 1.56 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:56 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.3% | 100% | 4.60 cm | 0.69 cm - 0.69 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:57 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.8% | 100% | 4.61 cm | 0.78 cm - 0.78 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:57 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.7% | 100% | 4.63 cm | 0.97 cm - 0.97 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:58 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.60 cm | 1.60 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.3% | 100% | 4.64 cm | 1.18 cm - 1.18 cm | 0.00 cm
| TR3 | 2018-09-05 | 11:59 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.7% | 100% | 4.63 cm | 1.30 cm - 1.30 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:00 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 25.29 cm | 0.35 cm - 0.35 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:02 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 4.04 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:02 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 25.69 cm | 0.52 cm - 0.52 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:03 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.1% | 100% | 4.10 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:03 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.3% | 100% | 4.25 cm | 1.34 cm - 1.34 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:03 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 4.50 cm | 4.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 26.39 cm | 1.39 cm - 1.39 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:04 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.9% | 0% | 4.38 cm | 1.73 cm - 1.73 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:04 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.1% | 0% | 4.63 cm | 2.10 cm - 2.10 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:04 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 5.50 cm | 5.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 26.65 cm | 1.88 cm - 1.88 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:05 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 6.50 cm | 6.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 26.87 cm | 2.40 cm - 2.40 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:06 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 4.02 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:07 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.7% | 100% | 4.12 cm | 0.77 cm - 0.77 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:08 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.1% | 100% | 4.22 cm | 1.39 cm - 1.39 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:08 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.4% | 100% | 4.18 cm | 1.20 cm - 1.20 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:09 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.2% | 0% | 4.51 cm | 1.98 cm - 1.98 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:10 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 25.29 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:11 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.9% | 100% | 4.01 cm | 0.42 cm - 0.42 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:12 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.1% | 100% | 4.04 cm | 0.52 cm - 0.52 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:12 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.2% | 100% | 4.11 cm | 0.99 cm - 0.99 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:12 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 25.86 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:13 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.2% | 0% | 4.20 cm | 1.48 cm - 1.48 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:13 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 4.00 cm | 4.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 26.31 cm | 1.45 cm - 1.45 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:14 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.5% | 0% | 4.26 cm | 1.69 cm - 1.69 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:14 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 4.50 cm | 4.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 26.45 cm | 1.72 cm - 1.72 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:16 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.3% | 100% | 4.01 cm | 0.57 cm - 0.57 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:16 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.6% | 100% | 4.03 cm | 0.68 cm - 0.68 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:16 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 5.50 cm | 5.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 26.60 cm | 2.20 cm - 2.20 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:17 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.4% | 100% | 4.07 cm | 1.06 cm - 1.06 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:17 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.3% | 0% | 4.08 cm | 1.33 cm - 1.33 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:18 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19% | 0% | 4.09 cm | 1.52 cm - 1.52 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:19 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 25.30 cm | 0.41 cm - 0.41 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:20 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.9% | 100% | 4.01 cm | 0.68 cm - 0.68 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:21 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.4% | 100% | 4.02 cm | 0.77 cm - 0.77 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:21 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.4% | 100% | 4.03 cm | 0.95 cm - 0.95 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:21 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 1.20 cm | 1.20 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 25.50 cm | 0.50 cm - 0.50 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:22 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.60 cm | 1.60 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16% | 100% | 4.02 cm | 1.16 cm - 1.16 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:22 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.5% | 0% | 4.01 cm | 1.28 cm - 1.28 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:22 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.92 cm | 0.98 cm - 0.98 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:23 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 3.50 cm | 3.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 26.20 cm | 1.54 cm - 1.54 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:25 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 4.00 cm | 4.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 26.27 cm | 1.78 cm - 1.78 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:27 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 3.12 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:27 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.3% | 100% | 3.18 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:28 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.4% | 0% | 3.38 cm | 1.28 cm - 1.28 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:29 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20% | 0% | 3.62 cm | 1.63 cm - 1.63 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:29 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.9% | 0% | 3.99 cm | 2.00 cm - 2.00 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:31 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.3% | 100% | 3.11 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:32 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.9% | 0% | 3.19 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:33 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19% | 0% | 3.38 cm | 1.31 cm - 1.31 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:33 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.8% | 0% | 3.52 cm | 1.51 cm - 1.51 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:34 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.9% | 0% | 3.88 cm | 1.87 cm - 1.87 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:36 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.6% | 100% | 3.10 cm | 0.43 cm - 0.43 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:36 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.3% | 100% | 3.12 cm | 0.52 cm - 0.52 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:37 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 0.80 cm | 0.80 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 25.39 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:37 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16% | 0% | 3.21 cm | 0.95 cm - 0.95 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:37 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.5% | 0% | 3.42 cm | 1.39 cm - 1.39 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:38 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.4% | 0% | 3.58 cm | 1.58 cm - 1.58 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:39 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 1.20 cm | 1.20 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.52 cm | 0.65 cm - 0.65 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:40 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.9% | 100% | 3.10 cm | 0.57 cm - 0.57 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:40 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.9% | 0% | 3.12 cm | 0.67 cm - 0.67 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:40 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.77 cm | 1.05 cm - 1.05 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:41 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.9% | 0% | 3.19 cm | 1.01 cm - 1.01 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:41 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.88 cm | 1.34 cm - 1.34 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:42 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.1% | 0% | 3.29 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:42 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.3% | 0% | 3.38 cm | 1.41 cm - 1.41 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:43 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.91 cm | 1.55 cm - 1.55 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:46 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 25.28 cm | 0.71 cm - 0.71 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:47 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 0.80 cm | 0.80 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 25.37 cm | 0.81 cm - 0.81 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:48 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 1.20 cm | 1.20 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 25.49 cm | 1.00 cm - 1.00 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:49 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.5% | 0% | 3.10 cm | 0.68 cm - 0.68 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:50 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.7% | 0% | 3.11 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:50 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.2% | 0% | 3.14 cm | 0.92 cm - 0.92 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:50 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 1.60 cm | 1.60 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 25.54 cm | 1.23 cm - 1.23 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:51 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.60 cm | 1.60 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.2% | 0% | 3.18 cm | 1.10 cm - 1.10 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:51 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 25.52 cm | 1.38 cm - 1.38 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:52 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.3% | 0% | 3.20 cm | 1.22 cm - 1.22 cm | 0.00 cm
| TR3 | 2018-09-04 | 12:59 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 20.22 cm | 0.36 cm - 0.36 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:00 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.54 cm | 0.52 cm - 0.52 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:01 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 4.50 cm | 4.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 21.07 cm | 1.37 cm - 1.37 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:02 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 5.50 cm | 5.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 21.29 cm | 1.86 cm - 1.86 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:03 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 6.50 cm | 6.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.9% | 100% | 21.45 cm | 2.38 cm - 2.38 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:07 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 20.22 cm | 0.37 cm - 0.37 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:08 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 20.67 cm | 0.74 cm - 0.74 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:09 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 4.00 cm | 4.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 21.02 cm | 1.41 cm - 1.41 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:10 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 4.50 cm | 4.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 21.13 cm | 1.68 cm - 1.68 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:12 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 5.50 cm | 5.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 21.26 cm | 2.17 cm - 2.17 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:15 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 20.23 cm | 0.41 cm - 0.41 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:16 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 1.20 cm | 1.20 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 20.38 cm | 0.49 cm - 0.49 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:17 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 20.71 cm | 0.96 cm - 0.96 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:18 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 3.50 cm | 3.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.93 cm | 1.51 cm - 1.51 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:19 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 4.00 cm | 4.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 21.00 cm | 1.74 cm - 1.74 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:23 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 0.80 cm | 0.80 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.29 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:24 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 1.20 cm | 1.20 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.39 cm | 0.64 cm - 0.64 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:25 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 20.58 cm | 1.02 cm - 1.02 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:27 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.67 cm | 1.31 cm - 1.31 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:28 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 20.70 cm | 1.51 cm - 1.51 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:31 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 20.20 cm | 0.70 cm - 0.70 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:32 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 0.80 cm | 0.80 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 20.27 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:33 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 1.20 cm | 1.20 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 20.35 cm | 0.98 cm - 0.98 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:35 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 1.60 cm | 1.60 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 20.40 cm | 1.20 cm - 1.20 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:36 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 20.38 cm | 1.34 cm - 1.34 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:48 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 17.15 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:50 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 17.43 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:51 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 4.50 cm | 4.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.1% | 100% | 17.89 cm | 1.38 cm - 1.38 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:52 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 5.50 cm | 5.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.6% | 100% | 18.08 cm | 1.88 cm - 1.88 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:53 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 6.50 cm | 6.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.8% | 100% | 18.22 cm | 2.39 cm - 2.39 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:56 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 17.15 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:57 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 17.53 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:58 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 4.00 cm | 4.00 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 17.84 cm | 1.44 cm - 1.44 cm | 0.00 cm
| TR3 | 2018-09-04 | 13:59 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 4.50 cm | 4.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.2% | 100% | 17.92 cm | 1.70 cm - 1.70 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:00 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 5.50 cm | 5.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.4% | 100% | 18.05 cm | 2.20 cm - 2.20 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:04 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 17.15 cm | 0.45 cm - 0.45 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:05 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 1.20 cm | 1.20 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 17.28 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:06 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 17.55 cm | 1.00 cm - 1.00 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:07 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 3.50 cm | 3.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 17.74 cm | 1.54 cm - 1.54 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:08 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 4.00 cm | 4.00 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.7% | 100% | 17.79 cm | 1.78 cm - 1.78 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:11 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 0.80 cm | 0.80 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 17.20 cm | 0.57 cm - 0.57 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:12 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 1.20 cm | 1.20 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 17.28 cm | 0.67 cm - 0.67 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:13 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 17.45 cm | 1.04 cm - 1.04 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:14 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 17.53 cm | 1.33 cm - 1.33 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:15 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 17.56 cm | 1.52 cm - 1.52 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:18 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 17.12 cm | 0.74 cm - 0.74 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:19 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 0.80 cm | 0.80 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 17.18 cm | 0.82 cm - 0.82 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:20 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 1.20 cm | 1.20 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 17.25 cm | 0.99 cm - 0.99 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:21 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 1.60 cm | 1.60 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 17.29 cm | 1.20 cm - 1.20 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:22 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 17.28 cm | 1.33 cm - 1.33 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:29 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 15.12 cm | 0.37 cm - 0.37 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:30 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 15.35 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:31 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.9% | 100% | 15.75 cm | 1.41 cm - 1.41 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:32 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.9% | 100% | 15.90 cm | 1.90 cm - 1.90 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:33 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 6.50 cm | 6.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.3% | 100% | 16.03 cm | 2.43 cm - 2.43 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:35 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 15.11 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:36 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 15.45 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:37 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.4% | 100% | 15.71 cm | 1.45 cm - 1.45 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:38 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.6% | 100% | 15.79 cm | 1.73 cm - 1.73 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:39 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.1% | 100% | 15.90 cm | 2.22 cm - 2.22 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:42 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 15.11 cm | 0.42 cm - 0.42 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:43 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 1.20 cm | 1.20 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 15.23 cm | 0.51 cm - 0.51 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:44 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 15.47 cm | 0.99 cm - 0.99 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:45 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 3.50 cm | 3.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.9% | 100% | 15.63 cm | 1.57 cm - 1.57 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:46 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.1% | 100% | 15.67 cm | 1.82 cm - 1.82 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:49 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 15.16 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:50 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 1.20 cm | 1.20 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 15.24 cm | 0.65 cm - 0.65 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:51 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 15.38 cm | 1.04 cm - 1.04 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:52 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.4% | 100% | 15.45 cm | 1.34 cm - 1.34 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:53 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3% | 100% | 15.46 cm | 1.57 cm - 1.57 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:55 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.4% | 100% | 15.09 cm | 0.70 cm - 0.70 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:56 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 15.15 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:57 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 1.20 cm | 1.20 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.7% | 100% | 15.21 cm | 0.98 cm - 0.98 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:58 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 1.60 cm | 1.60 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3% | 100% | 15.24 cm | 1.21 cm - 1.21 cm | 0.00 cm
| TR3 | 2018-09-04 | 14:59 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 15.23 cm | 1.36 cm - 1.36 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:05 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 13.04 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:06 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 13.24 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:07 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.5% | 100% | 13.60 cm | 1.39 cm - 1.39 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:08 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.1% | 100% | 13.73 cm | 1.88 cm - 1.88 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:09 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 6.50 cm | 6.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.9% | 100% | 13.85 cm | 2.40 cm - 2.40 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:11 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.2% | 100% | 13.03 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:12 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 13.32 cm | 0.77 cm - 0.77 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:13 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 13.55 cm | 1.45 cm - 1.45 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:14 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.9% | 100% | 13.62 cm | 1.72 cm - 1.72 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:15 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.9% | 100% | 13.71 cm | 2.22 cm - 2.22 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:17 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.2% | 100% | 13.03 cm | 0.44 cm - 0.44 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:18 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 1.20 cm | 1.20 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.2% | 100% | 13.13 cm | 0.53 cm - 0.53 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:19 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 13.34 cm | 0.99 cm - 0.99 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:20 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 3.50 cm | 3.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4% | 100% | 13.47 cm | 1.56 cm - 1.56 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:21 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.5% | 100% | 13.52 cm | 1.81 cm - 1.81 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:24 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 13.07 cm | 0.57 cm - 0.57 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:25 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 1.20 cm | 1.20 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 13.13 cm | 0.67 cm - 0.67 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:26 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 13.26 cm | 1.04 cm - 1.04 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:26 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.9% | 100% | 13.31 cm | 1.33 cm - 1.33 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:27 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.1% | 100% | 13.33 cm | 1.55 cm - 1.55 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:30 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 13.01 cm | 0.72 cm - 0.72 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:31 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 13.06 cm | 0.80 cm - 0.80 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:32 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 1.20 cm | 1.20 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.2% | 100% | 13.11 cm | 0.97 cm - 0.97 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:33 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 1.60 cm | 1.60 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.8% | 100% | 13.13 cm | 1.18 cm - 1.18 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:34 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 13.13 cm | 1.32 cm - 1.32 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:19 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 10.14 cm | 0.36 cm - 0.36 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:20 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 10.30 cm | 0.51 cm - 0.51 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:21 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.4% | 100% | 10.57 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:21 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.6% | 100% | 10.68 cm | 1.85 cm - 1.85 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:22 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 6.50 cm | 6.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.8% | 100% | 10.76 cm | 2.35 cm - 2.35 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:31 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 10.13 cm | 0.37 cm - 0.37 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:32 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 10.35 cm | 0.74 cm - 0.74 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:33 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.6% | 100% | 10.53 cm | 1.41 cm - 1.41 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:34 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.6% | 100% | 10.58 cm | 1.69 cm - 1.69 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:35 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.2% | 100% | 10.65 cm | 2.20 cm - 2.20 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:38 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 10.12 cm | 0.40 cm - 0.40 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:39 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 10.20 cm | 0.50 cm - 0.50 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:39 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.9% | 100% | 10.35 cm | 0.96 cm - 0.96 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:40 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.9% | 100% | 10.45 cm | 1.55 cm - 1.55 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:41 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.8% | 100% | 10.48 cm | 1.79 cm - 1.79 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:44 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 10.15 cm | 0.52 cm - 0.52 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:45 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 10.20 cm | 0.62 cm - 0.62 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:46 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.8% | 100% | 10.30 cm | 1.00 cm - 1.00 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:47 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5% | 100% | 10.34 cm | 1.29 cm - 1.29 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:48 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.7% | 100% | 10.36 cm | 1.52 cm - 1.52 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:50 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 10.11 cm | 0.66 cm - 0.66 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:51 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.9% | 100% | 10.15 cm | 0.74 cm - 0.74 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:52 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.7% | 100% | 10.19 cm | 0.93 cm - 0.93 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:53 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.60 cm | 1.60 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.8% | 100% | 10.21 cm | 1.14 cm - 1.14 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:54 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.6% | 100% | 10.20 cm | 1.29 cm - 1.29 cm | 0.00 cm
| TR3 | 2018-09-07 | 9:47 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 28.02 cm | 0.65 cm - 0.65 cm | 0.00 cm
| TR3 | 2018-09-07 | 9:48 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 28.55 cm | 0.92 cm - 0.92 cm | 0.00 cm
| TR3 | 2018-09-07 | 9:50 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 5.00 cm | 5.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 29.22 cm | 1.81 cm - 1.81 cm | 0.00 cm
| TR3 | 2018-09-07 | 9:51 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 5.50 cm | 5.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 29.36 cm | 2.06 cm - 2.06 cm | 0.00 cm
| TR3 | 2018-09-07 | 9:52 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 6.50 cm | 6.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 29.55 cm | 2.54 cm - 2.54 cm | 0.00 cm
| TR3 | 2018-09-07 | 9:56 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 28.04 cm | 0.75 cm - 0.75 cm | 0.00 cm
| TR3 | 2018-09-07 | 9:58 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 28.74 cm | 1.24 cm - 1.24 cm | 0.00 cm
| TR3 | 2018-09-07 | 9:59 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 4.50 cm | 4.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 29.15 cm | 1.92 cm - 1.92 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:00 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 5.00 cm | 5.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 29.24 cm | 2.17 cm - 2.17 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:02 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 5.50 cm | 5.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 29.31 cm | 2.37 cm - 2.37 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:05 | Electrons | 110.00 cm | 12 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 28.08 cm | 0.92 cm - 0.92 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:06 | Electrons | 110.00 cm | 12 MeV | 25x25 | Crossline | 1.20 cm | 1.20 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 28.27 cm | 1.00 cm - 1.00 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:07 | Electrons | 110.00 cm | 12 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 28.67 cm | 1.38 cm - 1.38 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:09 | Electrons | 110.00 cm | 12 MeV | 25x25 | Crossline | 3.50 cm | 3.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 28.92 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:10 | Electrons | 110.00 cm | 12 MeV | 25x25 | Crossline | 4.00 cm | 4.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 28.99 cm | 2.03 cm - 2.03 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:13 | Electrons | 110.00 cm | 9 MeV | 25x25 | Crossline | 0.80 cm | 0.80 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 28.17 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:14 | Electrons | 110.00 cm | 9 MeV | 25x25 | Crossline | 1.20 cm | 1.20 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 28.30 cm | 1.33 cm - 1.33 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:15 | Electrons | 110.00 cm | 9 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 28.54 cm | 1.62 cm - 1.62 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:17 | Electrons | 110.00 cm | 9 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 28.62 cm | 1.82 cm - 1.82 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:18 | Electrons | 110.00 cm | 9 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 28.62 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:21 | Electrons | 110.00 cm | 6 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.7% | 100% | 28.02 cm | 1.82 cm - 1.82 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:22 | Electrons | 110.00 cm | 6 MeV | 25x25 | Crossline | 0.80 cm | 0.80 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 28.13 cm | 1.89 cm - 1.89 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:24 | Electrons | 110.00 cm | 6 MeV | 25x25 | Crossline | 1.20 cm | 1.20 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 28.22 cm | 2.00 cm - 2.00 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:25 | Electrons | 110.00 cm | 6 MeV | 25x25 | Crossline | 1.60 cm | 1.60 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.7% | 100% | 28.25 cm | 2.10 cm - 2.10 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:27 | Electrons | 110.00 cm | 6 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.8% | 100% | 28.20 cm | 2.13 cm - 2.13 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:31 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 22.37 cm | 0.64 cm - 0.64 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:32 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 22.79 cm | 0.92 cm - 0.92 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:33 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 5.00 cm | 5.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.2% | 100% | 23.34 cm | 1.83 cm - 1.83 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:35 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 5.50 cm | 5.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.7% | 100% | 23.44 cm | 2.08 cm - 2.08 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:36 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 6.50 cm | 6.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4% | 100% | 23.61 cm | 2.57 cm - 2.57 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:39 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 22.38 cm | 0.74 cm - 0.74 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:40 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 22.95 cm | 1.24 cm - 1.24 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:41 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 4.50 cm | 4.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 23.29 cm | 1.93 cm - 1.93 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:42 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 5.00 cm | 5.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 23.38 cm | 2.18 cm - 2.18 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:44 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 5.50 cm | 5.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 23.43 cm | 2.38 cm - 2.38 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:46 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 22.41 cm | 0.91 cm - 0.91 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:47 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 1.20 cm | 1.20 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 22.57 cm | 1.00 cm - 1.00 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:48 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 22.90 cm | 1.38 cm - 1.38 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:50 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 3.50 cm | 3.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 23.10 cm | 1.85 cm - 1.85 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:51 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 4.00 cm | 4.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 23.16 cm | 2.04 cm - 2.04 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:53 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 0.80 cm | 0.80 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 22.48 cm | 1.24 cm - 1.24 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:54 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 1.20 cm | 1.20 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 22.58 cm | 1.33 cm - 1.33 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:56 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 22.77 cm | 1.63 cm - 1.63 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:57 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 22.85 cm | 1.83 cm - 1.83 cm | 0.00 cm
| TR3 | 2018-09-07 | 10:58 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 22.85 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:00 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 22.36 cm | 1.79 cm - 1.79 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:02 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 0.80 cm | 0.80 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 22.44 cm | 1.87 cm - 1.87 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:03 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 1.20 cm | 1.20 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.6% | 100% | 22.51 cm | 1.99 cm - 1.99 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:04 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 1.60 cm | 1.60 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.6% | 100% | 22.52 cm | 2.09 cm - 2.09 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:05 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.6% | 100% | 22.50 cm | 2.12 cm - 2.12 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:09 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 18.98 cm | 0.68 cm - 0.68 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:10 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 19.33 cm | 0.93 cm - 0.93 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:11 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 5.00 cm | 5.00 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.1% | 100% | 19.80 cm | 1.80 cm - 1.80 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:12 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 5.50 cm | 5.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.8% | 100% | 19.89 cm | 2.05 cm - 2.05 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:13 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 6.50 cm | 6.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.4% | 100% | 20.02 cm | 2.54 cm - 2.54 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:16 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.2% | 100% | 19.00 cm | 0.77 cm - 0.77 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:18 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 19.46 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:19 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 4.50 cm | 4.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 19.75 cm | 1.92 cm - 1.92 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:20 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 5.00 cm | 5.00 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.4% | 100% | 19.82 cm | 2.16 cm - 2.16 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:21 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 5.50 cm | 5.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.4% | 100% | 19.86 cm | 2.38 cm - 2.38 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:24 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 19.01 cm | 0.97 cm - 0.97 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:25 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 1.20 cm | 1.20 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 19.14 cm | 1.04 cm - 1.04 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:26 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 19.41 cm | 1.41 cm - 1.41 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:27 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 3.50 cm | 3.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.3% | 100% | 19.57 cm | 1.87 cm - 1.87 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:28 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 4.00 cm | 4.00 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3% | 100% | 19.62 cm | 2.06 cm - 2.06 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:33 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 0.80 cm | 0.80 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 19.07 cm | 1.27 cm - 1.27 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:34 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 1.20 cm | 1.20 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 19.15 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:35 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 19.31 cm | 1.63 cm - 1.63 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:36 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.3% | 100% | 19.36 cm | 1.83 cm - 1.83 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:37 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.8% | 100% | 19.38 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:40 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.1% | 100% | 18.97 cm | 1.80 cm - 1.80 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:42 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 0.80 cm | 0.80 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 19.04 cm | 1.86 cm - 1.86 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:43 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 1.20 cm | 1.20 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.6% | 100% | 19.10 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:44 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 1.60 cm | 1.60 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.1% | 100% | 19.11 cm | 2.06 cm - 2.06 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:45 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 17.00 cm | 17.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.1% | 100% | 19.09 cm | 2.09 cm - 2.09 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:50 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 16.74 cm | 0.65 cm - 0.65 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:51 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2% | 100% | 17.05 cm | 0.92 cm - 0.92 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:52 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 5.00 cm | 5.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.1% | 100% | 17.46 cm | 1.83 cm - 1.83 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:53 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.1% | 100% | 17.53 cm | 2.09 cm - 2.09 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:54 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 6.50 cm | 6.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.1% | 100% | 17.64 cm | 2.58 cm - 2.58 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:57 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 16.75 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:58 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 17.17 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR3 | 2018-09-07 | 11:59 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.9% | 100% | 17.42 cm | 1.94 cm - 1.94 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:00 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 5.00 cm | 5.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.1% | 100% | 17.49 cm | 2.19 cm - 2.19 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:01 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.1% | 100% | 17.51 cm | 2.41 cm - 2.41 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:03 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 16.78 cm | 0.91 cm - 0.91 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:05 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 1.20 cm | 1.20 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 16.90 cm | 1.00 cm - 1.00 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:05 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 17.14 cm | 1.39 cm - 1.39 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:06 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 3.50 cm | 3.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.5% | 100% | 17.28 cm | 1.87 cm - 1.87 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:07 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.5% | 100% | 17.32 cm | 2.07 cm - 2.07 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:10 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 16.83 cm | 1.23 cm - 1.23 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:11 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 1.20 cm | 1.20 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 16.91 cm | 1.32 cm - 1.32 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:12 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.9% | 100% | 17.05 cm | 1.61 cm - 1.61 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:13 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.7% | 100% | 17.10 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:14 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.3% | 100% | 17.09 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:16 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.5% | 100% | 16.76 cm | 1.75 cm - 1.75 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:17 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.8% | 100% | 16.82 cm | 1.82 cm - 1.82 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:18 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 1.20 cm | 1.20 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.2% | 100% | 16.88 cm | 1.94 cm - 1.94 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:19 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 1.60 cm | 1.60 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.6% | 100% | 16.88 cm | 2.06 cm - 2.06 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:20 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.9% | 100% | 16.86 cm | 2.10 cm - 2.10 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:25 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 14.44 cm | 0.67 cm - 0.67 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:26 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 14.72 cm | 0.92 cm - 0.92 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:27 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 5.00 cm | 5.00 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.9% | 100% | 15.08 cm | 1.81 cm - 1.81 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:28 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.2% | 100% | 15.14 cm | 2.07 cm - 2.07 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:29 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 6.50 cm | 6.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.7% | 100% | 15.24 cm | 2.57 cm - 2.57 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:31 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 14.45 cm | 0.77 cm - 0.77 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:32 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2% | 100% | 14.81 cm | 1.26 cm - 1.26 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:33 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.2% | 100% | 15.03 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:34 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 5.00 cm | 5.00 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.6% | 100% | 15.08 cm | 2.20 cm - 2.20 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:35 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.8% | 100% | 15.12 cm | 2.41 cm - 2.41 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:38 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 14.46 cm | 0.95 cm - 0.95 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:39 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 1.20 cm | 1.20 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 14.57 cm | 1.03 cm - 1.03 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:40 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.4% | 100% | 14.77 cm | 1.41 cm - 1.41 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:41 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 3.50 cm | 3.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.9% | 100% | 14.90 cm | 1.88 cm - 1.88 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:41 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.1% | 100% | 14.93 cm | 2.09 cm - 2.09 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:44 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2% | 100% | 14.50 cm | 1.27 cm - 1.27 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:45 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 1.20 cm | 1.20 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.3% | 100% | 14.57 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:46 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.7% | 100% | 14.70 cm | 1.64 cm - 1.64 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:47 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.9% | 100% | 14.74 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:48 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.7% | 100% | 14.74 cm | 1.98 cm - 1.98 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:50 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5% | 100% | 14.45 cm | 1.76 cm - 1.76 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:51 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.2% | 100% | 14.50 cm | 1.83 cm - 1.83 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:52 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 1.20 cm | 1.20 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.9% | 100% | 14.56 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:53 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 1.60 cm | 1.60 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.5% | 100% | 14.57 cm | 2.04 cm - 2.04 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:53 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 13.00 cm | 13.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.9% | 100% | 14.53 cm | 2.10 cm - 2.10 cm | 0.00 cm
| TR3 | 2018-09-07 | 12:59 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 11.21 cm | 0.63 cm - 0.63 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:00 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.1% | 100% | 11.43 cm | 0.89 cm - 0.89 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:01 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.1% | 100% | 11.71 cm | 1.78 cm - 1.78 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:01 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.7% | 100% | 11.76 cm | 2.03 cm - 2.03 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:02 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 6.50 cm | 6.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.5% | 100% | 11.84 cm | 2.51 cm - 2.51 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:04 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 11.21 cm | 0.72 cm - 0.72 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:05 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 11.50 cm | 1.21 cm - 1.21 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:06 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.9% | 100% | 11.67 cm | 1.91 cm - 1.91 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:07 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.6% | 100% | 11.70 cm | 2.17 cm - 2.17 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:08 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.9% | 100% | 11.73 cm | 2.38 cm - 2.38 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:10 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 11.23 cm | 0.89 cm - 0.89 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:11 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.2% | 100% | 11.31 cm | 0.97 cm - 0.97 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:12 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.6% | 100% | 11.46 cm | 1.37 cm - 1.37 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:12 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.8% | 100% | 11.55 cm | 1.85 cm - 1.85 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:13 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.3% | 100% | 11.57 cm | 2.07 cm - 2.07 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:15 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 11.26 cm | 1.21 cm - 1.21 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:16 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.8% | 100% | 11.30 cm | 1.29 cm - 1.29 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:17 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.9% | 100% | 11.40 cm | 1.58 cm - 1.58 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:18 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.5% | 100% | 11.43 cm | 1.80 cm - 1.80 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:18 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.5% | 100% | 11.44 cm | 1.94 cm - 1.94 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:20 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.2% | 100% | 11.23 cm | 1.67 cm - 1.67 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:21 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.7% | 100% | 11.27 cm | 1.75 cm - 1.75 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:22 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.5% | 100% | 11.31 cm | 1.87 cm - 1.87 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:23 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 1.60 cm | 1.60 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.3% | 100% | 11.31 cm | 1.98 cm - 1.98 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:24 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.7% | 100% | 11.30 cm | 2.03 cm - 2.03 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:27 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 8.88 cm | 0.66 cm - 0.66 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:28 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 9.05 cm | 0.90 cm - 0.90 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:28 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10% | 100% | 9.27 cm | 1.76 cm - 1.76 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:29 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.6% | 100% | 9.31 cm | 2.00 cm - 2.00 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:30 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 6.50 cm | 6.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.4% | 100% | 9.39 cm | 2.46 cm - 2.46 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:32 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 8.87 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:33 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.7% | 100% | 9.09 cm | 1.22 cm - 1.22 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:34 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.2% | 100% | 9.22 cm | 1.90 cm - 1.90 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:35 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.9% | 100% | 9.26 cm | 2.15 cm - 2.15 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:35 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.2% | 100% | 9.28 cm | 2.37 cm - 2.37 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:40 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 8.88 cm | 0.92 cm - 0.92 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:41 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.9% | 100% | 8.94 cm | 1.00 cm - 1.00 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:41 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.2% | 100% | 9.06 cm | 1.37 cm - 1.37 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:42 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11% | 100% | 9.13 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:43 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.6% | 100% | 9.14 cm | 2.06 cm - 2.06 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:45 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.4% | 100% | 8.89 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:46 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7% | 100% | 8.93 cm | 1.33 cm - 1.33 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:46 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.4% | 100% | 9.00 cm | 1.61 cm - 1.61 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:47 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.1% | 100% | 9.02 cm | 1.82 cm - 1.82 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:48 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.2% | 100% | 9.02 cm | 1.96 cm - 1.96 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:50 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.3% | 100% | 8.89 cm | 1.68 cm - 1.68 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:50 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.8% | 100% | 8.91 cm | 1.75 cm - 1.75 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:51 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.6% | 100% | 8.94 cm | 1.86 cm - 1.86 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:52 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 1.60 cm | 1.60 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.5% | 100% | 8.94 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:53 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 8.00 cm | 8.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.8% | 100% | 8.93 cm | 2.01 cm - 2.01 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:56 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2% | 100% | 7.77 cm | 0.67 cm - 0.67 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:57 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.3% | 100% | 7.92 cm | 0.90 cm - 0.90 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:58 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.7% | 100% | 8.13 cm | 1.75 cm - 1.75 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:58 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.3% | 100% | 8.17 cm | 1.98 cm - 1.98 cm | 0.00 cm
| TR3 | 2018-09-07 | 13:59 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 6.50 cm | 6.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.9% | 100% | 8.26 cm | 2.43 cm - 2.43 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:01 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.8% | 100% | 7.76 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:02 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.4% | 100% | 7.95 cm | 1.23 cm - 1.23 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:03 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.9% | 100% | 8.08 cm | 1.89 cm - 1.89 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:03 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.6% | 100% | 8.12 cm | 2.14 cm - 2.14 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:04 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.8% | 100% | 8.15 cm | 2.35 cm - 2.35 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:06 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.5% | 100% | 7.77 cm | 0.93 cm - 0.93 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:07 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.3% | 100% | 7.82 cm | 1.00 cm - 1.00 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:08 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9% | 100% | 7.93 cm | 1.37 cm - 1.37 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:08 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.9% | 100% | 7.99 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:09 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.5% | 100% | 8.02 cm | 2.05 cm - 2.05 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:11 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.3% | 100% | 7.78 cm | 1.26 cm - 1.26 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:12 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.1% | 100% | 7.81 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:13 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.4% | 100% | 7.87 cm | 1.62 cm - 1.62 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:13 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.1% | 100% | 7.89 cm | 1.83 cm - 1.83 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:14 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.2% | 100% | 7.90 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:16 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.4% | 100% | 7.77 cm | 1.69 cm - 1.69 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:17 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.9% | 100% | 7.79 cm | 1.76 cm - 1.76 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:17 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.7% | 100% | 7.83 cm | 1.87 cm - 1.87 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:18 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 1.60 cm | 1.60 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.4% | 100% | 7.83 cm | 1.98 cm - 1.98 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:19 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.8% | 100% | 7.82 cm | 2.02 cm - 2.02 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:23 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.4% | 100% | 6.72 cm | 0.62 cm - 0.62 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:24 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.1% | 100% | 6.86 cm | 0.87 cm - 0.87 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:24 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.4% | 100% | 7.06 cm | 1.73 cm - 1.73 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:25 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.9% | 100% | 7.10 cm | 1.96 cm - 1.96 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:26 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.2% | 100% | 7.20 cm | 2.39 cm - 2.39 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:28 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.5% | 100% | 6.72 cm | 0.72 cm - 0.72 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:29 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.9% | 100% | 6.89 cm | 1.20 cm - 1.20 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:30 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.6% | 100% | 7.02 cm | 1.86 cm - 1.86 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:31 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.1% | 100% | 7.06 cm | 2.11 cm - 2.11 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:32 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.3% | 100% | 7.09 cm | 2.32 cm - 2.32 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:34 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.8% | 100% | 6.73 cm | 0.88 cm - 0.88 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:34 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.8% | 100% | 6.78 cm | 0.97 cm - 0.97 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:35 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.7% | 100% | 6.88 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:36 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.7% | 100% | 6.94 cm | 1.83 cm - 1.83 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:36 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.2% | 100% | 6.95 cm | 2.05 cm - 2.05 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:38 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.6% | 100% | 6.75 cm | 1.23 cm - 1.23 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:39 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.3% | 100% | 6.78 cm | 1.30 cm - 1.30 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:40 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13% | 100% | 6.84 cm | 1.59 cm - 1.59 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:41 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.7% | 100% | 6.86 cm | 1.80 cm - 1.80 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:41 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16% | 100% | 6.86 cm | 1.96 cm - 1.96 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:43 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.3% | 100% | 6.80 cm | 1.61 cm - 1.61 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:44 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.9% | 100% | 6.83 cm | 1.68 cm - 1.68 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:45 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.7% | 100% | 6.86 cm | 1.80 cm - 1.80 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:45 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.60 cm | 1.60 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.5% | 100% | 6.87 cm | 1.91 cm - 1.91 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:46 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.1% | 100% | 6.85 cm | 1.98 cm - 1.98 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:50 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.2% | 100% | 5.56 cm | 0.66 cm - 0.66 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:51 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.6% | 100% | 5.68 cm | 0.89 cm - 0.89 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:52 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.5% | 100% | 5.88 cm | 1.70 cm - 1.70 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:52 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.7% | 100% | 5.94 cm | 1.91 cm - 1.91 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:53 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.6% | 100% | 6.10 cm | 2.31 cm - 2.31 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:55 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.7% | 100% | 5.56 cm | 0.75 cm - 0.75 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:56 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.6% | 100% | 5.71 cm | 1.21 cm - 1.21 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:56 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.6% | 100% | 5.85 cm | 1.83 cm - 1.83 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:57 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.7% | 100% | 5.93 cm | 2.05 cm - 2.05 cm | 0.00 cm
| TR3 | 2018-09-07 | 14:58 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.7% | 100% | 5.98 cm | 2.24 cm - 2.24 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:00 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.3% | 100% | 5.56 cm | 0.92 cm - 0.92 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:00 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.4% | 100% | 5.61 cm | 0.99 cm - 0.99 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:01 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.2% | 100% | 5.70 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:02 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.6% | 100% | 5.80 cm | 1.79 cm - 1.79 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:02 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18% | 100% | 5.84 cm | 1.99 cm - 1.99 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:04 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.7% | 100% | 5.59 cm | 1.26 cm - 1.26 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:05 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.4% | 100% | 5.62 cm | 1.34 cm - 1.34 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:05 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.7% | 100% | 5.68 cm | 1.60 cm - 1.60 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:06 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.1% | 100% | 5.71 cm | 1.80 cm - 1.80 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:07 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.1% | 100% | 5.73 cm | 1.94 cm - 1.94 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:09 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16% | 100% | 5.65 cm | 1.63 cm - 1.63 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:09 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.5% | 100% | 5.68 cm | 1.69 cm - 1.69 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:10 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.1% | 100% | 5.72 cm | 1.80 cm - 1.80 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:11 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.60 cm | 1.60 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.8% | 100% | 5.73 cm | 1.90 cm - 1.90 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:11 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 5.00 cm | 5.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.2% | 100% | 5.72 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:15 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.7% | 100% | 5.10 cm | 0.64 cm - 0.64 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:16 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.4% | 100% | 5.20 cm | 0.88 cm - 0.88 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:17 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.3% | 100% | 5.40 cm | 1.68 cm - 1.68 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:17 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.4% | 100% | 5.47 cm | 1.87 cm - 1.87 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:18 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.1% | 100% | 5.66 cm | 2.26 cm - 2.26 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:20 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.6% | 100% | 5.09 cm | 0.74 cm - 0.74 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:21 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.5% | 100% | 5.24 cm | 1.19 cm - 1.19 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:21 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.4% | 100% | 5.40 cm | 1.80 cm - 1.80 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:22 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.5% | 100% | 5.47 cm | 2.01 cm - 2.01 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:23 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.2% | 100% | 5.56 cm | 2.19 cm - 2.19 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:25 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.2% | 100% | 5.10 cm | 0.90 cm - 0.90 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:26 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.2% | 100% | 5.14 cm | 0.97 cm - 0.97 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:26 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.1% | 100% | 5.23 cm | 1.33 cm - 1.33 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:27 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.5% | 100% | 5.33 cm | 1.76 cm - 1.76 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:28 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.6% | 100% | 5.39 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:30 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.7% | 100% | 5.12 cm | 1.24 cm - 1.24 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:31 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.3% | 100% | 5.15 cm | 1.32 cm - 1.32 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:32 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.5% | 100% | 5.21 cm | 1.58 cm - 1.58 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:32 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.9% | 100% | 5.25 cm | 1.77 cm - 1.77 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:33 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.8% | 100% | 5.27 cm | 1.91 cm - 1.91 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:35 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.9% | 100% | 5.18 cm | 1.60 cm - 1.60 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:35 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.2% | 100% | 5.21 cm | 1.66 cm - 1.66 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:36 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.9% | 100% | 5.25 cm | 1.77 cm - 1.77 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:37 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.60 cm | 1.60 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.5% | 100% | 5.27 cm | 1.87 cm - 1.87 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:37 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.50 cm | 4.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19% | 100% | 5.27 cm | 1.92 cm - 1.92 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:42 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.3% | 100% | 4.47 cm | 0.64 cm - 0.64 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:42 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.1% | 100% | 4.56 cm | 0.88 cm - 0.88 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:43 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.8% | 100% | 4.73 cm | 1.65 cm - 1.65 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:44 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.7% | 100% | 4.80 cm | 1.83 cm - 1.83 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:44 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.9% | 0% | 5.01 cm | 2.19 cm - 2.19 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:46 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.2% | 100% | 4.45 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:47 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.2% | 100% | 4.57 cm | 1.18 cm - 1.18 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:48 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.7% | 100% | 4.72 cm | 1.75 cm - 1.75 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:48 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.5% | 0% | 4.81 cm | 1.94 cm - 1.94 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:49 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.1% | 0% | 4.92 cm | 2.11 cm - 2.11 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:51 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.1% | 100% | 4.46 cm | 0.89 cm - 0.89 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:52 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.1% | 100% | 4.49 cm | 0.97 cm - 0.97 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:52 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.7% | 100% | 4.56 cm | 1.31 cm - 1.31 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:53 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.7% | 100% | 4.66 cm | 1.70 cm - 1.70 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:54 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.7% | 0% | 4.72 cm | 1.88 cm - 1.88 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:55 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.6% | 100% | 4.45 cm | 1.23 cm - 1.23 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:56 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.2% | 100% | 4.47 cm | 1.31 cm - 1.31 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:57 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18% | 100% | 4.53 cm | 1.54 cm - 1.54 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:57 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.2% | 0% | 4.57 cm | 1.72 cm - 1.72 cm | 0.00 cm
| TR3 | 2018-09-07 | 15:58 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.9% | 0% | 4.60 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:00 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.6% | 0% | 4.49 cm | 1.58 cm - 1.58 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:01 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.9% | 0% | 4.52 cm | 1.63 cm - 1.63 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:01 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.4% | 0% | 4.55 cm | 1.73 cm - 1.73 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:02 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.60 cm | 1.60 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.9% | 0% | 4.58 cm | 1.82 cm - 1.82 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:03 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.1% | 0% | 4.60 cm | 1.86 cm - 1.86 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:06 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.4% | 100% | 3.45 cm | 0.63 cm - 0.63 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:06 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.1% | 100% | 3.54 cm | 0.86 cm - 0.86 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:07 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.3% | 0% | 3.82 cm | 1.55 cm - 1.55 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:08 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.8% | 0% | 3.94 cm | 1.72 cm - 1.72 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:08 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.7% | 0% | 4.27 cm | 2.07 cm - 2.07 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:10 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.5% | 100% | 3.45 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:11 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.8% | 0% | 3.58 cm | 1.14 cm - 1.14 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:12 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.9% | 0% | 3.86 cm | 1.64 cm - 1.64 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:12 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.5% | 0% | 4.03 cm | 1.82 cm - 1.82 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:13 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.8% | 0% | 4.21 cm | 1.98 cm - 1.98 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:15 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.2% | 0% | 3.45 cm | 0.88 cm - 0.88 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:15 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.2% | 0% | 3.49 cm | 0.95 cm - 0.95 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:16 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.1% | 0% | 3.60 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:17 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20% | 0% | 3.82 cm | 1.60 cm - 1.60 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:17 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.6% | 0% | 3.96 cm | 1.76 cm - 1.76 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:20 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.9% | 0% | 3.50 cm | 1.18 cm - 1.18 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:21 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.4% | 0% | 3.54 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:22 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.6% | 0% | 3.66 cm | 1.45 cm - 1.45 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:22 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.3% | 0% | 3.78 cm | 1.60 cm - 1.60 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:23 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.8% | 0% | 3.88 cm | 1.71 cm - 1.71 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:25 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.1% | 0% | 3.64 cm | 1.49 cm - 1.49 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:25 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.3% | 0% | 3.69 cm | 1.53 cm - 1.53 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:26 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.6% | 0% | 3.77 cm | 1.61 cm - 1.61 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:27 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.60 cm | 1.60 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.8% | 0% | 3.84 cm | 1.69 cm - 1.69 cm | 0.00 cm
| TR3 | 2018-09-07 | 16:27 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 21% | 0% | 3.88 cm | 1.73 cm - 1.73 cm | 0.00 cm

## TR3 circle Profile Parameters
ParameterGroup
Electrons profile analysis (current protocol: 'DataBook')
| Radiation device | Date | Time | Radiation type | SSD | Energy | Applicator | Scan type | Start position [De] | End position [De] | Field size inline | Field size crossline | Gantry angle | Collimator angle | Field detector | Penumbra margin | Scan direction | Scan mode | Flatness | Symmetry | FieldWidth | Penumbra | Center
| TR3 | 2018-09-05 | 14:14 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.8% | 100% | 3.41 cm | 0.71 cm - 0.71 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:15 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.3% | 0% | 3.40 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:15 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.9% | 0% | 3.38 cm | 0.96 cm - 0.96 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:16 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.60 cm | 1.60 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.2% | 0% | 3.35 cm | 1.14 cm - 1.14 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:17 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.4% | 0% | 3.32 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:48 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.6% | 100% | 3.92 cm | 0.69 cm - 0.69 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:49 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.2% | 100% | 3.91 cm | 0.77 cm - 0.77 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:50 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.1% | 100% | 3.89 cm | 0.96 cm - 0.96 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:50 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.60 cm | 1.60 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.8% | 0% | 3.84 cm | 1.17 cm - 1.17 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:51 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.3% | 0% | 3.80 cm | 1.29 cm - 1.29 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:17 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.6% | 100% | 5.97 cm | 0.72 cm - 0.72 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:18 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.8% | 100% | 5.97 cm | 0.81 cm - 0.81 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:19 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.5% | 100% | 5.95 cm | 1.01 cm - 1.01 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:19 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.60 cm | 1.60 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.9% | 100% | 5.91 cm | 1.23 cm - 1.23 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:20 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.4% | 100% | 5.87 cm | 1.36 cm - 1.36 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:35 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.1% | 100% | 7.43 cm | 0.72 cm - 0.72 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:36 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.2% | 100% | 7.44 cm | 0.81 cm - 0.81 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:36 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.4% | 100% | 7.44 cm | 1.00 cm - 1.00 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:37 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.60 cm | 1.60 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.6% | 100% | 7.41 cm | 1.21 cm - 1.21 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:38 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.9% | 100% | 7.38 cm | 1.32 cm - 1.32 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:26 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.3% | 100% | 8.96 cm | 0.69 cm - 0.69 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:27 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.7% | 100% | 8.98 cm | 0.77 cm - 0.77 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:28 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.2% | 100% | 8.99 cm | 0.94 cm - 0.94 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:29 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.60 cm | 1.60 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.8% | 100% | 8.98 cm | 1.15 cm - 1.15 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:30 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.1% | 100% | 8.95 cm | 1.30 cm - 1.30 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:07 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 11.86 cm | 0.71 cm - 0.71 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:08 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 11.90 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:09 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 1.20 cm | 1.20 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.4% | 100% | 11.93 cm | 0.97 cm - 0.97 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:10 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 1.60 cm | 1.60 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 11.93 cm | 1.17 cm - 1.17 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:11 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.9% | 100% | 11.91 cm | 1.31 cm - 1.31 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:09 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.1% | 100% | 3.44 cm | 0.59 cm - 0.59 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:10 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.2% | 100% | 3.43 cm | 0.69 cm - 0.69 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:11 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.6% | 0% | 3.40 cm | 1.03 cm - 1.03 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:11 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.9% | 0% | 3.42 cm | 1.27 cm - 1.27 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:12 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.4% | 0% | 3.45 cm | 1.44 cm - 1.44 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:44 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7% | 100% | 3.94 cm | 0.58 cm - 0.58 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:44 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.3% | 100% | 3.94 cm | 0.68 cm - 0.68 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:45 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15% | 100% | 3.91 cm | 1.05 cm - 1.05 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:46 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18% | 0% | 3.88 cm | 1.31 cm - 1.31 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:46 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.6% | 0% | 3.86 cm | 1.50 cm - 1.50 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:12 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 5.98 cm | 0.59 cm - 0.59 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:13 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.8% | 100% | 6.00 cm | 0.70 cm - 0.70 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:14 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.3% | 100% | 5.98 cm | 1.11 cm - 1.11 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:15 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.7% | 100% | 5.92 cm | 1.42 cm - 1.42 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:15 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.8% | 100% | 5.86 cm | 1.63 cm - 1.63 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:30 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 7.47 cm | 0.59 cm - 0.59 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:31 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.7% | 100% | 7.49 cm | 0.70 cm - 0.70 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:32 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.4% | 100% | 7.49 cm | 1.11 cm - 1.11 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:32 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.6% | 100% | 7.45 cm | 1.41 cm - 1.41 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:33 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.4% | 100% | 7.40 cm | 1.61 cm - 1.61 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:18 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 9.00 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:19 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 9.03 cm | 0.65 cm - 0.65 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:20 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.3% | 100% | 9.07 cm | 1.02 cm - 1.02 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:21 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.1% | 100% | 9.05 cm | 1.32 cm - 1.32 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:21 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.1% | 100% | 9.01 cm | 1.54 cm - 1.54 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:01 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 11.92 cm | 0.56 cm - 0.56 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:02 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 1.20 cm | 1.20 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 11.97 cm | 0.66 cm - 0.66 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:03 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.2% | 100% | 12.05 cm | 1.04 cm - 1.04 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:04 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.9% | 100% | 12.05 cm | 1.33 cm - 1.33 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:04 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.3% | 100% | 12.04 cm | 1.55 cm - 1.55 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:05 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.7% | 100% | 3.46 cm | 0.44 cm - 0.44 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:05 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.3% | 100% | 3.48 cm | 0.53 cm - 0.53 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:06 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.6% | 0% | 3.45 cm | 0.97 cm - 0.97 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:07 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.6% | 0% | 3.53 cm | 1.41 cm - 1.41 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:07 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.3% | 0% | 3.66 cm | 1.57 cm - 1.57 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:39 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 3.97 cm | 0.43 cm - 0.43 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:40 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.4% | 100% | 3.98 cm | 0.52 cm - 0.52 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:40 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14% | 100% | 3.96 cm | 0.99 cm - 0.99 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:41 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.7% | 0% | 3.99 cm | 1.46 cm - 1.46 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:42 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20% | 0% | 4.05 cm | 1.65 cm - 1.65 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:07 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 5.98 cm | 0.44 cm - 0.44 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:08 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 1.20 cm | 1.20 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3% | 100% | 6.02 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:09 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.2% | 100% | 6.04 cm | 1.04 cm - 1.04 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:10 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.1% | 100% | 5.98 cm | 1.61 cm - 1.61 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:10 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.2% | 100% | 5.93 cm | 1.86 cm - 1.86 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:25 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 7.47 cm | 0.45 cm - 0.45 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:26 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 7.52 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:27 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.5% | 100% | 7.56 cm | 1.04 cm - 1.04 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:27 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.4% | 100% | 7.51 cm | 1.63 cm - 1.63 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:28 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.5% | 100% | 7.45 cm | 1.88 cm - 1.88 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:12 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 8.99 cm | 0.42 cm - 0.42 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:13 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 1.20 cm | 1.20 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 9.05 cm | 0.51 cm - 0.51 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:14 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.9% | 100% | 9.13 cm | 0.97 cm - 0.97 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:15 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.1% | 100% | 9.11 cm | 1.56 cm - 1.56 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:16 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.3% | 100% | 9.08 cm | 1.82 cm - 1.82 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:55 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.2% | 100% | 11.89 cm | 0.44 cm - 0.44 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:56 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 1.20 cm | 1.20 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 11.98 cm | 0.52 cm - 0.52 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:56 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 12.12 cm | 0.99 cm - 0.99 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:57 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 3.50 cm | 3.50 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.4% | 100% | 12.17 cm | 1.56 cm - 1.56 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:58 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.1% | 100% | 12.15 cm | 1.82 cm - 1.82 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:00 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 3.48 cm | 0.40 cm - 0.40 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:01 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.1% | 100% | 3.49 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:01 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.9% | 0% | 3.53 cm | 1.33 cm - 1.33 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:02 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.9% | 0% | 3.61 cm | 1.52 cm - 1.52 cm | 0.00 cm
| TR3 | 2018-09-05 | 14:03 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.9% | 0% | 3.91 cm | 1.86 cm - 1.86 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:34 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2% | 100% | 3.98 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:35 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.2% | 100% | 4.01 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:35 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.7% | 0% | 4.01 cm | 1.36 cm - 1.36 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:36 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.2% | 0% | 4.06 cm | 1.57 cm - 1.57 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:37 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.6% | 0% | 4.28 cm | 1.92 cm - 1.92 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:02 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 5.99 cm | 0.40 cm - 0.40 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:03 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.6% | 100% | 6.08 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:04 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.5% | 100% | 6.07 cm | 1.46 cm - 1.46 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:04 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.6% | 100% | 6.06 cm | 1.72 cm - 1.72 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:05 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.6% | 100% | 6.05 cm | 2.17 cm - 2.17 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:20 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 7.47 cm | 0.40 cm - 0.40 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:21 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.6% | 100% | 7.61 cm | 0.80 cm - 0.80 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:22 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.8% | 100% | 7.61 cm | 1.49 cm - 1.49 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:22 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.1% | 100% | 7.59 cm | 1.76 cm - 1.76 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:23 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.7% | 100% | 7.53 cm | 2.26 cm - 2.26 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:06 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 9.00 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:07 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 9.16 cm | 0.74 cm - 0.74 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:08 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.7% | 100% | 9.21 cm | 1.42 cm - 1.42 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:09 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.9% | 100% | 9.21 cm | 1.69 cm - 1.69 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:10 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.9% | 100% | 9.14 cm | 2.23 cm - 2.23 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:46 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.2% | 100% | 11.89 cm | 0.40 cm - 0.40 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:47 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 12.13 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:48 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.6% | 100% | 12.26 cm | 1.45 cm - 1.45 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:49 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.4% | 100% | 12.27 cm | 1.73 cm - 1.73 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:49 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.8% | 100% | 12.26 cm | 2.24 cm - 2.24 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:55 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 3.49 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:56 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.1% | 100% | 3.53 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:57 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.2% | 0% | 3.55 cm | 1.29 cm - 1.29 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:57 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.1% | 0% | 3.70 cm | 1.62 cm - 1.62 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:58 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 3.50 cm | 3.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.6% | 0% | 4.01 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:29 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 3.99 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:30 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.4% | 100% | 4.04 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:30 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.1% | 0% | 4.05 cm | 1.32 cm - 1.32 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:31 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.5% | 0% | 4.15 cm | 1.67 cm - 1.67 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:32 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.3% | 0% | 4.39 cm | 2.03 cm - 2.03 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:57 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 6.01 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:58 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.4% | 100% | 6.09 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:58 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.4% | 100% | 6.14 cm | 1.38 cm - 1.38 cm | 0.00 cm
| TR3 | 2018-09-05 | 12:59 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.1% | 100% | 6.14 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR3 | 2018-09-05 | 13:00 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.7% | 100% | 6.19 cm | 2.28 cm - 2.28 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:15 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 7.48 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:16 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 7.59 cm | 0.56 cm - 0.56 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:17 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.9% | 100% | 7.68 cm | 1.41 cm - 1.41 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:17 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.8% | 100% | 7.68 cm | 1.88 cm - 1.88 cm | 0.00 cm
| TR3 | 2018-09-05 | 10:18 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 6.50 cm | 6.50 cm | 7.50 cm | 7.50 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.9% | 100% | 7.66 cm | 2.36 cm - 2.36 cm | 0.00 cm
| TR3 | 2018-09-04 | 16:59 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 9.01 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:00 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 9.14 cm | 0.52 cm - 0.52 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:01 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7% | 100% | 9.28 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:02 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.8% | 100% | 9.29 cm | 1.83 cm - 1.83 cm | 0.00 cm
| TR3 | 2018-09-04 | 17:02 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 6.50 cm | 6.50 cm | 9.00 cm | 9.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.3% | 100% | 9.25 cm | 2.35 cm - 2.35 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:39 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 11.90 cm | 0.37 cm - 0.37 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:40 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 12.08 cm | 0.53 cm - 0.53 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:41 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.5% | 100% | 12.32 cm | 1.39 cm - 1.39 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:41 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.6% | 100% | 12.37 cm | 1.89 cm - 1.89 cm | 0.00 cm
| TR3 | 2018-09-04 | 15:42 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 6.50 cm | 6.50 cm | 12.00 cm | 12.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.8% | 100% | 12.37 cm | 2.43 cm - 2.43 cm | 0.00 cm

## TR2 Profile Parameters
ParameterGroup
Electrons profile analysis (current protocol: 'DataBook')
| Radiation device | Date | Time | Radiation type | SSD | Energy | Applicator | Scan type | Start position [De] | End position [De] | Field size inline | Field size crossline | Gantry angle | Collimator angle | Field detector | Penumbra margin | Scan direction | Scan mode | Flatness | Symmetry | FieldWidth | Penumbra | Center
| TR2 | 2017-08-28 | 16:58 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.5% | 0% | 3.13 cm | 0.75 cm - 0.75 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:59 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.7% | 0% | 3.14 cm | 0.82 cm - 0.82 cm | 0.00 cm
| TR2 | 2017-08-28 | 17:00 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.40 cm | 1.40 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17% | 0% | 3.17 cm | 1.00 cm - 1.00 cm | 0.00 cm
| TR2 | 2017-08-28 | 17:00 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.3% | 0% | 3.21 cm | 1.21 cm - 1.21 cm | 0.00 cm
| TR2 | 2017-08-28 | 17:01 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.7% | 0% | 3.21 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR2 | 2017-08-28 | 17:01 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19% | 0% | 3.20 cm | 1.19 cm - 1.19 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:12 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.2% | 100% | 4.03 cm | 0.75 cm - 0.75 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:12 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.5% | 100% | 4.04 cm | 0.83 cm - 0.83 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:13 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.40 cm | 1.40 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.4% | 100% | 4.04 cm | 1.03 cm - 1.03 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:14 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.4% | 0% | 4.02 cm | 1.27 cm - 1.27 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:14 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.9% | 0% | 4.02 cm | 1.31 cm - 1.31 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:15 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.5% | 100% | 4.06 cm | 1.23 cm - 1.23 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:55 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.7% | 100% | 6.83 cm | 1.67 cm - 1.67 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:56 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14% | 100% | 6.84 cm | 1.72 cm - 1.72 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:57 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.40 cm | 1.40 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15% | 100% | 6.86 cm | 1.85 cm - 1.85 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:58 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16% | 100% | 6.83 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:58 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16% | 100% | 6.80 cm | 1.96 cm - 1.96 cm | 0.00 cm
| TR2 | 2017-08-29 | 17:00 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.3% | 100% | 6.81 cm | 1.90 cm - 1.90 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:23 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.9% | 100% | 6.09 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:23 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7% | 100% | 6.10 cm | 0.85 cm - 0.85 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:24 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.40 cm | 1.40 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.6% | 100% | 6.12 cm | 1.07 cm - 1.07 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:25 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.3% | 100% | 6.10 cm | 1.32 cm - 1.32 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:25 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.8% | 100% | 6.09 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:26 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.7% | 100% | 6.12 cm | 1.27 cm - 1.27 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:10 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.1% | 0% | 2.97 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:11 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.1% | 0% | 3.00 cm | 0.87 cm - 0.87 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:11 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.40 cm | 1.40 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.1% | 0% | 3.03 cm | 1.04 cm - 1.04 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:12 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.9% | 0% | 3.08 cm | 1.22 cm - 1.22 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:13 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20% | 0% | 3.07 cm | 1.24 cm - 1.24 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:13 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.1% | 0% | 3.10 cm | 1.19 cm - 1.19 cm | 0.00 cm
| TR2 | 2017-08-25 | 16:20 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4% | 100% | 6.99 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR2 | 2017-08-25 | 16:21 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.8% | 100% | 7.00 cm | 0.84 cm - 0.84 cm | 0.00 cm
| TR2 | 2017-08-25 | 16:22 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.40 cm | 1.40 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7% | 100% | 7.02 cm | 1.02 cm - 1.02 cm | 0.00 cm
| TR2 | 2017-08-25 | 16:23 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10% | 100% | 7.03 cm | 1.29 cm - 1.29 cm | 0.00 cm
| TR2 | 2017-08-25 | 16:24 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.6% | 100% | 7.02 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR2 | 2017-08-25 | 16:24 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.8% | 100% | 7.02 cm | 1.28 cm - 1.28 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:48 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.5% | 100% | 11.25 cm | 1.72 cm - 1.72 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:49 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.8% | 100% | 11.29 cm | 1.78 cm - 1.78 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:50 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 1.40 cm | 1.40 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.7% | 100% | 11.31 cm | 1.90 cm - 1.90 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:51 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.7% | 100% | 11.29 cm | 2.02 cm - 2.02 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:52 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.4% | 100% | 11.25 cm | 1.98 cm - 1.98 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:53 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.8% | 100% | 11.25 cm | 1.92 cm - 1.92 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:29 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 11.06 cm | 0.78 cm - 0.78 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:30 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 1.00 cm | 1.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2% | 100% | 11.09 cm | 0.86 cm - 0.86 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:31 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 1.40 cm | 1.40 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3% | 100% | 11.12 cm | 1.07 cm - 1.07 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:32 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.2% | 100% | 11.13 cm | 1.30 cm - 1.30 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:33 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.2% | 100% | 11.11 cm | 1.32 cm - 1.32 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:34 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.6% | 100% | 11.12 cm | 1.23 cm - 1.23 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:27 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.4% | 100% | 16.83 cm | 1.78 cm - 1.78 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:28 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 1.00 cm | 1.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.7% | 100% | 16.87 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:29 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 1.40 cm | 1.40 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.1% | 100% | 16.90 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:31 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.5% | 100% | 16.88 cm | 2.08 cm - 2.08 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:32 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.2% | 100% | 16.81 cm | 2.02 cm - 2.02 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:33 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.2% | 100% | 16.79 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:15 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.3% | 100% | 15.18 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:17 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 1.00 cm | 1.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.4% | 100% | 15.22 cm | 0.88 cm - 0.88 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:18 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 1.40 cm | 1.40 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.7% | 100% | 15.26 cm | 1.11 cm - 1.11 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:19 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3% | 100% | 15.26 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:20 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.8% | 100% | 15.23 cm | 1.37 cm - 1.37 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:21 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.9% | 100% | 15.22 cm | 1.28 cm - 1.28 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:14 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 0.80 cm | 0.80 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 20.27 cm | 0.81 cm - 0.81 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:15 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 1.00 cm | 1.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 20.31 cm | 0.90 cm - 0.90 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:17 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 1.40 cm | 1.40 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 20.38 cm | 1.12 cm - 1.12 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:18 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 20.37 cm | 1.34 cm - 1.34 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:19 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 20.34 cm | 1.34 cm - 1.34 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:20 | Electrons | 100.00 cm | 6 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 20.32 cm | 1.27 cm - 1.27 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:10 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 0.80 cm | 0.80 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 22.40 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:11 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 1.00 cm | 1.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 22.46 cm | 1.88 cm - 1.88 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:13 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 1.40 cm | 1.40 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 22.51 cm | 2.01 cm - 2.01 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:14 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.4% | 100% | 22.48 cm | 2.11 cm - 2.11 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:15 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.1% | 100% | 22.41 cm | 2.04 cm - 2.04 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:17 | Electrons | 110.00 cm | 6 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 22.35 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:05 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 0.80 cm | 0.80 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 21.06 cm | 0.77 cm - 0.77 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:06 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 1.00 cm | 1.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 21.11 cm | 0.85 cm - 0.85 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:07 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 1.40 cm | 1.40 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 21.19 cm | 1.04 cm - 1.04 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:08 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 21.22 cm | 1.29 cm - 1.29 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:10 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 21.20 cm | 1.33 cm - 1.33 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:11 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 21.14 cm | 1.26 cm - 1.26 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:24 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 0.80 cm | 0.80 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.24 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:25 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 1.00 cm | 1.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.31 cm | 0.86 cm - 0.86 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:27 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 1.40 cm | 1.40 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.39 cm | 1.07 cm - 1.07 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:28 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 25.43 cm | 1.34 cm - 1.34 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:29 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.37 cm | 1.38 cm - 1.38 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:31 | Electrons | 100.00 cm | 6 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 25.32 cm | 1.29 cm - 1.29 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:33 | Electrons | 110.00 cm | 6 MeV | 25x25 | Crossline | 0.80 cm | 0.80 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.4% | 100% | 27.99 cm | 1.87 cm - 1.87 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:34 | Electrons | 110.00 cm | 6 MeV | 25x25 | Crossline | 1.00 cm | 1.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.3% | 100% | 28.04 cm | 1.93 cm - 1.93 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:36 | Electrons | 110.00 cm | 6 MeV | 25x25 | Crossline | 1.40 cm | 1.40 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.4% | 100% | 28.12 cm | 2.04 cm - 2.04 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:37 | Electrons | 110.00 cm | 6 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 28.06 cm | 2.14 cm - 2.14 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:39 | Electrons | 110.00 cm | 6 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.3% | 100% | 27.97 cm | 2.06 cm - 2.06 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:40 | Electrons | 110.00 cm | 6 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 27.90 cm | 2.01 cm - 2.01 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:19 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.8% | 100% | 4.02 cm | 0.67 cm - 0.67 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:20 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.80 cm | 1.80 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.7% | 0% | 4.02 cm | 1.22 cm - 1.22 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:20 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.20 cm | 2.20 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.8% | 0% | 4.01 cm | 1.31 cm - 1.31 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:21 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.70 cm | 2.70 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.5% | 0% | 4.01 cm | 1.29 cm - 1.29 cm | 0.00 cm
| TR2 | 2017-09-19 | 14:55 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.4% | 100% | 4.48 cm | 1.56 cm - 1.56 cm | 0.00 cm
| TR2 | 2017-09-19 | 14:55 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.8% | 100% | 4.51 cm | 1.61 cm - 1.61 cm | 0.00 cm
| TR2 | 2017-09-19 | 14:56 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19% | 0% | 4.52 cm | 1.66 cm - 1.66 cm | 0.00 cm
| TR2 | 2017-09-19 | 14:57 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.40 cm | 1.40 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.6% | 0% | 4.56 cm | 1.76 cm - 1.76 cm | 0.00 cm
| TR2 | 2017-09-19 | 14:57 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.80 cm | 1.80 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20% | 0% | 4.57 cm | 1.83 cm - 1.83 cm | 0.00 cm
| TR2 | 2017-09-19 | 14:58 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.1% | 0% | 4.57 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR2 | 2017-09-19 | 14:59 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 2.20 cm | 2.20 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.1% | 0% | 4.56 cm | 1.85 cm - 1.85 cm | 0.00 cm
| TR2 | 2017-09-19 | 14:59 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.9% | 0% | 4.54 cm | 1.83 cm - 1.83 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:00 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 2.70 cm | 2.70 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20% | 0% | 4.53 cm | 1.81 cm - 1.81 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:38 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.8% | 100% | 6.08 cm | 0.68 cm - 0.68 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:38 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 1.80 cm | 1.80 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.7% | 100% | 6.12 cm | 1.27 cm - 1.27 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:39 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.20 cm | 2.20 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.7% | 100% | 6.11 cm | 1.36 cm - 1.36 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:40 | Electrons | 100.00 cm | 6 MeV | 6x6 | Crossline | 2.70 cm | 2.70 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.2% | 100% | 6.10 cm | 1.31 cm - 1.31 cm | 0.00 cm
| TR2 | 2017-09-19 | 13:54 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.9% | 100% | 6.79 cm | 1.59 cm - 1.59 cm | 0.00 cm
| TR2 | 2017-09-19 | 13:54 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 1.80 cm | 1.80 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.8% | 100% | 6.85 cm | 1.94 cm - 1.94 cm | 0.00 cm
| TR2 | 2017-09-19 | 13:55 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 2.20 cm | 2.20 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16% | 100% | 6.83 cm | 1.96 cm - 1.96 cm | 0.00 cm
| TR2 | 2017-09-19 | 13:56 | Electrons | 110.00 cm | 6 MeV | 6x6 | Crossline | 2.70 cm | 2.70 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.6% | 100% | 6.82 cm | 1.89 cm - 1.89 cm | 0.00 cm
| TR2 | 2017-09-18 | 14:19 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.80 cm | 1.80 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.3% | 100% | 10.20 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR2 | 2017-09-18 | 14:20 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 2.20 cm | 2.20 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.8% | 100% | 10.19 cm | 1.33 cm - 1.33 cm | 0.00 cm
| TR2 | 2017-09-18 | 14:21 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 2.70 cm | 2.70 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.4% | 100% | 10.17 cm | 1.28 cm - 1.28 cm | 0.00 cm
| TR2 | 2017-09-19 | 12:03 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7% | 100% | 11.22 cm | 1.65 cm - 1.65 cm | 0.00 cm
| TR2 | 2017-09-19 | 12:04 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 1.80 cm | 1.80 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.5% | 100% | 11.29 cm | 2.01 cm - 2.01 cm | 0.00 cm
| TR2 | 2017-09-19 | 12:05 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 2.20 cm | 2.20 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.6% | 100% | 11.27 cm | 2.01 cm - 2.01 cm | 0.00 cm
| TR2 | 2017-09-19 | 12:06 | Electrons | 110.00 cm | 6 MeV | 10x10 | Crossline | 2.70 cm | 2.70 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.1% | 100% | 11.24 cm | 1.93 cm - 1.93 cm | 0.00 cm
| TR2 | 2017-09-18 | 12:08 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.3% | 100% | 15.13 cm | 0.70 cm - 0.70 cm | 0.00 cm
| TR2 | 2017-09-18 | 12:09 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 1.80 cm | 1.80 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3% | 100% | 15.25 cm | 1.30 cm - 1.30 cm | 0.00 cm
| TR2 | 2017-09-18 | 12:10 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 2.20 cm | 2.20 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3% | 100% | 15.24 cm | 1.38 cm - 1.38 cm | 0.00 cm
| TR2 | 2017-09-18 | 12:12 | Electrons | 100.00 cm | 6 MeV | 15x15 | Crossline | 2.70 cm | 2.70 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.7% | 100% | 15.20 cm | 1.33 cm - 1.33 cm | 0.00 cm
| TR2 | 2017-09-19 | 11:14 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.3% | 100% | 16.76 cm | 1.72 cm - 1.72 cm | 0.00 cm
| TR2 | 2017-09-19 | 11:16 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 1.80 cm | 1.80 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.4% | 100% | 16.88 cm | 2.05 cm - 2.05 cm | 0.00 cm
| TR2 | 2017-09-19 | 11:17 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 2.20 cm | 2.20 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.5% | 100% | 16.84 cm | 2.07 cm - 2.07 cm | 0.00 cm
| TR2 | 2017-09-19 | 11:18 | Electrons | 110.00 cm | 6 MeV | 15x15 | Crossline | 2.70 cm | 2.70 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.1% | 100% | 16.78 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:20 | Electrons | 110.00 cm | 9 MeV | 25x25 | Crossline | 0.80 cm | 0.80 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 28.02 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:21 | Electrons | 110.00 cm | 9 MeV | 25x25 | Crossline | 1.00 cm | 1.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 28.10 cm | 1.28 cm - 1.28 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:23 | Electrons | 110.00 cm | 9 MeV | 25x25 | Crossline | 1.40 cm | 1.40 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 28.22 cm | 1.39 cm - 1.39 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:24 | Electrons | 110.00 cm | 9 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 28.40 cm | 1.62 cm - 1.62 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:25 | Electrons | 110.00 cm | 9 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 28.48 cm | 1.82 cm - 1.82 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:27 | Electrons | 110.00 cm | 9 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 28.50 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:28 | Electrons | 110.00 cm | 9 MeV | 25x25 | Crossline | 3.50 cm | 3.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 28.45 cm | 1.96 cm - 1.96 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:30 | Electrons | 110.00 cm | 9 MeV | 25x25 | Crossline | 4.00 cm | 4.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.1% | 100% | 28.43 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:10 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 0.80 cm | 0.80 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 25.26 cm | 0.53 cm - 0.53 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:11 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 1.00 cm | 1.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 25.33 cm | 0.57 cm - 0.57 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:13 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 1.40 cm | 1.40 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.45 cm | 0.70 cm - 0.70 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:14 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.63 cm | 1.00 cm - 1.00 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:15 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.74 cm | 1.28 cm - 1.28 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:17 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.80 cm | 1.51 cm - 1.51 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:18 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 3.50 cm | 3.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 25.79 cm | 1.60 cm - 1.60 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:20 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 4.00 cm | 4.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.3% | 100% | 25.78 cm | 1.54 cm - 1.54 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:53 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 0.80 cm | 0.80 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 21.08 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:54 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 1.00 cm | 1.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 21.13 cm | 0.58 cm - 0.58 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:55 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 1.40 cm | 1.40 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 21.23 cm | 0.70 cm - 0.70 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:57 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 21.39 cm | 0.98 cm - 0.98 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:58 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 21.49 cm | 1.26 cm - 1.26 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:59 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 21.54 cm | 1.48 cm - 1.48 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:00 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 3.50 cm | 3.50 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 21.54 cm | 1.58 cm - 1.58 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:01 | Electrons | 100.00 cm | 9 MeV | 25x25 | Crossline | 4.00 cm | 4.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2% | 100% | 21.54 cm | 1.53 cm - 1.53 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:58 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 0.80 cm | 0.80 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 22.45 cm | 1.21 cm - 1.21 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:59 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 1.00 cm | 1.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 22.50 cm | 1.26 cm - 1.26 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:00 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 1.40 cm | 1.40 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 22.60 cm | 1.36 cm - 1.36 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:02 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 22.75 cm | 1.60 cm - 1.60 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:03 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 22.83 cm | 1.81 cm - 1.81 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:04 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 22.83 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:05 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 3.50 cm | 3.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 22.81 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:07 | Electrons | 110.00 cm | 9 MeV | 20x20 | Crossline | 4.00 cm | 4.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 22.77 cm | 1.86 cm - 1.86 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:58 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 0.80 cm | 0.80 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.29 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:59 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 1.00 cm | 1.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.34 cm | 0.59 cm - 0.59 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:00 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 1.40 cm | 1.40 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 20.44 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:01 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.58 cm | 1.05 cm - 1.05 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:02 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.66 cm | 1.32 cm - 1.32 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:04 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 20.70 cm | 1.52 cm - 1.52 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:05 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 3.50 cm | 3.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 20.67 cm | 1.58 cm - 1.58 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:06 | Electrons | 100.00 cm | 9 MeV | 20x20 | Crossline | 4.00 cm | 4.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 20.67 cm | 1.51 cm - 1.51 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:15 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 16.84 cm | 1.22 cm - 1.22 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:17 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 1.00 cm | 1.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 16.87 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:18 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 1.40 cm | 1.40 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 16.96 cm | 1.36 cm - 1.36 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:19 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 17.06 cm | 1.60 cm - 1.60 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:20 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.5% | 100% | 17.12 cm | 1.82 cm - 1.82 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:21 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4% | 100% | 17.12 cm | 1.96 cm - 1.96 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:22 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 3.50 cm | 3.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.2% | 100% | 17.10 cm | 1.99 cm - 1.99 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:24 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.6% | 100% | 17.05 cm | 1.91 cm - 1.91 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:02 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 15.19 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:03 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 1.00 cm | 1.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 15.23 cm | 0.59 cm - 0.59 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:04 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 1.40 cm | 1.40 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 15.31 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:05 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 15.41 cm | 1.05 cm - 1.05 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:06 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.2% | 100% | 15.47 cm | 1.34 cm - 1.34 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:08 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.8% | 100% | 15.49 cm | 1.55 cm - 1.55 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:09 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 3.50 cm | 3.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.1% | 100% | 15.48 cm | 1.63 cm - 1.63 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:10 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 15.46 cm | 1.57 cm - 1.57 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:17 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 0.80 cm | 0.80 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 11.08 cm | 0.56 cm - 0.56 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:18 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 1.00 cm | 1.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 11.11 cm | 0.60 cm - 0.60 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:19 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 1.40 cm | 1.40 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 11.15 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:20 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.4% | 100% | 11.23 cm | 1.03 cm - 1.03 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:21 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.2% | 100% | 11.27 cm | 1.32 cm - 1.32 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:22 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.7% | 100% | 11.29 cm | 1.53 cm - 1.53 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:23 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 3.50 cm | 3.50 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.4% | 100% | 11.29 cm | 1.61 cm - 1.61 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:24 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.8% | 100% | 11.29 cm | 1.56 cm - 1.56 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:47 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 0.00 cm | 0.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.1% | 100% | 10.06 cm | 0.44 cm - 0.44 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:49 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 10.15 cm | 0.52 cm - 0.52 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:50 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 10.17 cm | 0.56 cm - 0.56 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:51 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 1.40 cm | 1.40 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 10.22 cm | 0.69 cm - 0.69 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:52 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.7% | 100% | 10.29 cm | 0.99 cm - 0.99 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:53 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.9% | 100% | 10.33 cm | 1.28 cm - 1.28 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:54 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.7% | 100% | 10.35 cm | 1.51 cm - 1.51 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:55 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.5% | 100% | 10.34 cm | 1.61 cm - 1.61 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:56 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7% | 100% | 10.34 cm | 1.56 cm - 1.56 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:33 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3% | 100% | 11.24 cm | 1.19 cm - 1.19 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:35 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 11.27 cm | 1.23 cm - 1.23 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:36 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 1.40 cm | 1.40 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4% | 100% | 11.32 cm | 1.33 cm - 1.33 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:37 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.8% | 100% | 11.39 cm | 1.57 cm - 1.57 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:38 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.4% | 100% | 11.43 cm | 1.79 cm - 1.79 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:39 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.5% | 100% | 11.43 cm | 1.94 cm - 1.94 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:40 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.9% | 100% | 11.42 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:41 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.5% | 100% | 11.40 cm | 1.90 cm - 1.90 cm | 0.00 cm
| TR2 | 2017-08-25 | 16:03 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 7.00 cm | 0.57 cm - 0.57 cm | 0.00 cm
| TR2 | 2017-08-25 | 16:03 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 7.02 cm | 0.60 cm - 0.60 cm | 0.00 cm
| TR2 | 2017-08-25 | 16:04 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 1.40 cm | 1.40 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 7.04 cm | 0.72 cm - 0.72 cm | 0.00 cm
| TR2 | 2017-08-25 | 16:05 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.5% | 100% | 7.09 cm | 1.00 cm - 1.00 cm | 0.00 cm
| TR2 | 2017-08-25 | 16:06 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.7% | 100% | 7.11 cm | 1.29 cm - 1.29 cm | 0.00 cm
| TR2 | 2017-08-25 | 16:07 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.1% | 100% | 7.12 cm | 1.53 cm - 1.53 cm | 0.00 cm
| TR2 | 2017-08-25 | 16:07 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.3% | 100% | 7.12 cm | 1.66 cm - 1.66 cm | 0.00 cm
| TR2 | 2017-08-25 | 16:08 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13% | 100% | 7.12 cm | 1.63 cm - 1.63 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:02 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11% | 0% | 2.97 cm | 0.59 cm - 0.59 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:03 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.8% | 0% | 2.98 cm | 0.63 cm - 0.63 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:03 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 1.40 cm | 1.40 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.2% | 0% | 2.99 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:04 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.4% | 0% | 3.06 cm | 1.01 cm - 1.01 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:05 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.3% | 0% | 3.16 cm | 1.23 cm - 1.23 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:05 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.4% | 0% | 3.27 cm | 1.40 cm - 1.40 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:06 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.8% | 0% | 3.31 cm | 1.46 cm - 1.46 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:07 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.7% | 0% | 3.28 cm | 1.45 cm - 1.45 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:15 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.4% | 100% | 6.09 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:16 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3% | 100% | 6.10 cm | 0.59 cm - 0.59 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:16 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 1.40 cm | 1.40 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.7% | 100% | 6.13 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:17 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.8% | 100% | 6.16 cm | 1.05 cm - 1.05 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:18 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.1% | 100% | 6.18 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:18 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.4% | 100% | 6.17 cm | 1.58 cm - 1.58 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:19 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.3% | 100% | 6.17 cm | 1.67 cm - 1.67 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:20 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15% | 100% | 6.18 cm | 1.62 cm - 1.62 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:43 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.3% | 100% | 6.76 cm | 1.21 cm - 1.21 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:44 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.7% | 100% | 6.78 cm | 1.24 cm - 1.24 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:45 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 1.40 cm | 1.40 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.7% | 100% | 6.80 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:46 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.9% | 100% | 6.84 cm | 1.58 cm - 1.58 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:47 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.7% | 100% | 6.86 cm | 1.81 cm - 1.81 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:48 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16% | 100% | 6.86 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:49 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.4% | 100% | 6.84 cm | 2.02 cm - 2.02 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:50 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.8% | 100% | 6.84 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:04 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.1% | 100% | 4.03 cm | 0.56 cm - 0.56 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:05 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.2% | 100% | 4.03 cm | 0.60 cm - 0.60 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:05 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 1.40 cm | 1.40 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.6% | 100% | 4.05 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:06 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.9% | 100% | 4.07 cm | 1.02 cm - 1.02 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:07 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.9% | 0% | 4.10 cm | 1.28 cm - 1.28 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:07 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.8% | 0% | 4.09 cm | 1.49 cm - 1.49 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:08 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.7% | 0% | 4.09 cm | 1.59 cm - 1.59 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:09 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.6% | 0% | 4.09 cm | 1.58 cm - 1.58 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:51 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 0.80 cm | 0.80 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.7% | 100% | 3.11 cm | 0.56 cm - 0.56 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:52 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.6% | 0% | 3.12 cm | 0.60 cm - 0.60 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:52 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 1.40 cm | 1.40 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.9% | 0% | 3.14 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:53 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.5% | 0% | 3.19 cm | 0.98 cm - 0.98 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:54 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.8% | 0% | 3.28 cm | 1.22 cm - 1.22 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:54 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20% | 0% | 3.38 cm | 1.39 cm - 1.39 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:55 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.7% | 0% | 3.43 cm | 1.47 cm - 1.47 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:56 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.6% | 0% | 3.40 cm | 1.47 cm - 1.47 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:22 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.3% | 100% | 4.01 cm | 0.52 cm - 0.52 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:24 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.4% | 0% | 4.20 cm | 1.59 cm - 1.59 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:01 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.1% | 100% | 4.44 cm | 1.17 cm - 1.17 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:03 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.6% | 100% | 4.46 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:03 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 1.50 cm | 1.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.7% | 100% | 4.49 cm | 1.36 cm - 1.36 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:04 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.9% | 0% | 4.52 cm | 1.53 cm - 1.53 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:05 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.1% | 0% | 4.56 cm | 1.70 cm - 1.70 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:05 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.9% | 0% | 4.58 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:06 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.3% | 0% | 4.58 cm | 1.88 cm - 1.88 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:07 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.9% | 0% | 4.58 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:07 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.2% | 0% | 4.70 cm | 1.89 cm - 1.89 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:41 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 6.08 cm | 0.51 cm - 0.51 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:42 | Electrons | 100.00 cm | 9 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.9% | 100% | 6.30 cm | 1.61 cm - 1.61 cm | 0.00 cm
| TR2 | 2017-09-19 | 13:57 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.9% | 100% | 6.73 cm | 1.17 cm - 1.17 cm | 0.00 cm
| TR2 | 2017-09-19 | 13:58 | Electrons | 110.00 cm | 9 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.7% | 100% | 6.98 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR2 | 2017-09-19 | 12:07 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.8% | 100% | 11.22 cm | 1.16 cm - 1.16 cm | 0.00 cm
| TR2 | 2017-09-19 | 12:09 | Electrons | 110.00 cm | 9 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8% | 100% | 11.48 cm | 1.98 cm - 1.98 cm | 0.00 cm
| TR2 | 2017-09-18 | 12:13 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 15.15 cm | 0.50 cm - 0.50 cm | 0.00 cm
| TR2 | 2017-09-18 | 12:15 | Electrons | 100.00 cm | 9 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.3% | 100% | 15.53 cm | 1.56 cm - 1.56 cm | 0.00 cm
| TR2 | 2017-09-19 | 11:19 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 16.79 cm | 1.17 cm - 1.17 cm | 0.00 cm
| TR2 | 2017-09-19 | 11:20 | Electrons | 110.00 cm | 9 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.8% | 100% | 17.08 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:03 | Electrons | 110.00 cm | 12 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 27.93 cm | 0.92 cm - 0.92 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:05 | Electrons | 110.00 cm | 12 MeV | 25x25 | Crossline | 1.00 cm | 1.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 28.08 cm | 0.97 cm - 0.97 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:06 | Electrons | 110.00 cm | 12 MeV | 25x25 | Crossline | 1.40 cm | 1.40 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 28.18 cm | 1.04 cm - 1.04 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:08 | Electrons | 110.00 cm | 12 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 28.37 cm | 1.19 cm - 1.19 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:09 | Electrons | 110.00 cm | 12 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 28.54 cm | 1.39 cm - 1.39 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:10 | Electrons | 110.00 cm | 12 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 28.69 cm | 1.61 cm - 1.61 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:12 | Electrons | 110.00 cm | 12 MeV | 25x25 | Crossline | 4.00 cm | 4.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 28.85 cm | 2.04 cm - 2.04 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:13 | Electrons | 110.00 cm | 12 MeV | 25x25 | Crossline | 4.50 cm | 4.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 28.88 cm | 2.14 cm - 2.14 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:15 | Electrons | 110.00 cm | 12 MeV | 25x25 | Crossline | 5.00 cm | 5.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 28.88 cm | 2.13 cm - 2.13 cm | 0.00 cm
| TR2 | 2017-08-29 | 11:16 | Electrons | 110.00 cm | 12 MeV | 25x25 | Crossline | 5.50 cm | 5.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.8% | 100% | 28.86 cm | 2.05 cm - 2.05 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:53 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 25.18 cm | 0.40 cm - 0.40 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:55 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 1.00 cm | 1.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 25.32 cm | 0.46 cm - 0.46 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:56 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 1.40 cm | 1.40 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.43 cm | 0.53 cm - 0.53 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:57 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.62 cm | 0.72 cm - 0.72 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:58 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.77 cm | 0.93 cm - 0.93 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:00 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.93 cm | 1.20 cm - 1.20 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:01 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 4.00 cm | 4.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 26.14 cm | 1.74 cm - 1.74 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:02 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 4.50 cm | 4.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 26.20 cm | 1.90 cm - 1.90 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:04 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 5.00 cm | 5.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 26.19 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:05 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 5.50 cm | 5.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.7% | 100% | 26.23 cm | 1.89 cm - 1.89 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:36 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 21.02 cm | 0.42 cm - 0.42 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:38 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 1.00 cm | 1.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 21.13 cm | 0.47 cm - 0.47 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:39 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 1.40 cm | 1.40 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 21.23 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:40 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 21.37 cm | 0.72 cm - 0.72 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:41 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 21.50 cm | 0.93 cm - 0.93 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:42 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 21.63 cm | 1.19 cm - 1.19 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:43 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 4.00 cm | 4.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 21.84 cm | 1.72 cm - 1.72 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:45 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 4.50 cm | 4.50 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 21.87 cm | 1.90 cm - 1.90 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:46 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 5.00 cm | 5.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 21.89 cm | 1.96 cm - 1.96 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:47 | Electrons | 100.00 cm | 12 MeV | 25x25 | Crossline | 5.50 cm | 5.50 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 21.91 cm | 1.92 cm - 1.92 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:43 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 22.38 cm | 0.91 cm - 0.91 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:44 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 1.00 cm | 1.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 22.49 cm | 0.95 cm - 0.95 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:45 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 1.40 cm | 1.40 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 22.58 cm | 1.03 cm - 1.03 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:47 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 22.73 cm | 1.18 cm - 1.18 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:48 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 22.87 cm | 1.38 cm - 1.38 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:49 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 22.99 cm | 1.61 cm - 1.61 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:50 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 4.00 cm | 4.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 23.14 cm | 2.04 cm - 2.04 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:52 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 4.50 cm | 4.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 23.16 cm | 2.15 cm - 2.15 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:53 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 5.00 cm | 5.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 23.14 cm | 2.16 cm - 2.16 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:54 | Electrons | 110.00 cm | 12 MeV | 20x20 | Crossline | 5.50 cm | 5.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 23.13 cm | 2.10 cm - 2.10 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:44 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.22 cm | 0.41 cm - 0.41 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:45 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 1.00 cm | 1.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.34 cm | 0.47 cm - 0.47 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:46 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 1.40 cm | 1.40 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 20.42 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:47 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 20.58 cm | 0.74 cm - 0.74 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:48 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.70 cm | 0.97 cm - 0.97 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:50 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.83 cm | 1.24 cm - 1.24 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:51 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 4.00 cm | 4.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 20.98 cm | 1.76 cm - 1.76 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:52 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 4.50 cm | 4.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 21.02 cm | 1.91 cm - 1.91 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:53 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 5.00 cm | 5.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 21.01 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:54 | Electrons | 100.00 cm | 12 MeV | 20x20 | Crossline | 5.50 cm | 5.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 21.03 cm | 1.88 cm - 1.88 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:00 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 16.79 cm | 0.91 cm - 0.91 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:01 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 1.00 cm | 1.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 16.87 cm | 0.95 cm - 0.95 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:03 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 1.40 cm | 1.40 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 16.95 cm | 1.03 cm - 1.03 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:04 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 17.05 cm | 1.19 cm - 1.19 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:06 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 17.15 cm | 1.38 cm - 1.38 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:07 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 17.24 cm | 1.62 cm - 1.62 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:08 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.4% | 100% | 17.34 cm | 2.08 cm - 2.08 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:09 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.1% | 100% | 17.34 cm | 2.20 cm - 2.20 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:10 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 5.00 cm | 5.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.1% | 100% | 17.34 cm | 2.23 cm - 2.23 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:12 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.6% | 100% | 17.31 cm | 2.17 cm - 2.17 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:49 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.2% | 100% | 15.14 cm | 0.42 cm - 0.42 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:50 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 1.00 cm | 1.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 15.23 cm | 0.48 cm - 0.48 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:51 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 1.40 cm | 1.40 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 15.30 cm | 0.56 cm - 0.56 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:52 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 15.41 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:53 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 15.50 cm | 0.99 cm - 0.99 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:54 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 15.59 cm | 1.27 cm - 1.27 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:55 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.9% | 100% | 15.70 cm | 1.82 cm - 1.82 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:56 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.9% | 100% | 15.72 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:58 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 5.00 cm | 5.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5% | 100% | 15.71 cm | 2.02 cm - 2.02 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:59 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.6% | 100% | 15.72 cm | 1.98 cm - 1.98 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:04 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.2% | 100% | 11.05 cm | 0.43 cm - 0.43 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:05 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 1.00 cm | 1.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.2% | 100% | 11.11 cm | 0.49 cm - 0.49 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:06 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 1.40 cm | 1.40 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 11.16 cm | 0.57 cm - 0.57 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:07 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 11.24 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:08 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2% | 100% | 11.30 cm | 0.99 cm - 0.99 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:09 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.7% | 100% | 11.36 cm | 1.27 cm - 1.27 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:10 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.5% | 100% | 11.44 cm | 1.81 cm - 1.81 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:11 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.8% | 100% | 11.46 cm | 1.98 cm - 1.98 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:12 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 5.00 cm | 5.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.1% | 100% | 11.45 cm | 2.03 cm - 2.03 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:13 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.9% | 100% | 11.48 cm | 1.99 cm - 1.99 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:22 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 0.00 cm | 0.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 10.07 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:23 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 10.12 cm | 0.40 cm - 0.40 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:23 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 10.17 cm | 0.45 cm - 0.45 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:24 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 1.50 cm | 1.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 10.23 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:25 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 10.29 cm | 0.72 cm - 0.72 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:26 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 10.34 cm | 0.95 cm - 0.95 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:27 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.6% | 100% | 10.40 cm | 1.23 cm - 1.23 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:28 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.8% | 100% | 10.47 cm | 1.79 cm - 1.79 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:29 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10% | 100% | 10.48 cm | 1.96 cm - 1.96 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:30 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.5% | 100% | 10.48 cm | 2.03 cm - 2.03 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:31 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.3% | 100% | 10.50 cm | 2.00 cm - 2.00 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:32 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 6.00 cm | 6.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.8% | 100% | 10.56 cm | 1.99 cm - 1.99 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:15 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 11.22 cm | 0.88 cm - 0.88 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:16 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 11.28 cm | 0.93 cm - 0.93 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:17 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 1.40 cm | 1.40 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.2% | 100% | 11.32 cm | 1.00 cm - 1.00 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:18 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.1% | 100% | 11.40 cm | 1.16 cm - 1.16 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:19 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.4% | 100% | 11.45 cm | 1.36 cm - 1.36 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:21 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.1% | 100% | 11.51 cm | 1.60 cm - 1.60 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:22 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.4% | 100% | 11.57 cm | 2.08 cm - 2.08 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:23 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.2% | 100% | 11.57 cm | 2.21 cm - 2.21 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:24 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.5% | 100% | 11.56 cm | 2.24 cm - 2.24 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:25 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.2% | 100% | 11.54 cm | 2.20 cm - 2.20 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:47 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 6.99 cm | 0.43 cm - 0.43 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:47 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 7.03 cm | 0.48 cm - 0.48 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:48 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 1.40 cm | 1.40 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 7.06 cm | 0.56 cm - 0.56 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:49 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 7.11 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:50 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.6% | 100% | 7.15 cm | 0.94 cm - 0.94 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:51 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.5% | 100% | 7.19 cm | 1.21 cm - 1.21 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:51 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.8% | 100% | 7.23 cm | 1.78 cm - 1.78 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:52 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.4% | 100% | 7.24 cm | 1.98 cm - 1.98 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:53 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.9% | 100% | 7.23 cm | 2.07 cm - 2.07 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:54 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.7% | 100% | 7.24 cm | 2.04 cm - 2.04 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:53 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.9% | 100% | 2.97 cm | 0.44 cm - 0.44 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:54 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.6% | 100% | 2.99 cm | 0.50 cm - 0.50 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:55 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 1.40 cm | 1.40 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.4% | 0% | 3.00 cm | 0.58 cm - 0.58 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:55 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.9% | 0% | 3.03 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:56 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.5% | 0% | 3.07 cm | 0.96 cm - 0.96 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:56 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.7% | 0% | 3.16 cm | 1.18 cm - 1.18 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:57 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.8% | 0% | 3.47 cm | 1.57 cm - 1.57 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:58 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 21% | 0% | 3.62 cm | 1.68 cm - 1.68 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:58 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 21.2% | 0% | 3.67 cm | 1.74 cm - 1.74 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:59 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.9% | 0% | 3.66 cm | 1.76 cm - 1.76 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:06 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 6.07 cm | 0.41 cm - 0.41 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:07 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 6.11 cm | 0.47 cm - 0.47 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:08 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 1.40 cm | 1.40 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 6.13 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:08 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.9% | 100% | 6.17 cm | 0.75 cm - 0.75 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:09 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.8% | 100% | 6.21 cm | 0.98 cm - 0.98 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:10 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.1% | 100% | 6.24 cm | 1.27 cm - 1.27 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:10 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16% | 100% | 6.27 cm | 1.83 cm - 1.83 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:11 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.2% | 100% | 6.28 cm | 2.00 cm - 2.00 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:12 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.7% | 100% | 6.26 cm | 2.06 cm - 2.06 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:13 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.6% | 100% | 6.28 cm | 2.03 cm - 2.03 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:31 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.5% | 100% | 6.75 cm | 0.88 cm - 0.88 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:32 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6% | 100% | 6.78 cm | 0.93 cm - 0.93 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:33 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 1.40 cm | 1.40 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.9% | 100% | 6.81 cm | 1.00 cm - 1.00 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:34 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.5% | 100% | 6.86 cm | 1.15 cm - 1.15 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:35 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.6% | 100% | 6.90 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:35 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.8% | 100% | 6.92 cm | 1.59 cm - 1.59 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:36 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.3% | 100% | 6.96 cm | 2.07 cm - 2.07 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:37 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.3% | 100% | 6.95 cm | 2.21 cm - 2.21 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:38 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.5% | 100% | 6.94 cm | 2.25 cm - 2.25 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:39 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.3% | 100% | 6.95 cm | 2.20 cm - 2.20 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:55 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 4.02 cm | 0.42 cm - 0.42 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:55 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4% | 100% | 4.04 cm | 0.47 cm - 0.47 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:56 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 1.40 cm | 1.40 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.8% | 100% | 4.06 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:57 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.4% | 100% | 4.09 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:57 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.8% | 100% | 4.12 cm | 0.95 cm - 0.95 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:58 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.7% | 100% | 4.15 cm | 1.20 cm - 1.20 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:58 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.3% | 0% | 4.27 cm | 1.66 cm - 1.66 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:59 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.3% | 0% | 4.32 cm | 1.81 cm - 1.81 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:00 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.5% | 0% | 4.33 cm | 1.87 cm - 1.87 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:00 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.5% | 0% | 4.30 cm | 1.88 cm - 1.88 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:42 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.5% | 100% | 3.11 cm | 0.42 cm - 0.42 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:42 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7% | 100% | 3.12 cm | 0.48 cm - 0.48 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:43 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 1.40 cm | 1.40 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.9% | 100% | 3.14 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:43 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.5% | 0% | 3.17 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:44 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.7% | 0% | 3.22 cm | 0.93 cm - 0.93 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:45 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.9% | 0% | 3.28 cm | 1.14 cm - 1.14 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:45 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.4% | 0% | 3.57 cm | 1.55 cm - 1.55 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:46 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.9% | 0% | 3.71 cm | 1.68 cm - 1.68 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:47 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.8% | 0% | 3.78 cm | 1.73 cm - 1.73 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:47 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.9% | 0% | 3.74 cm | 1.74 cm - 1.74 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:25 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.1% | 0% | 4.21 cm | 1.48 cm - 1.48 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:08 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.9% | 100% | 4.46 cm | 0.88 cm - 0.88 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:10 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.5% | 100% | 4.49 cm | 0.92 cm - 0.92 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:11 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 1.50 cm | 1.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.6% | 100% | 4.51 cm | 1.01 cm - 1.01 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:11 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14% | 100% | 4.54 cm | 1.13 cm - 1.13 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:12 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.7% | 100% | 4.56 cm | 1.30 cm - 1.30 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:12 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.2% | 100% | 4.60 cm | 1.49 cm - 1.49 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:13 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.7% | 100% | 4.66 cm | 1.70 cm - 1.70 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:14 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.7% | 0% | 4.72 cm | 1.88 cm - 1.88 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:14 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.3% | 0% | 4.75 cm | 2.00 cm - 2.00 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:15 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.4% | 0% | 4.76 cm | 2.03 cm - 2.03 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:16 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.3% | 0% | 4.75 cm | 2.04 cm - 2.04 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:43 | Electrons | 100.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.1% | 100% | 6.28 cm | 1.60 cm - 1.60 cm | 0.00 cm
| TR2 | 2017-09-19 | 13:59 | Electrons | 110.00 cm | 12 MeV | 6x6 | Crossline | 3.50 cm | 3.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.9% | 100% | 6.95 cm | 1.87 cm - 1.87 cm | 0.00 cm
| TR2 | 2017-09-19 | 12:10 | Electrons | 110.00 cm | 12 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.1% | 100% | 11.55 cm | 1.89 cm - 1.89 cm | 0.00 cm
| TR2 | 2017-09-18 | 12:16 | Electrons | 100.00 cm | 12 MeV | 15x15 | Crossline | 3.50 cm | 3.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.1% | 100% | 15.66 cm | 1.61 cm - 1.61 cm | 0.00 cm
| TR2 | 2017-09-19 | 11:21 | Electrons | 110.00 cm | 12 MeV | 15x15 | Crossline | 3.50 cm | 3.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.6% | 100% | 17.30 cm | 1.90 cm - 1.90 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:31 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4% | 100% | 3.12 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:31 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.2% | 100% | 3.13 cm | 0.43 cm - 0.43 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:32 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.6% | 100% | 3.18 cm | 0.59 cm - 0.59 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:33 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.5% | 0% | 3.20 cm | 0.74 cm - 0.74 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:33 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.1% | 0% | 3.23 cm | 0.91 cm - 0.91 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:34 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.8% | 0% | 3.38 cm | 1.29 cm - 1.29 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:34 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.8% | 0% | 3.52 cm | 1.49 cm - 1.49 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:35 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.5% | 0% | 3.69 cm | 1.69 cm - 1.69 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:36 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.9% | 0% | 3.89 cm | 1.86 cm - 1.86 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:36 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 6.00 cm | 6.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 21.3% | 0% | 4.08 cm | 2.00 cm - 2.00 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:37 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 21.4% | 0% | 4.21 cm | 2.08 cm - 2.08 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:38 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 7.00 cm | 7.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 21.4% | 0% | 4.25 cm | 2.13 cm - 2.13 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:43 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 4.03 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:43 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 4.05 cm | 0.42 cm - 0.42 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:44 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.4% | 100% | 4.10 cm | 0.60 cm - 0.60 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:45 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.3% | 100% | 4.12 cm | 0.74 cm - 0.74 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:45 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.1% | 100% | 4.15 cm | 0.93 cm - 0.93 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:46 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.9% | 100% | 4.22 cm | 1.36 cm - 1.36 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:46 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.4% | 0% | 4.29 cm | 1.58 cm - 1.58 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:47 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.5% | 0% | 4.39 cm | 1.79 cm - 1.79 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:48 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.2% | 0% | 4.52 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:49 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.6% | 0% | 4.70 cm | 2.19 cm - 2.19 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:14 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.1% | 100% | 6.73 cm | 0.71 cm - 0.71 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:15 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.6% | 100% | 6.77 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:16 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.4% | 100% | 6.84 cm | 0.91 cm - 0.91 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:17 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.1% | 100% | 6.88 cm | 1.04 cm - 1.04 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:18 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.9% | 100% | 6.92 cm | 1.21 cm - 1.21 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:19 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.9% | 100% | 7.00 cm | 1.65 cm - 1.65 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:20 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.8% | 100% | 7.04 cm | 1.91 cm - 1.91 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:21 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.5% | 100% | 7.07 cm | 2.15 cm - 2.15 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:22 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.5% | 100% | 7.09 cm | 2.35 cm - 2.35 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:23 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 6.00 cm | 6.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.4% | 100% | 7.10 cm | 2.50 cm - 2.50 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:24 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.7% | 100% | 7.10 cm | 2.58 cm - 2.58 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:25 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 7.00 cm | 7.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.8% | 100% | 7.10 cm | 2.60 cm - 2.60 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:56 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 6.08 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:56 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 6.11 cm | 0.41 cm - 0.41 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:57 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.4% | 100% | 6.18 cm | 0.60 cm - 0.60 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:58 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.4% | 100% | 6.22 cm | 0.75 cm - 0.75 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:58 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.9% | 100% | 6.26 cm | 0.94 cm - 0.94 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:59 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.3% | 100% | 6.33 cm | 1.43 cm - 1.43 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:00 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 4.50 cm | 4.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.6% | 100% | 6.37 cm | 1.71 cm - 1.71 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:00 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.4% | 100% | 6.40 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:01 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 5.50 cm | 5.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.7% | 100% | 6.44 cm | 2.18 cm - 2.18 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:02 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 6.50 cm | 6.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.2% | 100% | 6.47 cm | 2.44 cm - 2.44 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:44 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.2% | 100% | 2.98 cm | 0.41 cm - 0.41 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:44 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.5% | 100% | 3.00 cm | 0.44 cm - 0.44 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:45 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.9% | 0% | 3.03 cm | 0.62 cm - 0.62 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:45 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.7% | 0% | 3.06 cm | 0.77 cm - 0.77 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:46 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.1% | 0% | 3.09 cm | 0.94 cm - 0.94 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:47 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.4% | 0% | 3.26 cm | 1.32 cm - 1.32 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:47 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.2% | 0% | 3.40 cm | 1.51 cm - 1.51 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:48 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.8% | 0% | 3.60 cm | 1.71 cm - 1.71 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:48 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 21.2% | 0% | 3.81 cm | 1.89 cm - 1.89 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:49 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 6.00 cm | 6.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 21.5% | 0% | 4.03 cm | 2.01 cm - 2.01 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:50 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 6.50 cm | 6.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 21.8% | 0% | 4.17 cm | 2.11 cm - 2.11 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:50 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 7.00 cm | 7.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 21.8% | 0% | 4.21 cm | 2.18 cm - 2.18 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:26 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 7.00 cm | 0.40 cm - 0.40 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:27 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 7.04 cm | 0.44 cm - 0.44 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:27 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 7.11 cm | 0.60 cm - 0.60 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:28 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.1% | 100% | 7.15 cm | 0.74 cm - 0.74 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:29 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.2% | 100% | 7.19 cm | 0.93 cm - 0.93 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:30 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.2% | 100% | 7.28 cm | 1.40 cm - 1.40 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:30 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.6% | 100% | 7.32 cm | 1.68 cm - 1.68 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:31 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.4% | 100% | 7.35 cm | 1.94 cm - 1.94 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:32 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.1% | 100% | 7.38 cm | 2.18 cm - 2.18 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:33 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 6.00 cm | 6.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.1% | 100% | 7.40 cm | 2.37 cm - 2.37 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:34 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 6.50 cm | 6.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.8% | 100% | 7.41 cm | 2.47 cm - 2.47 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:34 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 7.00 cm | 7.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.9% | 100% | 7.43 cm | 2.52 cm - 2.52 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:57 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 11.20 cm | 0.72 cm - 0.72 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:58 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 11.26 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:59 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 11.37 cm | 0.92 cm - 0.92 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:00 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.3% | 100% | 11.43 cm | 1.05 cm - 1.05 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:01 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 11.48 cm | 1.23 cm - 1.23 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:02 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.4% | 100% | 11.61 cm | 1.69 cm - 1.69 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:04 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.1% | 100% | 11.66 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:05 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.7% | 100% | 11.69 cm | 2.20 cm - 2.20 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:06 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.1% | 100% | 11.72 cm | 2.40 cm - 2.40 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:07 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 6.00 cm | 6.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.1% | 100% | 11.72 cm | 2.56 cm - 2.56 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:08 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 6.50 cm | 6.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.4% | 100% | 11.72 cm | 2.63 cm - 2.63 cm | 0.00 cm
| TR2 | 2017-08-29 | 15:09 | Electrons | 110.00 cm | 16 MeV | 10x10 | Crossline | 7.00 cm | 7.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.6% | 100% | 11.71 cm | 2.66 cm - 2.66 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:56 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 0.00 cm | 0.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.2% | 100% | 10.07 cm | 0.35 cm - 0.35 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:57 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 10.12 cm | 0.37 cm - 0.37 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:58 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 10.18 cm | 0.40 cm - 0.40 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:59 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 1.50 cm | 1.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 10.23 cm | 0.47 cm - 0.47 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:00 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 10.28 cm | 0.57 cm - 0.57 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:01 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 10.34 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:02 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.2% | 100% | 10.40 cm | 0.91 cm - 0.91 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:02 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.7% | 100% | 10.46 cm | 1.15 cm - 1.15 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:03 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.6% | 100% | 10.51 cm | 1.41 cm - 1.41 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:04 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.7% | 100% | 10.57 cm | 1.70 cm - 1.70 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:05 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.7% | 100% | 10.61 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:06 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.4% | 100% | 10.65 cm | 2.21 cm - 2.21 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:08 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 6.50 cm | 6.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.1% | 100% | 10.67 cm | 2.49 cm - 2.49 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:49 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 11.05 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:50 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 1.00 cm | 1.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 11.11 cm | 0.44 cm - 0.44 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:51 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 11.22 cm | 0.61 cm - 0.61 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:52 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 11.29 cm | 0.77 cm - 0.77 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:53 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2% | 100% | 11.35 cm | 0.96 cm - 0.96 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:54 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5% | 100% | 11.48 cm | 1.46 cm - 1.46 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:55 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.9% | 100% | 11.53 cm | 1.74 cm - 1.74 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:56 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 5.00 cm | 5.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.8% | 100% | 11.59 cm | 2.02 cm - 2.02 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:57 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.3% | 100% | 11.62 cm | 2.25 cm - 2.25 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:58 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 6.00 cm | 6.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.3% | 100% | 11.64 cm | 2.42 cm - 2.42 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:59 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 6.50 cm | 6.50 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.9% | 100% | 11.63 cm | 2.51 cm - 2.51 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:00 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 7.00 cm | 7.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.2% | 100% | 11.65 cm | 2.54 cm - 2.54 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:42 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 16.77 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:43 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 1.00 cm | 1.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 16.84 cm | 0.78 cm - 0.78 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:44 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 17.01 cm | 0.94 cm - 0.94 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:45 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 17.09 cm | 1.08 cm - 1.08 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:46 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 17.20 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:47 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.9% | 100% | 17.36 cm | 1.71 cm - 1.71 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:49 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4% | 100% | 17.44 cm | 1.97 cm - 1.97 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:50 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 5.00 cm | 5.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.1% | 100% | 17.49 cm | 2.21 cm - 2.21 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:51 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.1% | 100% | 17.54 cm | 2.42 cm - 2.42 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:52 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 6.00 cm | 6.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.8% | 100% | 17.53 cm | 2.57 cm - 2.57 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:53 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 6.50 cm | 6.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.2% | 100% | 17.54 cm | 2.64 cm - 2.64 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:55 | Electrons | 110.00 cm | 16 MeV | 15x15 | Crossline | 7.00 cm | 7.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.3% | 100% | 17.51 cm | 2.66 cm - 2.66 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:33 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.2% | 100% | 15.14 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:35 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 1.00 cm | 1.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.1% | 100% | 15.22 cm | 0.42 cm - 0.42 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:36 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 15.38 cm | 0.60 cm - 0.60 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:37 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 15.47 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:38 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 15.56 cm | 0.96 cm - 0.96 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:39 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.4% | 100% | 15.74 cm | 1.47 cm - 1.47 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:40 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 4.50 cm | 4.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.6% | 100% | 15.82 cm | 1.74 cm - 1.74 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:41 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 5.00 cm | 5.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.9% | 100% | 15.89 cm | 2.02 cm - 2.02 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:42 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 5.50 cm | 5.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.2% | 100% | 15.92 cm | 2.26 cm - 2.26 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:44 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 6.50 cm | 6.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.6% | 100% | 15.93 cm | 2.52 cm - 2.52 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:27 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.22 cm | 0.37 cm - 0.37 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:28 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 1.00 cm | 1.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.32 cm | 0.41 cm - 0.41 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:30 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 20.55 cm | 0.60 cm - 0.60 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:31 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 20.65 cm | 0.75 cm - 0.75 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:32 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 20.78 cm | 0.94 cm - 0.94 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:33 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 4.00 cm | 4.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 21.01 cm | 1.44 cm - 1.44 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:34 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 4.50 cm | 4.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 21.13 cm | 1.71 cm - 1.71 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:35 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 5.00 cm | 5.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 21.20 cm | 1.98 cm - 1.98 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:36 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 5.50 cm | 5.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 21.27 cm | 2.21 cm - 2.21 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:37 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 6.00 cm | 6.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 21.28 cm | 2.37 cm - 2.37 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:39 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 6.50 cm | 6.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.7% | 100% | 21.30 cm | 2.44 cm - 2.44 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:40 | Electrons | 100.00 cm | 16 MeV | 20x20 | Crossline | 7.00 cm | 7.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.7% | 100% | 21.27 cm | 2.47 cm - 2.47 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:25 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 22.35 cm | 0.74 cm - 0.74 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:26 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 1.00 cm | 1.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 22.46 cm | 0.77 cm - 0.77 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:27 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 22.68 cm | 0.94 cm - 0.94 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:29 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 22.80 cm | 1.07 cm - 1.07 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:30 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 22.92 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:31 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 4.00 cm | 4.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 23.17 cm | 1.70 cm - 1.70 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:32 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 4.50 cm | 4.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 23.28 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:34 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 5.00 cm | 5.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2% | 100% | 23.35 cm | 2.19 cm - 2.19 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:35 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 5.50 cm | 5.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 23.40 cm | 2.40 cm - 2.40 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:36 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 6.00 cm | 6.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.1% | 100% | 23.42 cm | 2.53 cm - 2.53 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:38 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 6.50 cm | 6.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 23.40 cm | 2.60 cm - 2.60 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:39 | Electrons | 110.00 cm | 16 MeV | 20x20 | Crossline | 7.00 cm | 7.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 23.38 cm | 2.61 cm - 2.61 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:01 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 21.01 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:02 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 1.00 cm | 1.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 21.13 cm | 0.41 cm - 0.41 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:03 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 21.35 cm | 0.58 cm - 0.58 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:04 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 21.47 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:06 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 21.59 cm | 0.91 cm - 0.91 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:07 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 4.00 cm | 4.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 21.86 cm | 1.38 cm - 1.38 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:08 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 4.50 cm | 4.50 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 21.97 cm | 1.67 cm - 1.67 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:09 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 5.00 cm | 5.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 22.07 cm | 1.93 cm - 1.93 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:10 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 5.50 cm | 5.50 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.2% | 100% | 22.14 cm | 2.18 cm - 2.18 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:11 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 6.00 cm | 6.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.8% | 100% | 22.18 cm | 2.35 cm - 2.35 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:13 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 6.50 cm | 6.50 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.1% | 100% | 22.19 cm | 2.46 cm - 2.46 cm | 0.00 cm
| TR2 | 2017-08-25 | 13:14 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 7.00 cm | 7.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 22.18 cm | 2.49 cm - 2.49 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:32 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 25.17 cm | 0.36 cm - 0.36 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:33 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 1.00 cm | 1.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 25.31 cm | 0.40 cm - 0.40 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:34 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 25.57 cm | 0.57 cm - 0.57 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:36 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 25.73 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:37 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 25.87 cm | 0.91 cm - 0.91 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:38 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 4.00 cm | 4.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 26.18 cm | 1.41 cm - 1.41 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:40 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 4.50 cm | 4.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 26.31 cm | 1.68 cm - 1.68 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:41 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 5.00 cm | 5.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 26.43 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:42 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 5.50 cm | 5.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 26.49 cm | 2.18 cm - 2.18 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:44 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 6.00 cm | 6.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 26.54 cm | 2.36 cm - 2.36 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:45 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 6.50 cm | 6.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 26.53 cm | 2.44 cm - 2.44 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:46 | Electrons | 100.00 cm | 16 MeV | 25x25 | Crossline | 7.00 cm | 7.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 26.53 cm | 2.47 cm - 2.47 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:44 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 27.89 cm | 0.74 cm - 0.74 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:45 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 1.00 cm | 1.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 28.03 cm | 0.78 cm - 0.78 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:47 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 28.30 cm | 0.95 cm - 0.95 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:48 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 28.46 cm | 1.07 cm - 1.07 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:49 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 28.60 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:51 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 4.00 cm | 4.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 28.91 cm | 1.70 cm - 1.70 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:52 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 4.50 cm | 4.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 29.03 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:54 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 5.00 cm | 5.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 29.13 cm | 2.18 cm - 2.18 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:55 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 5.50 cm | 5.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 29.18 cm | 2.38 cm - 2.38 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:56 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 6.00 cm | 6.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 29.21 cm | 2.51 cm - 2.51 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:58 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 6.50 cm | 6.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 29.19 cm | 2.57 cm - 2.57 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:59 | Electrons | 110.00 cm | 16 MeV | 25x25 | Crossline | 7.00 cm | 7.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 29.17 cm | 2.57 cm - 2.57 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:20 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 3.14 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:21 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4% | 100% | 3.15 cm | 0.41 cm - 0.41 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:22 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.8% | 100% | 3.19 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:22 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.3% | 100% | 3.21 cm | 0.65 cm - 0.65 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:23 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.8% | 0% | 3.24 cm | 0.78 cm - 0.78 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:23 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.9% | 0% | 3.32 cm | 1.10 cm - 1.10 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:24 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.4% | 0% | 3.50 cm | 1.45 cm - 1.45 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:25 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 6.00 cm | 6.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.5% | 0% | 3.81 cm | 1.82 cm - 1.82 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:25 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 7.00 cm | 7.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 21.4% | 0% | 4.23 cm | 2.17 cm - 2.17 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:26 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 8.00 cm | 8.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 22.3% | 0% | 4.60 cm | 2.51 cm - 2.51 cm | 0.00 cm
| TR2 | 2017-08-28 | 16:27 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 9.00 cm | 9.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 22.6% | 0% | 4.74 cm | 2.97 cm - 2.97 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:33 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 4.04 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:33 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 4.07 cm | 0.41 cm - 0.41 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:34 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.8% | 100% | 4.11 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:35 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7% | 100% | 4.14 cm | 0.65 cm - 0.65 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:35 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.5% | 100% | 4.16 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:36 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.4% | 100% | 4.22 cm | 1.14 cm - 1.14 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:36 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.9% | 0% | 4.31 cm | 1.54 cm - 1.54 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:37 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 6.00 cm | 6.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.7% | 0% | 4.50 cm | 1.92 cm - 1.92 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:38 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 7.00 cm | 7.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.5% | 0% | 4.80 cm | 2.28 cm - 2.28 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:38 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 8.00 cm | 8.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.7% | 0% | 5.08 cm | 2.54 cm - 2.54 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:39 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 9.00 cm | 9.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.4% | 0% | 5.17 cm | 2.81 cm - 2.81 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:02 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2% | 100% | 6.74 cm | 0.63 cm - 0.63 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:03 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 6.77 cm | 0.66 cm - 0.66 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:04 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.8% | 100% | 6.84 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:05 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5% | 100% | 6.88 cm | 0.89 cm - 0.89 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:05 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.5% | 100% | 6.91 cm | 1.01 cm - 1.01 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:06 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10% | 100% | 7.00 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:07 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.7% | 100% | 7.08 cm | 1.79 cm - 1.79 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:08 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 6.00 cm | 6.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.5% | 100% | 7.17 cm | 2.24 cm - 2.24 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:09 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 7.00 cm | 7.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.4% | 100% | 7.26 cm | 2.62 cm - 2.62 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:10 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 8.00 cm | 8.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.3% | 100% | 7.31 cm | 2.87 cm - 2.87 cm | 0.00 cm
| TR2 | 2017-08-29 | 16:11 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 9.00 cm | 9.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.5% | 100% | 7.33 cm | 2.97 cm - 2.97 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:46 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.2% | 100% | 6.09 cm | 0.37 cm - 0.37 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:46 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 6.12 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:47 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 6.19 cm | 0.53 cm - 0.53 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:48 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 2.50 cm | 2.50 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.9% | 100% | 6.23 cm | 0.64 cm - 0.64 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:48 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.6% | 100% | 6.26 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:49 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.2% | 100% | 6.34 cm | 1.17 cm - 1.17 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:49 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.6% | 100% | 6.43 cm | 1.63 cm - 1.63 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:50 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 6.00 cm | 6.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 16.8% | 100% | 6.53 cm | 2.10 cm - 2.10 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:51 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 7.00 cm | 7.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.6% | 100% | 6.63 cm | 2.49 cm - 2.49 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:51 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 8.00 cm | 8.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.5% | 100% | 6.71 cm | 2.73 cm - 2.73 cm | 0.00 cm
| TR2 | 2017-08-28 | 14:52 | Electrons | 100.00 cm | 20 MeV | 6x6 | Crossline | 9.00 cm | 9.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.7% | 100% | 6.74 cm | 2.85 cm - 2.85 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:31 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.1% | 100% | 3.00 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:31 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.1% | 100% | 3.02 cm | 0.42 cm - 0.42 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:32 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.8% | 100% | 3.05 cm | 0.55 cm - 0.55 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:33 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.2% | 0% | 3.08 cm | 0.66 cm - 0.66 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:33 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.7% | 0% | 3.10 cm | 0.80 cm - 0.80 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:34 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.7% | 0% | 3.19 cm | 1.12 cm - 1.12 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:34 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.8% | 0% | 3.38 cm | 1.47 cm - 1.47 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:35 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 6.00 cm | 6.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.9% | 0% | 3.73 cm | 1.85 cm - 1.85 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:36 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 7.00 cm | 7.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 21.7% | 0% | 4.16 cm | 2.23 cm - 2.23 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:36 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 8.00 cm | 8.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 22.7% | 0% | 4.57 cm | 2.53 cm - 2.53 cm | 0.00 cm
| TR2 | 2017-08-28 | 11:37 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 9.00 cm | 9.00 cm | 3.00 cm | 3.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 23.1% | 0% | 4.74 cm | 2.93 cm - 2.93 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:04 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 7.01 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:04 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 7.05 cm | 0.41 cm - 0.41 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:05 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 7.13 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:06 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.2% | 100% | 7.16 cm | 0.64 cm - 0.64 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:07 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.6% | 100% | 7.20 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:08 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.4% | 100% | 7.28 cm | 1.15 cm - 1.15 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:08 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.7% | 100% | 7.38 cm | 1.60 cm - 1.60 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:09 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 6.00 cm | 6.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.1% | 100% | 7.47 cm | 2.09 cm - 2.09 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:10 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 7.00 cm | 7.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.4% | 100% | 7.56 cm | 2.53 cm - 2.53 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:11 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 8.00 cm | 8.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.7% | 100% | 7.62 cm | 2.82 cm - 2.82 cm | 0.00 cm
| TR2 | 2017-08-25 | 15:12 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 9.00 cm | 9.00 cm | 7.00 cm | 7.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19% | 100% | 7.66 cm | 2.97 cm - 2.97 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:42 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 11.20 cm | 0.63 cm - 0.63 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:43 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 11.25 cm | 0.66 cm - 0.66 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:44 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 11.36 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:45 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 11.42 cm | 0.89 cm - 0.89 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:46 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 11.47 cm | 1.02 cm - 1.02 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:47 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.6% | 100% | 11.59 cm | 1.38 cm - 1.38 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:48 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.4% | 100% | 11.70 cm | 1.83 cm - 1.83 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:50 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 6.00 cm | 6.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.5% | 100% | 11.80 cm | 2.33 cm - 2.33 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:51 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 7.00 cm | 7.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.7% | 100% | 11.86 cm | 2.75 cm - 2.75 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:52 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 8.00 cm | 8.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.3% | 100% | 11.88 cm | 3.04 cm - 3.04 cm | 0.00 cm
| TR2 | 2017-08-29 | 14:53 | Electrons | 110.00 cm | 20 MeV | 10x10 | Crossline | 9.00 cm | 9.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.8% | 100% | 11.90 cm | 3.16 cm - 3.16 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:24 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 0.00 cm | 0.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.2% | 100% | 10.08 cm | 0.34 cm - 0.34 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:25 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 10.14 cm | 0.36 cm - 0.36 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:26 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 10.18 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:27 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 1.50 cm | 1.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 10.24 cm | 0.43 cm - 0.43 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:28 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 10.29 cm | 0.51 cm - 0.51 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:28 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 10.35 cm | 0.61 cm - 0.61 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:29 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 10.40 cm | 0.75 cm - 0.75 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:30 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 10.46 cm | 0.92 cm - 0.92 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:31 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 4.00 cm | 4.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.6% | 100% | 10.52 cm | 1.12 cm - 1.12 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:32 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.2% | 100% | 10.57 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:33 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 5.00 cm | 5.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.9% | 100% | 10.63 cm | 1.60 cm - 1.60 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:34 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 5.50 cm | 5.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.7% | 100% | 10.69 cm | 1.85 cm - 1.85 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:34 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 6.00 cm | 6.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.4% | 100% | 10.73 cm | 2.12 cm - 2.12 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:35 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 6.50 cm | 6.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.9% | 100% | 10.78 cm | 2.38 cm - 2.38 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:36 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 7.00 cm | 7.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.3% | 100% | 10.81 cm | 2.61 cm - 2.61 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:37 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 8.00 cm | 8.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.9% | 100% | 10.83 cm | 2.92 cm - 2.92 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:38 | Electrons | 100.00 cm | 20 MeV | 10x10 | Crossline | 9.00 cm | 9.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.5% | 100% | 10.83 cm | 3.05 cm - 3.05 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:35 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 11.06 cm | 0.37 cm - 0.37 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:36 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 1.00 cm | 1.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 11.12 cm | 0.40 cm - 0.40 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:37 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 11.23 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:38 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 11.29 cm | 0.65 cm - 0.65 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:39 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 11.35 cm | 0.80 cm - 0.80 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:40 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.4% | 100% | 11.48 cm | 1.19 cm - 1.19 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:41 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 5.00 cm | 5.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.5% | 100% | 11.60 cm | 1.67 cm - 1.67 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:42 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 6.00 cm | 6.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.7% | 100% | 11.71 cm | 2.19 cm - 2.19 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:43 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 7.00 cm | 7.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.4% | 100% | 11.78 cm | 2.67 cm - 2.67 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:44 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 8.00 cm | 8.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14% | 100% | 11.80 cm | 2.98 cm - 2.98 cm | 0.00 cm
| TR2 | 2017-08-28 | 13:45 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 9.00 cm | 9.00 cm | 11.00 cm | 11.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.8% | 100% | 11.83 cm | 3.13 cm - 3.13 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:19 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.2% | 100% | 15.14 cm | 0.37 cm - 0.37 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:20 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 1.00 cm | 1.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.4% | 100% | 15.22 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:21 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 15.38 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:22 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 15.45 cm | 0.65 cm - 0.65 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:24 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 15.54 cm | 0.81 cm - 0.81 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:25 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.1% | 100% | 15.70 cm | 1.19 cm - 1.19 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:26 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 5.00 cm | 5.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.7% | 100% | 15.88 cm | 1.68 cm - 1.68 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:27 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 6.00 cm | 6.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.1% | 100% | 16.02 cm | 2.22 cm - 2.22 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:28 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 7.00 cm | 7.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.5% | 100% | 16.11 cm | 2.69 cm - 2.69 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:29 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 8.00 cm | 8.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10% | 100% | 16.12 cm | 2.98 cm - 2.98 cm | 0.00 cm
| TR2 | 2017-08-28 | 12:30 | Electrons | 100.00 cm | 20 MeV | 15x15 | Crossline | 9.00 cm | 9.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11% | 100% | 16.09 cm | 3.14 cm - 3.14 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:25 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 0.50 cm | 0.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 16.75 cm | 0.64 cm - 0.64 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:26 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 1.00 cm | 1.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 16.84 cm | 0.68 cm - 0.68 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:27 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 2.00 cm | 2.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 16.99 cm | 0.81 cm - 0.81 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:28 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 2.50 cm | 2.50 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 17.08 cm | 0.92 cm - 0.92 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:29 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 3.00 cm | 3.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 17.15 cm | 1.06 cm - 1.06 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:30 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 4.00 cm | 4.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 17.33 cm | 1.41 cm - 1.41 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:32 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 5.00 cm | 5.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4% | 100% | 17.48 cm | 1.86 cm - 1.86 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:33 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 6.00 cm | 6.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6% | 100% | 17.63 cm | 2.36 cm - 2.36 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:34 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 7.00 cm | 7.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.1% | 100% | 17.71 cm | 2.79 cm - 2.79 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:35 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 8.00 cm | 8.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.5% | 100% | 17.72 cm | 3.10 cm - 3.10 cm | 0.00 cm
| TR2 | 2017-08-29 | 13:36 | Electrons | 110.00 cm | 20 MeV | 15x15 | Crossline | 9.00 cm | 9.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.3% | 100% | 17.64 cm | 3.24 cm - 3.24 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:08 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 22.35 cm | 0.64 cm - 0.64 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:09 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 1.00 cm | 1.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 22.44 cm | 0.68 cm - 0.68 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:10 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 22.66 cm | 0.80 cm - 0.80 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:12 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 22.77 cm | 0.91 cm - 0.91 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:13 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 22.89 cm | 1.04 cm - 1.04 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:14 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 4.00 cm | 4.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 23.12 cm | 1.40 cm - 1.40 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:15 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 5.00 cm | 5.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.9% | 100% | 23.35 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:16 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 6.00 cm | 6.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3% | 100% | 23.54 cm | 2.34 cm - 2.34 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:18 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 7.00 cm | 7.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.5% | 100% | 23.64 cm | 2.78 cm - 2.78 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:19 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 8.00 cm | 8.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.5% | 100% | 23.68 cm | 3.07 cm - 3.07 cm | 0.00 cm
| TR2 | 2017-08-29 | 12:20 | Electrons | 110.00 cm | 20 MeV | 20x20 | Crossline | 9.00 cm | 9.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.4% | 100% | 23.60 cm | 3.19 cm - 3.19 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:12 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 0.50 cm | 0.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.23 cm | 0.35 cm - 0.35 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:13 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 1.00 cm | 1.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 20.32 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:14 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 2.00 cm | 2.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 20.54 cm | 0.53 cm - 0.53 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:15 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 2.50 cm | 2.50 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 20.64 cm | 0.64 cm - 0.64 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:16 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 3.00 cm | 3.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 20.75 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:17 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 4.00 cm | 4.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 20.97 cm | 1.17 cm - 1.17 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:19 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 5.00 cm | 5.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.6% | 100% | 21.21 cm | 1.65 cm - 1.65 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:20 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 6.00 cm | 6.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3% | 100% | 21.40 cm | 2.18 cm - 2.18 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:21 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 7.00 cm | 7.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.9% | 100% | 21.52 cm | 2.65 cm - 2.65 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:22 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 8.00 cm | 8.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.1% | 100% | 21.53 cm | 2.95 cm - 2.95 cm | 0.00 cm
| TR2 | 2017-08-28 | 10:23 | Electrons | 100.00 cm | 20 MeV | 20x20 | Crossline | 9.00 cm | 9.00 cm | 20.00 cm | 20.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.2% | 100% | 21.48 cm | 3.13 cm - 3.13 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:45 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 21.02 cm | 0.36 cm - 0.36 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:47 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 1.00 cm | 1.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.6% | 100% | 21.13 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:48 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 21.35 cm | 0.51 cm - 0.51 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:49 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 21.47 cm | 0.62 cm - 0.62 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:50 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 21.58 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:51 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 4.00 cm | 4.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 21.82 cm | 1.13 cm - 1.13 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:52 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 5.00 cm | 5.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 22.07 cm | 1.60 cm - 1.60 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:54 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 6.00 cm | 6.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 22.29 cm | 2.13 cm - 2.13 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:55 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 7.00 cm | 7.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.2% | 100% | 22.44 cm | 2.62 cm - 2.62 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:56 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 8.00 cm | 8.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.4% | 100% | 22.50 cm | 2.94 cm - 2.94 cm | 0.00 cm
| TR2 | 2017-08-25 | 12:57 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 9.00 cm | 9.00 cm | 21.00 cm | 21.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.5% | 100% | 22.48 cm | 3.13 cm - 3.13 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:26 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 27.89 cm | 0.64 cm - 0.64 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:27 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 1.00 cm | 1.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 28.01 cm | 0.67 cm - 0.67 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:29 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 28.28 cm | 0.81 cm - 0.81 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:30 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 28.42 cm | 0.91 cm - 0.91 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:31 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 28.56 cm | 1.04 cm - 1.04 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:33 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 4.00 cm | 4.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 28.84 cm | 1.39 cm - 1.39 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:34 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 5.00 cm | 5.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 29.14 cm | 1.84 cm - 1.84 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:36 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 6.00 cm | 6.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 29.36 cm | 2.32 cm - 2.32 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:37 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 7.00 cm | 7.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 29.51 cm | 2.74 cm - 2.74 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:38 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 8.00 cm | 8.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.1% | 100% | 29.53 cm | 3.03 cm - 3.03 cm | 0.00 cm
| TR2 | 2017-08-29 | 10:40 | Electrons | 110.00 cm | 20 MeV | 25x25 | Crossline | 9.00 cm | 9.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4% | 100% | 29.48 cm | 3.14 cm - 3.14 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:12 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 0.50 cm | 0.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 25.18 cm | 0.34 cm - 0.34 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:14 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 1.00 cm | 1.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.3% | 100% | 25.31 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:15 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 2.00 cm | 2.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 25.57 cm | 0.51 cm - 0.51 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:16 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 2.50 cm | 2.50 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.1% | 100% | 25.71 cm | 0.62 cm - 0.62 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:17 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 3.00 cm | 3.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 25.84 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:19 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 4.00 cm | 4.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 26.14 cm | 1.14 cm - 1.14 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:20 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 5.00 cm | 5.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 26.42 cm | 1.61 cm - 1.61 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:21 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 6.00 cm | 6.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 26.69 cm | 2.13 cm - 2.13 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:23 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 7.00 cm | 7.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.6% | 100% | 26.84 cm | 2.61 cm - 2.61 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:24 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 8.00 cm | 8.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.5% | 100% | 26.90 cm | 2.93 cm - 2.93 cm | 0.00 cm
| TR2 | 2017-08-25 | 11:25 | Electrons | 100.00 cm | 20 MeV | 25x25 | Crossline | 9.00 cm | 9.00 cm | 25.00 cm | 25.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.2% | 100% | 26.86 cm | 3.07 cm - 3.07 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:24 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.4% | 100% | 4.47 cm | 0.64 cm - 0.64 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:25 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.8% | 100% | 4.49 cm | 0.67 cm - 0.67 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:26 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.8% | 100% | 4.54 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:26 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.8% | 100% | 4.59 cm | 1.00 cm - 1.00 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:27 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.1% | 100% | 4.64 cm | 1.30 cm - 1.30 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:28 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18% | 100% | 4.73 cm | 1.68 cm - 1.68 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:28 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 6.00 cm | 6.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.4% | 100% | 4.90 cm | 2.05 cm - 2.05 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:29 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 7.00 cm | 7.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.2% | 100% | 5.16 cm | 2.37 cm - 2.37 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:30 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 8.00 cm | 8.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.3% | 100% | 5.38 cm | 2.60 cm - 2.60 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:30 | Electrons | 110.00 cm | 20 MeV | 6x6 | Crossline | 9.00 cm | 9.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.8% | 100% | 5.44 cm | 2.82 cm - 2.82 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:17 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 0.80 cm | 0.80 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.7% | 100% | 10.14 cm | 0.74 cm - 0.74 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:18 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.00 cm | 1.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.1% | 100% | 10.16 cm | 0.83 cm - 0.83 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:19 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 1.40 cm | 1.40 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.1% | 100% | 10.19 cm | 1.03 cm - 1.03 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:20 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.6% | 100% | 10.19 cm | 1.29 cm - 1.29 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:20 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.6% | 100% | 10.17 cm | 1.31 cm - 1.31 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:21 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.9% | 100% | 10.19 cm | 1.25 cm - 1.25 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:16 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 0.00 cm | 0.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.2% | 100% | 10.05 cm | 0.58 cm - 0.58 cm | 0.00 cm
| TR2 | 2017-09-18 | 14:18 | Electrons | 100.00 cm | 6 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 10.11 cm | 0.66 cm - 0.66 cm | 0.00 cm
| TR2 | 2017-09-18 | 14:22 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.3% | 100% | 10.12 cm | 0.48 cm - 0.48 cm | 0.00 cm
| TR2 | 2017-09-18 | 14:23 | Electrons | 100.00 cm | 9 MeV | 10x10 | Crossline | 4.50 cm | 4.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6% | 100% | 10.41 cm | 1.55 cm - 1.55 cm | 0.00 cm
| TR2 | 2017-09-18 | 14:24 | Electrons | 100.00 cm | 12 MeV | 10x10 | Crossline | 3.50 cm | 3.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7% | 100% | 10.45 cm | 1.56 cm - 1.56 cm | 0.00 cm
| TR2 | 2017-09-18 | 14:26 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 6.00 cm | 6.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.5% | 100% | 10.66 cm | 2.41 cm - 2.41 cm | 0.00 cm
| TR2 | 2017-09-18 | 14:28 | Electrons | 100.00 cm | 16 MeV | 10x10 | Crossline | 7.00 cm | 7.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.2% | 100% | 10.67 cm | 2.53 cm - 2.53 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:28 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 7.00 cm | 7.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.6% | 0% | 4.72 cm | 2.24 cm - 2.24 cm | 0.00 cm
| TR2 | 2017-09-18 | 15:26 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 6.00 cm | 6.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.6% | 0% | 4.62 cm | 2.11 cm - 2.11 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:02 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 6.00 cm | 6.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 18.6% | 100% | 6.46 cm | 2.34 cm - 2.34 cm | 0.00 cm
| TR2 | 2017-08-28 | 15:03 | Electrons | 100.00 cm | 16 MeV | 6x6 | Crossline | 7.00 cm | 7.00 cm | 6.00 cm | 6.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.2% | 100% | 6.48 cm | 2.48 cm - 2.48 cm | 0.00 cm
| TR2 | 2017-09-18 | 12:20 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 7.00 cm | 7.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.9% | 100% | 15.91 cm | 2.55 cm - 2.55 cm | 0.00 cm
| TR2 | 2017-09-18 | 12:18 | Electrons | 100.00 cm | 16 MeV | 15x15 | Crossline | 6.00 cm | 6.00 cm | 15.00 cm | 15.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.2% | 100% | 15.93 cm | 2.43 cm - 2.43 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:16 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 0.50 cm | 0.50 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.2% | 100% | 4.46 cm | 0.72 cm - 0.72 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:19 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 1.00 cm | 1.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.7% | 100% | 4.48 cm | 0.76 cm - 0.76 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:19 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 2.00 cm | 2.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.9% | 100% | 4.53 cm | 0.91 cm - 0.91 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:20 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 3.00 cm | 3.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 14.2% | 100% | 4.57 cm | 1.18 cm - 1.18 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:21 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 4.00 cm | 4.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 17.5% | 100% | 4.65 cm | 1.56 cm - 1.56 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:22 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 5.00 cm | 5.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 19.5% | 0% | 4.81 cm | 1.95 cm - 1.95 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:22 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 6.00 cm | 6.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.5% | 0% | 5.01 cm | 2.23 cm - 2.23 cm | 0.00 cm
| TR2 | 2017-09-19 | 15:23 | Electrons | 110.00 cm | 16 MeV | 6x6 | Crossline | 7.00 cm | 7.00 cm | 4.00 cm | 4.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 20.5% | 0% | 5.06 cm | 2.35 cm - 2.35 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:58 | Electrons | 100.00 cm | 9 MeV | 10x10 | Inline | 4.50 cm | 4.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.8% | 100% | 10.39 cm | 1.57 cm - 1.57 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:59 | Electrons | 100.00 cm | 9 MeV | 10x10 | Inline | 4.00 cm | 4.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.2% | 100% | 10.32 cm | 1.59 cm - 1.59 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:00 | Electrons | 100.00 cm | 9 MeV | 10x10 | Inline | 3.50 cm | 3.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.8% | 100% | 10.31 cm | 1.65 cm - 1.65 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:01 | Electrons | 100.00 cm | 9 MeV | 10x10 | Inline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.3% | 100% | 10.33 cm | 1.57 cm - 1.57 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:02 | Electrons | 100.00 cm | 9 MeV | 10x10 | Inline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.5% | 100% | 10.30 cm | 1.35 cm - 1.35 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:03 | Electrons | 100.00 cm | 9 MeV | 10x10 | Inline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.3% | 100% | 10.27 cm | 1.06 cm - 1.06 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:04 | Electrons | 100.00 cm | 9 MeV | 10x10 | Inline | 1.40 cm | 1.40 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 10.19 cm | 0.73 cm - 0.73 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:05 | Electrons | 100.00 cm | 9 MeV | 10x10 | Inline | 1.00 cm | 1.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 10.15 cm | 0.59 cm - 0.59 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:06 | Electrons | 100.00 cm | 9 MeV | 10x10 | Inline | 0.80 cm | 0.80 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.5% | 100% | 10.12 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:33 | Electrons | 100.00 cm | 12 MeV | 10x10 | Inline | 6.00 cm | 6.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.7% | 100% | 10.51 cm | 2.05 cm - 2.05 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:33 | Electrons | 100.00 cm | 12 MeV | 10x10 | Inline | 5.50 cm | 5.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.9% | 100% | 10.42 cm | 2.06 cm - 2.06 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:34 | Electrons | 100.00 cm | 12 MeV | 10x10 | Inline | 5.00 cm | 5.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.3% | 100% | 10.42 cm | 2.10 cm - 2.10 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:35 | Electrons | 100.00 cm | 12 MeV | 10x10 | Inline | 4.50 cm | 4.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.9% | 100% | 10.41 cm | 2.05 cm - 2.05 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:36 | Electrons | 100.00 cm | 12 MeV | 10x10 | Inline | 4.00 cm | 4.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.8% | 100% | 10.42 cm | 1.89 cm - 1.89 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:37 | Electrons | 100.00 cm | 12 MeV | 10x10 | Inline | 3.50 cm | 3.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.9% | 100% | 10.40 cm | 1.64 cm - 1.64 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:38 | Electrons | 100.00 cm | 12 MeV | 10x10 | Inline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.7% | 100% | 10.36 cm | 1.34 cm - 1.34 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:38 | Electrons | 100.00 cm | 12 MeV | 10x10 | Inline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.7% | 100% | 10.31 cm | 1.03 cm - 1.03 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:39 | Electrons | 100.00 cm | 12 MeV | 10x10 | Inline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.5% | 100% | 10.26 cm | 0.79 cm - 0.79 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:40 | Electrons | 100.00 cm | 12 MeV | 10x10 | Inline | 1.50 cm | 1.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 10.20 cm | 0.60 cm - 0.60 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:41 | Electrons | 100.00 cm | 12 MeV | 10x10 | Inline | 1.00 cm | 1.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.5% | 100% | 10.15 cm | 0.49 cm - 0.49 cm | 0.00 cm
| TR2 | 2017-08-24 | 16:42 | Electrons | 100.00 cm | 12 MeV | 10x10 | Inline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 10.09 cm | 0.43 cm - 0.43 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:09 | Electrons | 100.00 cm | 16 MeV | 10x10 | Inline | 7.00 cm | 7.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.7% | 100% | 10.63 cm | 2.57 cm - 2.57 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:10 | Electrons | 100.00 cm | 16 MeV | 10x10 | Inline | 6.50 cm | 6.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.6% | 100% | 10.63 cm | 2.54 cm - 2.54 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:11 | Electrons | 100.00 cm | 16 MeV | 10x10 | Inline | 6.00 cm | 6.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.1% | 100% | 10.62 cm | 2.45 cm - 2.45 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:12 | Electrons | 100.00 cm | 16 MeV | 10x10 | Inline | 5.50 cm | 5.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12% | 100% | 10.61 cm | 2.29 cm - 2.29 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:13 | Electrons | 100.00 cm | 16 MeV | 10x10 | Inline | 5.00 cm | 5.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 10.5% | 100% | 10.59 cm | 2.06 cm - 2.06 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:14 | Electrons | 100.00 cm | 16 MeV | 10x10 | Inline | 4.50 cm | 4.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 8.5% | 100% | 10.54 cm | 1.78 cm - 1.78 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:15 | Electrons | 100.00 cm | 16 MeV | 10x10 | Inline | 4.00 cm | 4.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 6.5% | 100% | 10.50 cm | 1.50 cm - 1.50 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:16 | Electrons | 100.00 cm | 16 MeV | 10x10 | Inline | 3.50 cm | 3.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.5% | 100% | 10.43 cm | 1.23 cm - 1.23 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:16 | Electrons | 100.00 cm | 16 MeV | 10x10 | Inline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.9% | 100% | 10.39 cm | 0.98 cm - 0.98 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:17 | Electrons | 100.00 cm | 16 MeV | 10x10 | Inline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.8% | 100% | 10.32 cm | 0.78 cm - 0.78 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:18 | Electrons | 100.00 cm | 16 MeV | 10x10 | Inline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.2% | 100% | 10.27 cm | 0.62 cm - 0.62 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:23 | Electrons | 100.00 cm | 16 MeV | 10x10 | Inline | 1.00 cm | 1.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.9% | 100% | 10.15 cm | 0.42 cm - 0.42 cm | 0.00 cm
| TR2 | 2017-08-24 | 17:24 | Electrons | 100.00 cm | 16 MeV | 10x10 | Inline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.7% | 100% | 10.10 cm | 0.39 cm - 0.39 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:39 | Electrons | 100.00 cm | 20 MeV | 10x10 | Inline | 9.00 cm | 9.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.9% | 100% | 10.81 cm | 3.07 cm - 3.07 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:40 | Electrons | 100.00 cm | 20 MeV | 10x10 | Inline | 8.00 cm | 8.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 15.3% | 100% | 10.79 cm | 2.96 cm - 2.96 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:41 | Electrons | 100.00 cm | 20 MeV | 10x10 | Inline | 7.00 cm | 7.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 13.9% | 100% | 10.77 cm | 2.66 cm - 2.66 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:42 | Electrons | 100.00 cm | 20 MeV | 10x10 | Inline | 6.50 cm | 6.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 12.6% | 100% | 10.74 cm | 2.45 cm - 2.45 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:42 | Electrons | 100.00 cm | 20 MeV | 10x10 | Inline | 6.00 cm | 6.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 11.2% | 100% | 10.71 cm | 2.20 cm - 2.20 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:43 | Electrons | 100.00 cm | 20 MeV | 10x10 | Inline | 5.50 cm | 5.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 9.5% | 100% | 10.65 cm | 1.94 cm - 1.94 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:44 | Electrons | 100.00 cm | 20 MeV | 10x10 | Inline | 5.00 cm | 5.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 7.8% | 100% | 10.61 cm | 1.68 cm - 1.68 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:45 | Electrons | 100.00 cm | 20 MeV | 10x10 | Inline | 4.50 cm | 4.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.9% | 100% | 10.54 cm | 1.42 cm - 1.42 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:47 | Electrons | 100.00 cm | 20 MeV | 10x10 | Inline | 4.00 cm | 4.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.4% | 100% | 10.50 cm | 1.19 cm - 1.19 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:47 | Electrons | 100.00 cm | 20 MeV | 10x10 | Inline | 3.50 cm | 3.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3% | 100% | 10.43 cm | 0.99 cm - 0.99 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:48 | Electrons | 100.00 cm | 20 MeV | 10x10 | Inline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.2% | 100% | 10.39 cm | 0.81 cm - 0.81 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:49 | Electrons | 100.00 cm | 20 MeV | 10x10 | Inline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.7% | 100% | 10.32 cm | 0.65 cm - 0.65 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:50 | Electrons | 100.00 cm | 20 MeV | 10x10 | Inline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1.4% | 100% | 10.27 cm | 0.54 cm - 0.54 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:51 | Electrons | 100.00 cm | 20 MeV | 10x10 | Inline | 1.00 cm | 1.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 1% | 100% | 10.16 cm | 0.40 cm - 0.40 cm | 0.00 cm
| TR2 | 2017-08-25 | 14:52 | Electrons | 100.00 cm | 20 MeV | 10x10 | Inline | 0.50 cm | 0.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 0.8% | 100% | 10.11 cm | 0.38 cm - 0.38 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:22 | Electrons | 100.00 cm | 6 MeV | 10x10 | Inline | 0.00 cm | 0.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.2% | 100% | 10.03 cm | 0.60 cm - 0.60 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:23 | Electrons | 100.00 cm | 6 MeV | 10x10 | Inline | 0.80 cm | 0.80 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 2.8% | 100% | 10.11 cm | 0.74 cm - 0.74 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:24 | Electrons | 100.00 cm | 6 MeV | 10x10 | Inline | 1.00 cm | 1.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 3.2% | 100% | 10.14 cm | 0.83 cm - 0.83 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:24 | Electrons | 100.00 cm | 6 MeV | 10x10 | Inline | 1.40 cm | 1.40 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 4.2% | 100% | 10.16 cm | 1.04 cm - 1.04 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:25 | Electrons | 100.00 cm | 6 MeV | 10x10 | Inline | 2.00 cm | 2.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.8% | 100% | 10.17 cm | 1.31 cm - 1.31 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:26 | Electrons | 100.00 cm | 6 MeV | 10x10 | Inline | 2.50 cm | 2.50 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.9% | 100% | 10.15 cm | 1.34 cm - 1.34 cm | 0.00 cm
| TR2 | 2017-08-24 | 15:27 | Electrons | 100.00 cm | 6 MeV | 10x10 | Inline | 3.00 cm | 3.00 cm | 10.00 cm | 10.00 cm | 0 ° | 0 ° | EFD3G | Absolute | Crossline | Continuous | 5.1% | 100% | 10.17 cm | 1.28 cm - 1.28 cm | 0.00 cm

## TR2 Profile Date Sequence Check
Date | Time | Start position [De] | Scan type | Energy | Field size inline | Applicator | SSD
2017-08-24 | 15:16 | 0.00 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:17 | 0.50 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:17 | 0.80 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:18 | 1.00 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:19 | 1.40 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:20 | 2.00 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:20 | 2.50 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:21 | 3.00 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | FALSE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:22 | 0.00 cm | Inline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:23 | 0.80 cm | Inline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:24 | 1.00 cm | Inline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:24 | 1.40 cm | Inline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:25 | 2.00 cm | Inline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:26 | 2.50 cm | Inline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:27 | 3.00 cm | Inline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-24 | 15:47 | 0.00 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:48 | 0.50 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:49 | 0.80 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:50 | 1.00 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:51 | 1.40 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:52 | 2.00 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:53 | 2.50 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:54 | 3.00 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:55 | 3.50 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:56 | 4.00 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 15:57 | 4.50 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-24 | 16:22 | 0.00 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 16:23 | 0.50 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 16:23 | 1.00 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 16:24 | 1.50 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 16:25 | 2.00 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 16:26 | 2.50 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 16:27 | 3.00 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 16:28 | 3.50 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 16:28 | 4.00 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 16:29 | 4.50 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 16:30 | 5.00 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 16:31 | 5.50 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 16:32 | 6.00 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-24 | 16:56 | 0.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 16:57 | 0.50 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 16:58 | 1.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 16:59 | 1.50 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 17:00 | 2.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 17:01 | 2.50 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 17:02 | 3.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 17:02 | 3.50 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 17:03 | 4.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 17:04 | 4.50 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 17:05 | 5.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 17:06 | 5.50 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 17:07 | 6.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 17:08 | 6.50 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-24 | 17:08 | 7.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | FALSE | FALSE | TRUE
2017-08-25 | 11:12 | 0.50 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:14 | 1.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:15 | 2.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:16 | 2.50 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:17 | 3.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:19 | 4.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:20 | 5.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:21 | 6.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:23 | 7.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:24 | 8.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:25 | 9.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-25 | 11:32 | 0.50 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:33 | 1.00 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:34 | 2.00 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:36 | 2.50 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:37 | 3.00 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:38 | 4.00 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:40 | 4.50 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:41 | 5.00 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:42 | 5.50 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:44 | 6.00 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:45 | 6.50 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:46 | 7.00 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-25 | 11:53 | 0.50 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:55 | 1.00 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:56 | 1.40 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:57 | 2.00 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 11:58 | 2.50 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:00 | 3.00 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:01 | 4.00 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:02 | 4.50 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:04 | 5.00 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:05 | 5.50 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-25 | 12:10 | 0.80 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:11 | 1.00 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:13 | 1.40 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:14 | 2.00 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:15 | 2.50 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:17 | 3.00 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:18 | 3.50 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:20 | 4.00 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-25 | 12:24 | 0.80 cm | Crossline | 6 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:25 | 1.00 cm | Crossline | 6 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:27 | 1.40 cm | Crossline | 6 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:28 | 2.00 cm | Crossline | 6 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:29 | 2.50 cm | Crossline | 6 MeV | 25.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:31 | 3.00 cm | Crossline | 6 MeV | 25.00 cm | 25x25 | 100.00 cm | FALSE | FALSE | FALSE | TRUE | TRUE
2017-08-25 | 12:45 | 0.50 cm | Crossline | 20 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:47 | 1.00 cm | Crossline | 20 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:48 | 2.00 cm | Crossline | 20 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:49 | 2.50 cm | Crossline | 20 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:50 | 3.00 cm | Crossline | 20 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:51 | 4.00 cm | Crossline | 20 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:52 | 5.00 cm | Crossline | 20 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:54 | 6.00 cm | Crossline | 20 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:55 | 7.00 cm | Crossline | 20 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:56 | 8.00 cm | Crossline | 20 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 12:57 | 9.00 cm | Crossline | 20 MeV | 21.00 cm | 25x25 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-25 | 13:01 | 0.50 cm | Crossline | 16 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:02 | 1.00 cm | Crossline | 16 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:03 | 2.00 cm | Crossline | 16 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:04 | 2.50 cm | Crossline | 16 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:06 | 3.00 cm | Crossline | 16 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:07 | 4.00 cm | Crossline | 16 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:08 | 4.50 cm | Crossline | 16 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:09 | 5.00 cm | Crossline | 16 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:10 | 5.50 cm | Crossline | 16 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:11 | 6.00 cm | Crossline | 16 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:13 | 6.50 cm | Crossline | 16 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:14 | 7.00 cm | Crossline | 16 MeV | 21.00 cm | 25x25 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-25 | 13:36 | 0.50 cm | Crossline | 12 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:38 | 1.00 cm | Crossline | 12 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:39 | 1.40 cm | Crossline | 12 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:40 | 2.00 cm | Crossline | 12 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:41 | 2.50 cm | Crossline | 12 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:42 | 3.00 cm | Crossline | 12 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:43 | 4.00 cm | Crossline | 12 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:45 | 4.50 cm | Crossline | 12 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:46 | 5.00 cm | Crossline | 12 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:47 | 5.50 cm | Crossline | 12 MeV | 21.00 cm | 25x25 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-25 | 13:53 | 0.80 cm | Crossline | 9 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:54 | 1.00 cm | Crossline | 9 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:55 | 1.40 cm | Crossline | 9 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:57 | 2.00 cm | Crossline | 9 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:58 | 2.50 cm | Crossline | 9 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 13:59 | 3.00 cm | Crossline | 9 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:00 | 3.50 cm | Crossline | 9 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:01 | 4.00 cm | Crossline | 9 MeV | 21.00 cm | 25x25 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-25 | 14:05 | 0.80 cm | Crossline | 6 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:06 | 1.00 cm | Crossline | 6 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:07 | 1.40 cm | Crossline | 6 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:08 | 2.00 cm | Crossline | 6 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:10 | 2.50 cm | Crossline | 6 MeV | 21.00 cm | 25x25 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:11 | 3.00 cm | Crossline | 6 MeV | 21.00 cm | 25x25 | 100.00 cm | FALSE | FALSE | FALSE | FALSE | TRUE
2017-08-25 | 14:24 | 0.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:25 | 0.50 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:26 | 1.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:27 | 1.50 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:28 | 2.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:28 | 2.50 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:29 | 3.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:30 | 3.50 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:31 | 4.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:32 | 4.50 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:33 | 5.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:34 | 5.50 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:34 | 6.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:35 | 6.50 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:36 | 7.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:37 | 8.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 14:38 | 9.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 100.00 cm | FALSE | TRUE | FALSE | TRUE | TRUE
2017-08-25 | 15:04 | 0.50 cm | Crossline | 20 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:04 | 1.00 cm | Crossline | 20 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:05 | 2.00 cm | Crossline | 20 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:06 | 2.50 cm | Crossline | 20 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:07 | 3.00 cm | Crossline | 20 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:08 | 4.00 cm | Crossline | 20 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:08 | 5.00 cm | Crossline | 20 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:09 | 6.00 cm | Crossline | 20 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:10 | 7.00 cm | Crossline | 20 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:11 | 8.00 cm | Crossline | 20 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:12 | 9.00 cm | Crossline | 20 MeV | 7.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-25 | 15:26 | 0.50 cm | Crossline | 16 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:27 | 1.00 cm | Crossline | 16 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:27 | 2.00 cm | Crossline | 16 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:28 | 2.50 cm | Crossline | 16 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:29 | 3.00 cm | Crossline | 16 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:30 | 4.00 cm | Crossline | 16 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:30 | 4.50 cm | Crossline | 16 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:31 | 5.00 cm | Crossline | 16 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:32 | 5.50 cm | Crossline | 16 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:33 | 6.00 cm | Crossline | 16 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:34 | 6.50 cm | Crossline | 16 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:34 | 7.00 cm | Crossline | 16 MeV | 7.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-25 | 15:47 | 0.50 cm | Crossline | 12 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:47 | 1.00 cm | Crossline | 12 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:48 | 1.40 cm | Crossline | 12 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:49 | 2.00 cm | Crossline | 12 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:50 | 2.50 cm | Crossline | 12 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:51 | 3.00 cm | Crossline | 12 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:51 | 4.00 cm | Crossline | 12 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:52 | 4.50 cm | Crossline | 12 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:53 | 5.00 cm | Crossline | 12 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 15:54 | 5.50 cm | Crossline | 12 MeV | 7.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-25 | 16:03 | 0.80 cm | Crossline | 9 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 16:03 | 1.00 cm | Crossline | 9 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 16:04 | 1.40 cm | Crossline | 9 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 16:05 | 2.00 cm | Crossline | 9 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 16:06 | 2.50 cm | Crossline | 9 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 16:07 | 3.00 cm | Crossline | 9 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 16:07 | 3.50 cm | Crossline | 9 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 16:08 | 4.00 cm | Crossline | 9 MeV | 7.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-25 | 16:20 | 0.80 cm | Crossline | 6 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 16:21 | 1.00 cm | Crossline | 6 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 16:22 | 1.40 cm | Crossline | 6 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 16:23 | 2.00 cm | Crossline | 6 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 16:24 | 2.50 cm | Crossline | 6 MeV | 7.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-25 | 16:24 | 3.00 cm | Crossline | 6 MeV | 7.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | FALSE | FALSE | TRUE
2017-08-28 | 10:12 | 0.50 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:13 | 1.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:14 | 2.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:15 | 2.50 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:16 | 3.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:17 | 4.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:19 | 5.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:20 | 6.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:21 | 7.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:22 | 8.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:23 | 9.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 10:27 | 0.50 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:28 | 1.00 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:30 | 2.00 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:31 | 2.50 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:32 | 3.00 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:33 | 4.00 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:34 | 4.50 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:35 | 5.00 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:36 | 5.50 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:37 | 6.00 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:39 | 6.50 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:40 | 7.00 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 10:44 | 0.50 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:45 | 1.00 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:46 | 1.40 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:47 | 2.00 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:48 | 2.50 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:50 | 3.00 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:51 | 4.00 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:52 | 4.50 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:53 | 5.00 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:54 | 5.50 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 10:58 | 0.80 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 10:59 | 1.00 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:00 | 1.40 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:01 | 2.00 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:02 | 2.50 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:04 | 3.00 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:05 | 3.50 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:06 | 4.00 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 11:14 | 0.80 cm | Crossline | 6 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:15 | 1.00 cm | Crossline | 6 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:17 | 1.40 cm | Crossline | 6 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:18 | 2.00 cm | Crossline | 6 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:19 | 2.50 cm | Crossline | 6 MeV | 20.00 cm | 20x20 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:20 | 3.00 cm | Crossline | 6 MeV | 20.00 cm | 20x20 | 100.00 cm | FALSE | FALSE | FALSE | FALSE | TRUE
2017-08-28 | 11:31 | 0.50 cm | Crossline | 20 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:31 | 1.00 cm | Crossline | 20 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:32 | 2.00 cm | Crossline | 20 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:33 | 2.50 cm | Crossline | 20 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:33 | 3.00 cm | Crossline | 20 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:34 | 4.00 cm | Crossline | 20 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:34 | 5.00 cm | Crossline | 20 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:35 | 6.00 cm | Crossline | 20 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:36 | 7.00 cm | Crossline | 20 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:36 | 8.00 cm | Crossline | 20 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:37 | 9.00 cm | Crossline | 20 MeV | 3.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 11:44 | 0.50 cm | Crossline | 16 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:44 | 1.00 cm | Crossline | 16 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:45 | 2.00 cm | Crossline | 16 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:45 | 2.50 cm | Crossline | 16 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:46 | 3.00 cm | Crossline | 16 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:47 | 4.00 cm | Crossline | 16 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:47 | 4.50 cm | Crossline | 16 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:48 | 5.00 cm | Crossline | 16 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:48 | 5.50 cm | Crossline | 16 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:49 | 6.00 cm | Crossline | 16 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:50 | 6.50 cm | Crossline | 16 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:50 | 7.00 cm | Crossline | 16 MeV | 3.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 11:53 | 0.50 cm | Crossline | 12 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:54 | 1.00 cm | Crossline | 12 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:55 | 1.40 cm | Crossline | 12 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:55 | 2.00 cm | Crossline | 12 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:56 | 2.50 cm | Crossline | 12 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:56 | 3.00 cm | Crossline | 12 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:57 | 4.00 cm | Crossline | 12 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:58 | 4.50 cm | Crossline | 12 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:58 | 5.00 cm | Crossline | 12 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 11:59 | 5.50 cm | Crossline | 12 MeV | 3.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 12:02 | 0.80 cm | Crossline | 9 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:03 | 1.00 cm | Crossline | 9 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:03 | 1.40 cm | Crossline | 9 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:04 | 2.00 cm | Crossline | 9 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:05 | 2.50 cm | Crossline | 9 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:05 | 3.00 cm | Crossline | 9 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:06 | 3.50 cm | Crossline | 9 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:07 | 4.00 cm | Crossline | 9 MeV | 3.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 12:10 | 0.80 cm | Crossline | 6 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:11 | 1.00 cm | Crossline | 6 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:11 | 1.40 cm | Crossline | 6 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:12 | 2.00 cm | Crossline | 6 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:13 | 2.50 cm | Crossline | 6 MeV | 3.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:13 | 3.00 cm | Crossline | 6 MeV | 3.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | FALSE | FALSE | TRUE
2017-08-28 | 12:19 | 0.50 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:20 | 1.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:21 | 2.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:22 | 2.50 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:24 | 3.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:25 | 4.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:26 | 5.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:27 | 6.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:28 | 7.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:29 | 8.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:30 | 9.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 12:33 | 0.50 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:35 | 1.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:36 | 2.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:37 | 2.50 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:38 | 3.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:39 | 4.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:40 | 4.50 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:41 | 5.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:42 | 5.50 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:43 | 6.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:44 | 6.50 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:45 | 7.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 12:49 | 0.50 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:50 | 1.00 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:51 | 1.40 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:52 | 2.00 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:53 | 2.50 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:54 | 3.00 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:55 | 4.00 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:56 | 4.50 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:58 | 5.00 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 12:59 | 5.50 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 13:02 | 0.80 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:03 | 1.00 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:04 | 1.40 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:05 | 2.00 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:06 | 2.50 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:08 | 3.00 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:09 | 3.50 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:10 | 4.00 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 13:15 | 0.80 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:17 | 1.00 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:18 | 1.40 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:19 | 2.00 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:20 | 2.50 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:21 | 3.00 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 100.00 cm | FALSE | FALSE | FALSE | TRUE | TRUE
2017-08-28 | 13:35 | 0.50 cm | Crossline | 20 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:36 | 1.00 cm | Crossline | 20 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:37 | 2.00 cm | Crossline | 20 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:38 | 2.50 cm | Crossline | 20 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:39 | 3.00 cm | Crossline | 20 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:40 | 4.00 cm | Crossline | 20 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:41 | 5.00 cm | Crossline | 20 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:42 | 6.00 cm | Crossline | 20 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:43 | 7.00 cm | Crossline | 20 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:44 | 8.00 cm | Crossline | 20 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:45 | 9.00 cm | Crossline | 20 MeV | 11.00 cm | 15x15 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 13:49 | 0.50 cm | Crossline | 16 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:50 | 1.00 cm | Crossline | 16 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:51 | 2.00 cm | Crossline | 16 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:52 | 2.50 cm | Crossline | 16 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:53 | 3.00 cm | Crossline | 16 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:54 | 4.00 cm | Crossline | 16 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:55 | 4.50 cm | Crossline | 16 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:56 | 5.00 cm | Crossline | 16 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:57 | 5.50 cm | Crossline | 16 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:58 | 6.00 cm | Crossline | 16 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 13:59 | 6.50 cm | Crossline | 16 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:00 | 7.00 cm | Crossline | 16 MeV | 11.00 cm | 15x15 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 14:04 | 0.50 cm | Crossline | 12 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:05 | 1.00 cm | Crossline | 12 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:06 | 1.40 cm | Crossline | 12 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:07 | 2.00 cm | Crossline | 12 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:08 | 2.50 cm | Crossline | 12 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:09 | 3.00 cm | Crossline | 12 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:10 | 4.00 cm | Crossline | 12 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:11 | 4.50 cm | Crossline | 12 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:12 | 5.00 cm | Crossline | 12 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:13 | 5.50 cm | Crossline | 12 MeV | 11.00 cm | 15x15 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 14:17 | 0.80 cm | Crossline | 9 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:18 | 1.00 cm | Crossline | 9 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:19 | 1.40 cm | Crossline | 9 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:20 | 2.00 cm | Crossline | 9 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:21 | 2.50 cm | Crossline | 9 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:22 | 3.00 cm | Crossline | 9 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:23 | 3.50 cm | Crossline | 9 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:24 | 4.00 cm | Crossline | 9 MeV | 11.00 cm | 15x15 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 14:29 | 0.80 cm | Crossline | 6 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:30 | 1.00 cm | Crossline | 6 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:31 | 1.40 cm | Crossline | 6 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:32 | 2.00 cm | Crossline | 6 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:33 | 2.50 cm | Crossline | 6 MeV | 11.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:34 | 3.00 cm | Crossline | 6 MeV | 11.00 cm | 15x15 | 100.00 cm | FALSE | FALSE | FALSE | FALSE | TRUE
2017-08-28 | 14:46 | 0.50 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:46 | 1.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:47 | 2.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:48 | 2.50 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:48 | 3.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:49 | 4.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:49 | 5.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:50 | 6.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:51 | 7.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:51 | 8.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:52 | 9.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 14:56 | 0.50 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:56 | 1.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:57 | 2.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:58 | 2.50 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:58 | 3.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 14:59 | 4.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:00 | 4.50 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:00 | 5.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:01 | 5.50 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:02 | 6.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:02 | 6.50 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:03 | 7.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 15:06 | 0.50 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:07 | 1.00 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:08 | 1.40 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:08 | 2.00 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:09 | 2.50 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:10 | 3.00 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:10 | 4.00 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:11 | 4.50 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:12 | 5.00 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:13 | 5.50 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 15:15 | 0.80 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:16 | 1.00 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:16 | 1.40 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:17 | 2.00 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:18 | 2.50 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:18 | 3.00 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:19 | 3.50 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:20 | 4.00 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 15:23 | 0.80 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:23 | 1.00 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:24 | 1.40 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:25 | 2.00 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:25 | 2.50 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:26 | 3.00 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | FALSE | TRUE | TRUE
2017-08-28 | 15:33 | 0.50 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:33 | 1.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:34 | 2.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:35 | 2.50 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:35 | 3.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:36 | 4.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:36 | 5.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:37 | 6.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:38 | 7.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:38 | 8.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:39 | 9.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 15:43 | 0.50 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:43 | 1.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:44 | 2.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:45 | 2.50 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:45 | 3.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:46 | 4.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:46 | 4.50 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:47 | 5.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:48 | 5.50 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:48 | 6.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:49 | 6.50 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:50 | 7.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 15:55 | 0.50 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:55 | 1.00 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:56 | 1.40 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:57 | 2.00 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:57 | 2.50 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:58 | 3.00 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:58 | 4.00 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 15:59 | 4.50 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:00 | 5.00 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:00 | 5.50 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 16:04 | 0.80 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:05 | 1.00 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:05 | 1.40 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:06 | 2.00 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:07 | 2.50 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:07 | 3.00 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:08 | 3.50 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:09 | 4.00 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 16:12 | 0.80 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:12 | 1.00 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:13 | 1.40 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:14 | 2.00 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:14 | 2.50 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:15 | 3.00 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | FALSE | TRUE | TRUE
2017-08-28 | 16:20 | 0.50 cm | Crossline | 20 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:21 | 1.00 cm | Crossline | 20 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:22 | 2.00 cm | Crossline | 20 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:22 | 2.50 cm | Crossline | 20 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:23 | 3.00 cm | Crossline | 20 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:23 | 4.00 cm | Crossline | 20 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:24 | 5.00 cm | Crossline | 20 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:25 | 6.00 cm | Crossline | 20 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:25 | 7.00 cm | Crossline | 20 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:26 | 8.00 cm | Crossline | 20 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:27 | 9.00 cm | Crossline | 20 MeV | 3.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 16:31 | 0.50 cm | Crossline | 16 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:31 | 1.00 cm | Crossline | 16 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:32 | 2.00 cm | Crossline | 16 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:33 | 2.50 cm | Crossline | 16 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:33 | 3.00 cm | Crossline | 16 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:34 | 4.00 cm | Crossline | 16 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:34 | 4.50 cm | Crossline | 16 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:35 | 5.00 cm | Crossline | 16 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:36 | 5.50 cm | Crossline | 16 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:36 | 6.00 cm | Crossline | 16 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:37 | 6.50 cm | Crossline | 16 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:38 | 7.00 cm | Crossline | 16 MeV | 3.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 16:42 | 0.50 cm | Crossline | 12 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:42 | 1.00 cm | Crossline | 12 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:43 | 1.40 cm | Crossline | 12 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:43 | 2.00 cm | Crossline | 12 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:44 | 2.50 cm | Crossline | 12 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:45 | 3.00 cm | Crossline | 12 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:45 | 4.00 cm | Crossline | 12 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:46 | 4.50 cm | Crossline | 12 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:47 | 5.00 cm | Crossline | 12 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:47 | 5.50 cm | Crossline | 12 MeV | 3.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 16:51 | 0.80 cm | Crossline | 9 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:52 | 1.00 cm | Crossline | 9 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:52 | 1.40 cm | Crossline | 9 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:53 | 2.00 cm | Crossline | 9 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:54 | 2.50 cm | Crossline | 9 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:54 | 3.00 cm | Crossline | 9 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:55 | 3.50 cm | Crossline | 9 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:56 | 4.00 cm | Crossline | 9 MeV | 3.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-28 | 16:58 | 0.80 cm | Crossline | 6 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 16:59 | 1.00 cm | Crossline | 6 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 17:00 | 1.40 cm | Crossline | 6 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 17:00 | 2.00 cm | Crossline | 6 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 17:01 | 2.50 cm | Crossline | 6 MeV | 3.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-28 | 17:01 | 3.00 cm | Crossline | 6 MeV | 3.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | FALSE | FALSE | FALSE
2017-08-29 | 10:26 | 0.50 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:27 | 1.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:29 | 2.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:30 | 2.50 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:31 | 3.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:33 | 4.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:34 | 5.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:36 | 6.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:37 | 7.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:38 | 8.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:40 | 9.00 cm | Crossline | 20 MeV | 25.00 cm | 25x25 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 10:44 | 0.50 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:45 | 1.00 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:47 | 2.00 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:48 | 2.50 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:49 | 3.00 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:51 | 4.00 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:52 | 4.50 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:54 | 5.00 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:55 | 5.50 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:56 | 6.00 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:58 | 6.50 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 10:59 | 7.00 cm | Crossline | 16 MeV | 25.00 cm | 25x25 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 11:03 | 0.50 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:05 | 1.00 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:06 | 1.40 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:08 | 2.00 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:09 | 2.50 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:10 | 3.00 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:12 | 4.00 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:13 | 4.50 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:15 | 5.00 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:16 | 5.50 cm | Crossline | 12 MeV | 25.00 cm | 25x25 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 11:20 | 0.80 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:21 | 1.00 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:23 | 1.40 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:24 | 2.00 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:25 | 2.50 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:27 | 3.00 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:28 | 3.50 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:30 | 4.00 cm | Crossline | 9 MeV | 25.00 cm | 25x25 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 11:33 | 0.80 cm | Crossline | 6 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:34 | 1.00 cm | Crossline | 6 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:36 | 1.40 cm | Crossline | 6 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:37 | 2.00 cm | Crossline | 6 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:39 | 2.50 cm | Crossline | 6 MeV | 25.00 cm | 25x25 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 11:40 | 3.00 cm | Crossline | 6 MeV | 25.00 cm | 25x25 | 110.00 cm | FALSE | FALSE | FALSE | FALSE | TRUE
2017-08-29 | 12:08 | 0.50 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:09 | 1.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:10 | 2.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:12 | 2.50 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:13 | 3.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:14 | 4.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:15 | 5.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:16 | 6.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:18 | 7.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:19 | 8.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:20 | 9.00 cm | Crossline | 20 MeV | 20.00 cm | 20x20 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 12:25 | 0.50 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:26 | 1.00 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:27 | 2.00 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:29 | 2.50 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:30 | 3.00 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:31 | 4.00 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:32 | 4.50 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:34 | 5.00 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:35 | 5.50 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:36 | 6.00 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:38 | 6.50 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:39 | 7.00 cm | Crossline | 16 MeV | 20.00 cm | 20x20 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 12:43 | 0.50 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:44 | 1.00 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:45 | 1.40 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:47 | 2.00 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:48 | 2.50 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:49 | 3.00 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:50 | 4.00 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:52 | 4.50 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:53 | 5.00 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:54 | 5.50 cm | Crossline | 12 MeV | 20.00 cm | 20x20 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 12:58 | 0.80 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 12:59 | 1.00 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:00 | 1.40 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:02 | 2.00 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:03 | 2.50 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:04 | 3.00 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:05 | 3.50 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:07 | 4.00 cm | Crossline | 9 MeV | 20.00 cm | 20x20 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 13:10 | 0.80 cm | Crossline | 6 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:11 | 1.00 cm | Crossline | 6 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:13 | 1.40 cm | Crossline | 6 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:14 | 2.00 cm | Crossline | 6 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:15 | 2.50 cm | Crossline | 6 MeV | 20.00 cm | 20x20 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:17 | 3.00 cm | Crossline | 6 MeV | 20.00 cm | 20x20 | 110.00 cm | FALSE | FALSE | FALSE | FALSE | TRUE
2017-08-29 | 13:25 | 0.50 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:26 | 1.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:27 | 2.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:28 | 2.50 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:29 | 3.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:30 | 4.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:32 | 5.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:33 | 6.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:34 | 7.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:35 | 8.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:36 | 9.00 cm | Crossline | 20 MeV | 15.00 cm | 15x15 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 13:42 | 0.50 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:43 | 1.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:44 | 2.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:45 | 2.50 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:46 | 3.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:47 | 4.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:49 | 4.50 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:50 | 5.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:51 | 5.50 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:52 | 6.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:53 | 6.50 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 13:55 | 7.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 14:00 | 0.50 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:01 | 1.00 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:03 | 1.40 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:04 | 2.00 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:06 | 2.50 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:07 | 3.00 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:08 | 4.00 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:09 | 4.50 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:10 | 5.00 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:12 | 5.50 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 14:15 | 0.80 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:17 | 1.00 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:18 | 1.40 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:19 | 2.00 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:20 | 2.50 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:21 | 3.00 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:22 | 3.50 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:24 | 4.00 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 14:27 | 0.80 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:28 | 1.00 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:29 | 1.40 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:31 | 2.00 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:32 | 2.50 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:33 | 3.00 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 110.00 cm | FALSE | FALSE | FALSE | FALSE | TRUE
2017-08-29 | 14:42 | 0.50 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:43 | 1.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:44 | 2.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:45 | 2.50 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:46 | 3.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:47 | 4.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:48 | 5.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:50 | 6.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:51 | 7.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:52 | 8.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:53 | 9.00 cm | Crossline | 20 MeV | 10.00 cm | 10x10 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 14:57 | 0.50 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:58 | 1.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 14:59 | 2.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:00 | 2.50 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:01 | 3.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:02 | 4.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:04 | 4.50 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:05 | 5.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:06 | 5.50 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:07 | 6.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:08 | 6.50 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:09 | 7.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 15:15 | 0.50 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:16 | 1.00 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:17 | 1.40 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:18 | 2.00 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:19 | 2.50 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:21 | 3.00 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:22 | 4.00 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:23 | 4.50 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:24 | 5.00 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:25 | 5.50 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 15:33 | 0.80 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:35 | 1.00 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:36 | 1.40 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:37 | 2.00 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:38 | 2.50 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:39 | 3.00 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:40 | 3.50 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:41 | 4.00 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 15:48 | 0.80 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:49 | 1.00 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:50 | 1.40 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:51 | 2.00 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:52 | 2.50 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 15:53 | 3.00 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 110.00 cm | FALSE | FALSE | FALSE | FALSE | TRUE
2017-08-29 | 16:02 | 0.50 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:03 | 1.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:04 | 2.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:05 | 2.50 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:05 | 3.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:06 | 4.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:07 | 5.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:08 | 6.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:09 | 7.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:10 | 8.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:11 | 9.00 cm | Crossline | 20 MeV | 6.00 cm | 6x6 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 16:14 | 0.50 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:15 | 1.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:16 | 2.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:17 | 2.50 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:18 | 3.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:19 | 4.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:20 | 4.50 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:21 | 5.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:22 | 5.50 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:23 | 6.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:24 | 6.50 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:25 | 7.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 16:31 | 0.50 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:32 | 1.00 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:33 | 1.40 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:34 | 2.00 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:35 | 2.50 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:35 | 3.00 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:36 | 4.00 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:37 | 4.50 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:38 | 5.00 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:39 | 5.50 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 16:43 | 0.80 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:44 | 1.00 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:45 | 1.40 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:46 | 2.00 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:47 | 2.50 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:48 | 3.00 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:49 | 3.50 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:50 | 4.00 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-08-29 | 16:55 | 0.80 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:56 | 1.00 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:57 | 1.40 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:58 | 2.00 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 16:58 | 2.50 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-08-29 | 17:00 | 3.00 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 110.00 cm | FALSE | TRUE | FALSE | FALSE | FALSE
2017-09-18 | 12:08 | 0.50 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 12:09 | 1.80 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 12:10 | 2.20 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 12:12 | 2.70 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-18 | 12:13 | 0.50 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 12:15 | 4.50 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-18 | 12:16 | 3.50 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | FALSE | TRUE | TRUE | TRUE
2017-09-18 | 12:18 | 6.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | FALSE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 12:18 | 6.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 12:20 | 7.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | FALSE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 12:20 | 7.00 cm | Crossline | 16 MeV | 15.00 cm | 15x15 | 100.00 cm | FALSE | FALSE | FALSE | FALSE | TRUE
2017-09-18 | 14:18 | 0.50 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 14:19 | 1.80 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 14:20 | 2.20 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 14:21 | 2.70 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-18 | 14:22 | 0.50 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 14:23 | 4.50 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-18 | 14:24 | 3.50 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | FALSE | TRUE | TRUE | TRUE
2017-09-18 | 14:26 | 6.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | FALSE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 14:26 | 6.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 14:28 | 7.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | FALSE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 14:28 | 7.00 cm | Crossline | 16 MeV | 10.00 cm | 10x10 | 100.00 cm | FALSE | FALSE | FALSE | FALSE | TRUE
2017-09-18 | 15:19 | 0.50 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 15:20 | 1.80 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 15:20 | 2.20 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 15:21 | 2.70 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-18 | 15:22 | 0.50 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 15:24 | 4.50 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-18 | 15:25 | 3.50 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | FALSE | TRUE | TRUE | TRUE
2017-09-18 | 15:26 | 6.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | FALSE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 15:26 | 6.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 15:28 | 7.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | FALSE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 15:28 | 7.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | FALSE | TRUE | TRUE
2017-09-18 | 15:38 | 0.50 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 15:38 | 1.80 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 15:39 | 2.20 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 15:40 | 2.70 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-18 | 15:41 | 0.50 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 15:42 | 4.50 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-18 | 15:43 | 3.50 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | FALSE | TRUE | TRUE | TRUE
2017-09-18 | 15:44 | 6.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | FALSE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 15:44 | 6.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 15:46 | 7.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | FALSE | TRUE | TRUE | TRUE | TRUE
2017-09-18 | 15:46 | 7.00 cm | Crossline | 16 MeV | 6.00 cm | 6x6 | 100.00 cm | FALSE | FALSE | FALSE | FALSE | FALSE
2017-09-19 | 11:14 | 0.50 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 11:16 | 1.80 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 11:17 | 2.20 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 11:18 | 2.70 cm | Crossline | 6 MeV | 15.00 cm | 15x15 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-19 | 11:19 | 0.50 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 11:20 | 4.50 cm | Crossline | 9 MeV | 15.00 cm | 15x15 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-19 | 11:21 | 3.50 cm | Crossline | 12 MeV | 15.00 cm | 15x15 | 110.00 cm | FALSE | FALSE | FALSE | FALSE | TRUE
2017-09-19 | 12:03 | 0.50 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 12:04 | 1.80 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 12:05 | 2.20 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 12:06 | 2.70 cm | Crossline | 6 MeV | 10.00 cm | 10x10 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-19 | 12:07 | 0.50 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 12:09 | 4.50 cm | Crossline | 9 MeV | 10.00 cm | 10x10 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-19 | 12:10 | 3.50 cm | Crossline | 12 MeV | 10.00 cm | 10x10 | 110.00 cm | FALSE | FALSE | FALSE | FALSE | TRUE
2017-09-19 | 13:54 | 0.50 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 13:54 | 1.80 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 13:55 | 2.20 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 13:56 | 2.70 cm | Crossline | 6 MeV | 6.00 cm | 6x6 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-19 | 13:57 | 0.50 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 13:58 | 4.50 cm | Crossline | 9 MeV | 6.00 cm | 6x6 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-19 | 13:59 | 3.50 cm | Crossline | 12 MeV | 6.00 cm | 6x6 | 110.00 cm | FALSE | FALSE | FALSE | TRUE | TRUE
2017-09-19 | 14:55 | 0.50 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 14:55 | 0.80 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 14:56 | 1.00 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 14:57 | 1.40 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 14:57 | 1.80 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 14:58 | 2.00 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 14:59 | 2.20 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 14:59 | 2.50 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:00 | 2.70 cm | Crossline | 6 MeV | 4.00 cm | 6x6 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-19 | 15:01 | 0.50 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:03 | 1.00 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:03 | 1.50 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:04 | 2.00 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:05 | 2.50 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:05 | 3.00 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:06 | 3.50 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:07 | 4.00 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:07 | 4.50 cm | Crossline | 9 MeV | 4.00 cm | 6x6 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-19 | 15:08 | 0.50 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:10 | 1.00 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:11 | 1.50 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:11 | 2.00 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:12 | 2.50 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:12 | 3.00 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:13 | 3.50 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:14 | 4.00 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:14 | 4.50 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:15 | 5.00 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:16 | 5.50 cm | Crossline | 12 MeV | 4.00 cm | 6x6 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-19 | 15:16 | 0.50 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | FALSE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:16 | 0.50 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:19 | 1.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:19 | 2.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | FALSE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:19 | 1.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:19 | 2.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:20 | 3.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | FALSE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:20 | 3.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:21 | 4.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | FALSE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:21 | 4.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:22 | 5.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:22 | 6.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | FALSE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:22 | 5.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:22 | 6.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:23 | 7.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | FALSE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:23 | 7.00 cm | Crossline | 16 MeV | 4.00 cm | 6x6 | 110.00 cm | FALSE | FALSE | TRUE | TRUE | TRUE
2017-09-19 | 15:24 | 0.50 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:25 | 1.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:26 | 2.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:26 | 3.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:27 | 4.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:28 | 5.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:28 | 6.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:29 | 7.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:30 | 8.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 110.00 cm | TRUE | TRUE | TRUE | TRUE | TRUE
2017-09-19 | 15:30 | 9.00 cm | Crossline | 20 MeV | 4.00 cm | 6x6 | 110.00 cm | FALSE | FALSE | FALSE | FALSE | FALSE

## Duplicate Check
[10 empty] | SSD | Energy | Applicator | Field size inline | Start position [De] | Date | Time | Count
Scan type | Crossline | Scan type | Crossline | [6 empty] | 100.00 cm | 6 MeV | 6x6 | 3.00 cm | 0.80 cm | 2017-08-28 | 16:58 | 1
[10 empty] | 100.00 cm | 6 MeV | 6x6 | 3.00 cm | 1.00 cm | 2017-08-28 | 16:59 | 1
Row Labels | Count of Radiation type | SSD | Energy | Applicator | Field size inline | Start position [De] | Date | Time | Count of Radiation type | 100.00 cm | 6 MeV | 6x6 | 3.00 cm | 1.40 cm | 2017-08-28 | 17:00 | 1
100.00 cm |  | 100.00 cm | 6 MeV | 6x6 | 3.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 3.00 cm | 2.00 cm | 2017-08-28 | 17:00 | 1
12 MeV |  | 100.00 cm | 6 MeV | 6x6 | 3.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 3.00 cm | 2.50 cm | 2017-08-28 | 17:01 | 1
10x10 |  | 100.00 cm | 6 MeV | 6x6 | 3.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 3.00 cm | 3.00 cm | 2017-08-28 | 17:01 | 1
10.00 cm |  | 100.00 cm | 6 MeV | 6x6 | 3.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 0.50 cm | 2017-09-18 | 15:19 | 1
0.00 cm | 1 | 100.00 cm | 6 MeV | 6x6 | 3.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 0.80 cm | 2017-08-28 | 16:12 | 1
2017-08-24 |  | 100.00 cm | 6 MeV | 6x6 | 3.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 1.00 cm | 2017-08-28 | 16:12 | 1
16:22 | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 1.40 cm | 2017-08-28 | 16:13 | 1
0.50 cm | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 1.80 cm | 2017-09-18 | 15:20 | 1
2017-08-24 |  | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.00 cm | 2017-08-28 | 16:14 | 1
16:23 | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.20 cm | 2017-09-18 | 15:20 | 1
1.00 cm | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 1.80 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.50 cm | 2017-08-28 | 16:14 | 1
2017-08-24 |  | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.70 cm | 2017-09-18 | 15:21 | 1
16:23 | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.20 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 3.00 cm | 2017-08-28 | 16:15 | 1
1.50 cm | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 0.50 cm | 2017-09-18 | 15:38 | 1
2017-08-24 |  | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.70 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 0.80 cm | 2017-08-28 | 15:23 | 1
16:24 | 1 | 100.00 cm | 6 MeV | 6x6 | 4.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 1.00 cm | 2017-08-28 | 15:23 | 1
2.00 cm | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 1.40 cm | 2017-08-28 | 15:24 | 1
2017-08-24 |  | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 1.80 cm | 2017-09-18 | 15:38 | 1
16:25 | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.00 cm | 2017-08-28 | 15:25 | 1
2.50 cm | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.20 cm | 2017-09-18 | 15:39 | 1
2017-08-24 |  | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 1.80 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.50 cm | 2017-08-28 | 15:25 | 1
16:26 | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.70 cm | 2017-09-18 | 15:40 | 1
3.00 cm | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.20 cm |  |  | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 3.00 cm | 2017-08-28 | 15:26 | 1
2017-08-24 |  | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 3.00 cm | 0.80 cm | 2017-08-28 | 12:10 | 1
16:27 | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.70 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 3.00 cm | 1.00 cm | 2017-08-28 | 12:11 | 1
3.50 cm | 1 | 100.00 cm | 6 MeV | 6x6 | 6.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 3.00 cm | 1.40 cm | 2017-08-28 | 12:11 | 1
2017-09-18 |  | 100.00 cm | 6 MeV | 10x10 | 3.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 3.00 cm | 2.00 cm | 2017-08-28 | 12:12 | 1
14:24 | 1 | 100.00 cm | 6 MeV | 10x10 | 3.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 3.00 cm | 2.50 cm | 2017-08-28 | 12:13 | 1
4.00 cm | 1 | 100.00 cm | 6 MeV | 10x10 | 3.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 3.00 cm | 3.00 cm | 2017-08-28 | 12:13 | 1
2017-08-24 |  | 100.00 cm | 6 MeV | 10x10 | 3.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 7.00 cm | 0.80 cm | 2017-08-25 | 16:20 | 1
16:28 | 1 | 100.00 cm | 6 MeV | 10x10 | 3.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 7.00 cm | 1.00 cm | 2017-08-25 | 16:21 | 1
4.50 cm | 1 | 100.00 cm | 6 MeV | 10x10 | 3.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 7.00 cm | 1.40 cm | 2017-08-25 | 16:22 | 1
2017-08-24 |  | 100.00 cm | 6 MeV | 10x10 | 7.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 7.00 cm | 2.00 cm | 2017-08-25 | 16:23 | 1
16:29 | 1 | 100.00 cm | 6 MeV | 10x10 | 7.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 7.00 cm | 2.50 cm | 2017-08-25 | 16:24 | 1
5.00 cm | 1 | 100.00 cm | 6 MeV | 10x10 | 7.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 7.00 cm | 3.00 cm | 2017-08-25 | 16:24 | 1
2017-08-24 |  | 100.00 cm | 6 MeV | 10x10 | 7.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 0.00 cm | 2017-08-24 | 15:16 | 1
16:30 | 1 | 100.00 cm | 6 MeV | 10x10 | 7.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-08-24 | 15:17 | 1
5.50 cm | 1 | 100.00 cm | 6 MeV | 10x10 | 7.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-09-18 | 14:18 | 1
2017-08-24 |  | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 0.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 0.80 cm | 2017-08-24 | 15:17 | 1
16:31 | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 1.00 cm | 2017-08-24 | 15:18 | 1
6.00 cm | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 1.40 cm | 2017-08-24 | 15:19 | 1
2017-08-24 |  | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 1.80 cm | 2017-09-18 | 14:19 | 1
16:32 | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.00 cm | 2017-08-24 | 15:20 | 1
3.00 cm |  | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 1.80 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.20 cm | 2017-09-18 | 14:20 | 1
0.50 cm | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.50 cm | 2017-08-24 | 15:20 | 1
2017-08-28 |  | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.20 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.70 cm | 2017-09-18 | 14:21 | 1
11:53 | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 3.00 cm | 2017-08-24 | 15:21 | 1
1.00 cm | 1 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.70 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 11.00 cm | 0.80 cm | 2017-08-28 | 14:29 | 1
2017-08-28 |  | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 11.00 cm | 1.00 cm | 2017-08-28 | 14:30 | 1
11:54 | 1 | 100.00 cm | 6 MeV | 15x15 | 11.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 11.00 cm | 1.40 cm | 2017-08-28 | 14:31 | 1
1.40 cm | 1 | 100.00 cm | 6 MeV | 15x15 | 11.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 11.00 cm | 2.00 cm | 2017-08-28 | 14:32 | 1
2017-08-28 |  | 100.00 cm | 6 MeV | 15x15 | 11.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 11.00 cm | 2.50 cm | 2017-08-28 | 14:33 | 1
11:55 | 1 | 100.00 cm | 6 MeV | 15x15 | 11.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 11.00 cm | 3.00 cm | 2017-08-28 | 14:34 | 1
2.00 cm | 1 | 100.00 cm | 6 MeV | 15x15 | 11.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 0.50 cm | 2017-09-18 | 12:08 | 1
2017-08-28 |  | 100.00 cm | 6 MeV | 15x15 | 11.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 0.80 cm | 2017-08-28 | 13:15 | 1
11:55 | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 1.00 cm | 2017-08-28 | 13:17 | 1
2.50 cm | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 1.40 cm | 2017-08-28 | 13:18 | 1
2017-08-28 |  | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 1.80 cm | 2017-09-18 | 12:09 | 1
11:56 | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.00 cm | 2017-08-28 | 13:19 | 1
3.00 cm | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 1.80 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.20 cm | 2017-09-18 | 12:10 | 1
2017-08-28 |  | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.50 cm | 2017-08-28 | 13:20 | 1
11:56 | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.20 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.70 cm | 2017-09-18 | 12:12 | 1
4.00 cm | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 3.00 cm | 2017-08-28 | 13:21 | 1
2017-08-28 |  | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.70 cm |  |  | 1 | 100.00 cm | 6 MeV | 20x20 | 20.00 cm | 0.80 cm | 2017-08-28 | 11:14 | 1
11:57 | 1 | 100.00 cm | 6 MeV | 15x15 | 15.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 20x20 | 20.00 cm | 1.00 cm | 2017-08-28 | 11:15 | 1
4.50 cm | 1 | 100.00 cm | 6 MeV | 20x20 | 20.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 6 MeV | 20x20 | 20.00 cm | 1.40 cm | 2017-08-28 | 11:17 | 1
2017-08-28 |  | 100.00 cm | 6 MeV | 20x20 | 20.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 20x20 | 20.00 cm | 2.00 cm | 2017-08-28 | 11:18 | 1
11:58 | 1 | 100.00 cm | 6 MeV | 20x20 | 20.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 6 MeV | 20x20 | 20.00 cm | 2.50 cm | 2017-08-28 | 11:19 | 1
5.00 cm | 1 | 100.00 cm | 6 MeV | 20x20 | 20.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 20x20 | 20.00 cm | 3.00 cm | 2017-08-28 | 11:20 | 1
2017-08-28 |  | 100.00 cm | 6 MeV | 20x20 | 20.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 6 MeV | 25x25 | 21.00 cm | 0.80 cm | 2017-08-25 | 14:05 | 1
11:58 | 1 | 100.00 cm | 6 MeV | 20x20 | 20.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 25x25 | 21.00 cm | 1.00 cm | 2017-08-25 | 14:06 | 1
5.50 cm | 1 | 100.00 cm | 6 MeV | 25x25 | 21.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 6 MeV | 25x25 | 21.00 cm | 1.40 cm | 2017-08-25 | 14:07 | 1
2017-08-28 |  | 100.00 cm | 6 MeV | 25x25 | 21.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 25x25 | 21.00 cm | 2.00 cm | 2017-08-25 | 14:08 | 1
11:59 | 1 | 100.00 cm | 6 MeV | 25x25 | 21.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 6 MeV | 25x25 | 21.00 cm | 2.50 cm | 2017-08-25 | 14:10 | 1
7.00 cm |  | 100.00 cm | 6 MeV | 25x25 | 21.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 25x25 | 21.00 cm | 3.00 cm | 2017-08-25 | 14:11 | 1
0.50 cm | 1 | 100.00 cm | 6 MeV | 25x25 | 21.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 6 MeV | 25x25 | 25.00 cm | 0.80 cm | 2017-08-25 | 12:24 | 1
2017-08-25 |  | 100.00 cm | 6 MeV | 25x25 | 21.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 25x25 | 25.00 cm | 1.00 cm | 2017-08-25 | 12:25 | 1
15:47 | 1 | 100.00 cm | 6 MeV | 25x25 | 25.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 6 MeV | 25x25 | 25.00 cm | 1.40 cm | 2017-08-25 | 12:27 | 1
1.00 cm | 1 | 100.00 cm | 6 MeV | 25x25 | 25.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 25x25 | 25.00 cm | 2.00 cm | 2017-08-25 | 12:28 | 1
2017-08-25 |  | 100.00 cm | 6 MeV | 25x25 | 25.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 6 MeV | 25x25 | 25.00 cm | 2.50 cm | 2017-08-25 | 12:29 | 1
15:47 | 1 | 100.00 cm | 6 MeV | 25x25 | 25.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 6 MeV | 25x25 | 25.00 cm | 3.00 cm | 2017-08-25 | 12:31 | 1
1.40 cm | 1 | 100.00 cm | 6 MeV | 25x25 | 25.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 0.80 cm | 2017-08-28 | 16:51 | 1
2017-08-25 |  | 100.00 cm | 6 MeV | 25x25 | 25.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 1.00 cm | 2017-08-28 | 16:52 | 1
15:48 | 1 | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 1.40 cm | 2017-08-28 | 16:52 | 1
2.00 cm | 1 | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 2.00 cm | 2017-08-28 | 16:53 | 1
2017-08-25 |  | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 2.50 cm | 2017-08-28 | 16:54 | 1
15:49 | 1 | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 3.00 cm | 2017-08-28 | 16:54 | 1
2.50 cm | 1 | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 3.50 cm | 2017-08-28 | 16:55 | 1
2017-08-25 |  | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 4.00 cm | 2017-08-28 | 16:56 | 1
15:50 | 1 | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 0.50 cm | 2017-09-18 | 15:22 | 1
3.00 cm | 1 | 100.00 cm | 9 MeV | 6x6 | 3.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 0.80 cm | 2017-08-28 | 16:04 | 1
2017-08-25 |  | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 1.00 cm | 2017-08-28 | 16:05 | 1
15:51 | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 1.40 cm | 2017-08-28 | 16:05 | 1
4.00 cm | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 2.00 cm | 2017-08-28 | 16:06 | 1
2017-08-25 |  | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 2.50 cm | 2017-08-28 | 16:07 | 1
15:51 | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 3.00 cm | 2017-08-28 | 16:07 | 1
4.50 cm | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 3.50 cm | 2017-08-28 | 16:08 | 1
2017-08-25 |  | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 4.00 cm | 2017-08-28 | 16:09 | 1
15:52 | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 4.50 cm | 2017-09-18 | 15:24 | 1
5.00 cm | 1 | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 0.50 cm | 2017-09-18 | 15:41 | 1
2017-08-25 |  | 100.00 cm | 9 MeV | 6x6 | 4.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 0.80 cm | 2017-08-28 | 15:15 | 1
15:53 | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 1.00 cm | 2017-08-28 | 15:16 | 1
5.50 cm | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 1.40 cm | 2017-08-28 | 15:16 | 1
2017-08-25 |  | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 2.00 cm | 2017-08-28 | 15:17 | 1
15:54 | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 2.50 cm | 2017-08-28 | 15:18 | 1
15x15 |  | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 3.00 cm | 2017-08-28 | 15:18 | 1
11.00 cm |  | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 3.50 cm | 2017-08-28 | 15:19 | 1
0.50 cm | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 4.00 cm | 2017-08-28 | 15:20 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 4.50 cm | 2017-09-18 | 15:42 | 1
14:04 | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 0.80 cm | 2017-08-28 | 12:02 | 1
1.00 cm | 1 | 100.00 cm | 9 MeV | 6x6 | 6.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 1.00 cm | 2017-08-28 | 12:03 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 1.40 cm | 2017-08-28 | 12:03 | 1
14:05 | 1 | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 2.00 cm | 2017-08-28 | 12:04 | 1
1.40 cm | 1 | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 2.50 cm | 2017-08-28 | 12:05 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 3.00 cm | 2017-08-28 | 12:05 | 1
14:06 | 1 | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 3.50 cm | 2017-08-28 | 12:06 | 1
2.00 cm | 1 | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 4.00 cm | 2017-08-28 | 12:07 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 0.80 cm | 2017-08-25 | 16:03 | 1
14:07 | 1 | 100.00 cm | 9 MeV | 10x10 | 3.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 1.00 cm | 2017-08-25 | 16:03 | 1
2.50 cm | 1 | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 1.40 cm | 2017-08-25 | 16:04 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 2.00 cm | 2017-08-25 | 16:05 | 1
14:08 | 1 | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 2.50 cm | 2017-08-25 | 16:06 | 1
3.00 cm | 1 | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 3.00 cm | 2017-08-25 | 16:07 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 3.50 cm | 2017-08-25 | 16:07 | 1
14:09 | 1 | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 4.00 cm | 2017-08-25 | 16:08 | 1
4.00 cm | 1 | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 0.00 cm | 2017-08-24 | 15:47 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 10x10 | 7.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-08-24 | 15:48 | 1
14:10 | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 0.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-09-18 | 14:22 | 1
4.50 cm | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 0.80 cm | 2017-08-24 | 15:49 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 1.00 cm | 2017-08-24 | 15:50 | 1
14:11 | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 1.40 cm | 2017-08-24 | 15:51 | 1
5.00 cm | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 2.00 cm | 2017-08-24 | 15:52 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 2.50 cm | 2017-08-24 | 15:53 | 1
14:12 | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 3.00 cm | 2017-08-24 | 15:54 | 1
5.50 cm | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 3.50 cm | 2017-08-24 | 15:55 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 4.00 cm | 2017-08-24 | 15:56 | 1
14:13 | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 4.50 cm | 2017-08-24 | 15:57 | 1
15.00 cm |  | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 4.50 cm | 2017-09-18 | 14:23 | 1
0.50 cm | 1 | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 0.80 cm | 2017-08-28 | 14:17 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 1.00 cm | 2017-08-28 | 14:18 | 1
12:49 | 1 | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 1.40 cm | 2017-08-28 | 14:19 | 1
1.00 cm | 1 | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 2.00 cm | 2017-08-28 | 14:20 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 2.50 cm | 2017-08-28 | 14:21 | 1
12:50 | 1 | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 3.00 cm | 2017-08-28 | 14:22 | 1
1.40 cm | 1 | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 3.50 cm | 2017-08-28 | 14:23 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 11.00 cm | 4.00 cm | 2017-08-28 | 14:24 | 1
12:51 | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 0.50 cm | 2017-09-18 | 12:13 | 1
2.00 cm | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 0.80 cm | 2017-08-28 | 13:02 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 1.00 cm | 2017-08-28 | 13:03 | 1
12:52 | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 1.40 cm | 2017-08-28 | 13:04 | 1
2.50 cm | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 2.00 cm | 2017-08-28 | 13:05 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 2.50 cm | 2017-08-28 | 13:06 | 1
12:53 | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 3.00 cm | 2017-08-28 | 13:08 | 1
3.00 cm | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 3.50 cm | 2017-08-28 | 13:09 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 4.00 cm | 2017-08-28 | 13:10 | 1
12:54 | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 15x15 | 15.00 cm | 4.50 cm | 2017-09-18 | 12:15 | 1
3.50 cm | 1 | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 0.80 cm | 2017-08-28 | 10:58 | 1
2017-09-18 |  | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 1.00 cm | 2017-08-28 | 10:59 | 1
12:16 | 1 | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 1.40 cm | 2017-08-28 | 11:00 | 1
4.00 cm | 1 | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 2.00 cm | 2017-08-28 | 11:01 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 2.50 cm | 2017-08-28 | 11:02 | 1
12:55 | 1 | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 3.00 cm | 2017-08-28 | 11:04 | 1
4.50 cm | 1 | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 3.50 cm | 2017-08-28 | 11:05 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 20x20 | 20.00 cm | 4.00 cm | 2017-08-28 | 11:06 | 1
12:56 | 1 | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 0.80 cm | 2017-08-25 | 13:53 | 1
5.00 cm | 1 | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 1.00 cm | 2017-08-25 | 13:54 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 1.40 cm | 2017-08-25 | 13:55 | 1
12:58 | 1 | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 2.00 cm | 2017-08-25 | 13:57 | 1
5.50 cm | 1 | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 2.50 cm | 2017-08-25 | 13:58 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 3.00 cm | 2017-08-25 | 13:59 | 1
12:59 | 1 | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 3.50 cm | 2017-08-25 | 14:00 | 1
20x20 |  | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 21.00 cm | 4.00 cm | 2017-08-25 | 14:01 | 1
20.00 cm |  | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 0.80 cm | 2017-08-25 | 12:10 | 1
0.50 cm | 1 | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 1.00 cm | 2017-08-25 | 12:11 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 1.40 cm | 2017-08-25 | 12:13 | 1
10:44 | 1 | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 2.00 cm | 2017-08-25 | 12:14 | 1
1.00 cm | 1 | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 2.50 cm | 2017-08-25 | 12:15 | 1
2017-08-28 |  | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 3.00 cm | 2017-08-25 | 12:17 | 1
10:45 | 1 | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 3.50 cm | 2017-08-25 | 12:18 | 1
1.40 cm | 1 | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 9 MeV | 25x25 | 25.00 cm | 4.00 cm | 2017-08-25 | 12:20 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 0.50 cm | 2017-08-28 | 16:42 | 1
10:46 | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 1.00 cm | 2017-08-28 | 16:42 | 1
2.00 cm | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 1.40 cm | 2017-08-28 | 16:43 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 2.00 cm | 2017-08-28 | 16:43 | 1
10:47 | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 2.50 cm | 2017-08-28 | 16:44 | 1
2.50 cm | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 3.00 cm | 2017-08-28 | 16:45 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 4.00 cm | 2017-08-28 | 16:45 | 1
10:48 | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 4.50 cm | 2017-08-28 | 16:46 | 1
3.00 cm | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 5.00 cm | 2017-08-28 | 16:47 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 3.00 cm | 5.50 cm | 2017-08-28 | 16:47 | 1
10:50 | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 0.50 cm | 2017-08-28 | 15:55 | 1
4.00 cm | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 1.00 cm | 2017-08-28 | 15:55 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 1.40 cm | 2017-08-28 | 15:56 | 1
10:51 | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 2.00 cm | 2017-08-28 | 15:57 | 1
4.50 cm | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 2.50 cm | 2017-08-28 | 15:57 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 3.00 cm | 2017-08-28 | 15:58 | 1
10:52 | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 3.50 cm | 2017-09-18 | 15:25 | 1
5.00 cm | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 4.00 cm | 2017-08-28 | 15:58 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 4.50 cm | 2017-08-28 | 15:59 | 1
10:53 | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 5.00 cm | 2017-08-28 | 16:00 | 1
5.50 cm | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 4.00 cm | 5.50 cm | 2017-08-28 | 16:00 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 0.50 cm | 2017-08-28 | 15:06 | 1
10:54 | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 1.00 cm | 2017-08-28 | 15:07 | 1
25x25 |  | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 1.40 cm | 2017-08-28 | 15:08 | 1
21.00 cm |  | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 2.00 cm | 2017-08-28 | 15:08 | 1
0.50 cm | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 2.50 cm | 2017-08-28 | 15:09 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 3.00 cm | 2017-08-28 | 15:10 | 1
13:36 | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 3.50 cm | 2017-09-18 | 15:43 | 1
1.00 cm | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 4.00 cm | 2017-08-28 | 15:10 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 4.50 cm | 2017-08-28 | 15:11 | 1
13:38 | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 5.00 cm | 2017-08-28 | 15:12 | 1
1.40 cm | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 6x6 | 6.00 cm | 5.50 cm | 2017-08-28 | 15:13 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 0.50 cm | 2017-08-28 | 11:53 | 1
13:39 | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 1.00 cm | 2017-08-28 | 11:54 | 1
2.00 cm | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 1.40 cm | 2017-08-28 | 11:55 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 2.00 cm | 2017-08-28 | 11:55 | 1
13:40 | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 2.50 cm | 2017-08-28 | 11:56 | 1
2.50 cm | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 3.00 cm | 2017-08-28 | 11:56 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 4.00 cm | 2017-08-28 | 11:57 | 1
13:41 | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 4.50 cm | 2017-08-28 | 11:58 | 1
3.00 cm | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 5.00 cm | 2017-08-28 | 11:58 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 3.00 cm | 5.50 cm | 2017-08-28 | 11:59 | 1
13:42 | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 0.50 cm | 2017-08-25 | 15:47 | 1
4.00 cm | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 1.00 cm | 2017-08-25 | 15:47 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 1.40 cm | 2017-08-25 | 15:48 | 1
13:43 | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 2.00 cm | 2017-08-25 | 15:49 | 1
4.50 cm | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 2.50 cm | 2017-08-25 | 15:50 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 3.00 cm | 2017-08-25 | 15:51 | 1
13:45 | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 4.00 cm | 2017-08-25 | 15:51 | 1
5.00 cm | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 4.50 cm | 2017-08-25 | 15:52 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 5.00 cm | 2017-08-25 | 15:53 | 1
13:46 | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 7.00 cm | 5.50 cm | 2017-08-25 | 15:54 | 1
5.50 cm | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 0.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 0.00 cm | 2017-08-24 | 16:22 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-08-24 | 16:23 | 1
13:47 | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 1.00 cm | 2017-08-24 | 16:23 | 1
25.00 cm |  | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 1.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 1.50 cm | 2017-08-24 | 16:24 | 1
0.50 cm | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 2.00 cm | 2017-08-24 | 16:25 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 2.50 cm | 2017-08-24 | 16:26 | 1
11:53 | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 3.00 cm | 2017-08-24 | 16:27 | 1
1.00 cm | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 3.50 cm | 2017-08-24 | 16:28 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 3.50 cm | 2017-09-18 | 14:24 | 1
11:55 | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 4.00 cm | 2017-08-24 | 16:28 | 1
1.40 cm | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 4.50 cm | 2017-08-24 | 16:29 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 5.00 cm | 2017-08-24 | 16:30 | 1
11:56 | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 5.50 cm | 2017-08-24 | 16:31 | 1
2.00 cm | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 6.00 cm | 2017-08-24 | 16:32 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 0.50 cm | 2017-08-28 | 14:04 | 1
11:57 | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 1.00 cm | 2017-08-28 | 14:05 | 1
2.50 cm | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 1.40 cm | 2017-08-28 | 14:06 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 2.00 cm | 2017-08-28 | 14:07 | 1
11:58 | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 2.50 cm | 2017-08-28 | 14:08 | 1
3.00 cm | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 3.00 cm | 2017-08-28 | 14:09 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 4.00 cm | 2017-08-28 | 14:10 | 1
12:00 | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 4.50 cm | 2017-08-28 | 14:11 | 1
4.00 cm | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 5.00 cm | 2017-08-28 | 14:12 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 11.00 cm | 5.50 cm | 2017-08-28 | 14:13 | 1
12:01 | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 0.50 cm | 2017-08-28 | 12:49 | 1
4.50 cm | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 1.00 cm | 2017-08-28 | 12:50 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 1.40 cm | 2017-08-28 | 12:51 | 1
12:02 | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 2.00 cm | 2017-08-28 | 12:52 | 1
5.00 cm | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 2.50 cm | 2017-08-28 | 12:53 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 3.00 cm | 2017-08-28 | 12:54 | 1
12:04 | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 3.50 cm | 2017-09-18 | 12:16 | 1
5.50 cm | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 4.00 cm | 2017-08-28 | 12:55 | 1
2017-08-25 |  | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 4.50 cm | 2017-08-28 | 12:56 | 1
12:05 | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 5.00 cm | 2017-08-28 | 12:58 | 1
6x6 |  | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 15x15 | 15.00 cm | 5.50 cm | 2017-08-28 | 12:59 | 1
3.00 cm |  | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 0.50 cm | 2017-08-28 | 10:44 | 1
0.50 cm | 1 | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 1.00 cm | 2017-08-28 | 10:45 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 1.40 cm | 2017-08-28 | 10:46 | 1
16:42 | 1 | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 2.00 cm | 2017-08-28 | 10:47 | 1
1.00 cm | 1 | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 2.50 cm | 2017-08-28 | 10:48 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 3.00 cm | 2017-08-28 | 10:50 | 1
16:42 | 1 | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 4.00 cm | 2017-08-28 | 10:51 | 1
1.40 cm | 1 | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 4.50 cm | 2017-08-28 | 10:52 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 5.00 cm | 2017-08-28 | 10:53 | 1
16:43 | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 20x20 | 20.00 cm | 5.50 cm | 2017-08-28 | 10:54 | 1
2.00 cm | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 0.50 cm | 2017-08-25 | 13:36 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 1.00 cm | 2017-08-25 | 13:38 | 1
16:43 | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 1.40 cm | 2017-08-25 | 13:39 | 1
2.50 cm | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 2.00 cm | 2017-08-25 | 13:40 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 2.50 cm | 2017-08-25 | 13:41 | 1
16:44 | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 3.00 cm | 2017-08-25 | 13:42 | 1
3.00 cm | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 4.00 cm | 2017-08-25 | 13:43 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 4.50 cm | 2017-08-25 | 13:45 | 1
16:45 | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 5.00 cm | 2017-08-25 | 13:46 | 1
4.00 cm | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 21.00 cm | 5.50 cm | 2017-08-25 | 13:47 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 0.50 cm | 2017-08-25 | 11:53 | 1
16:45 | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 1.00 cm | 2017-08-25 | 11:55 | 1
4.50 cm | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 1.40 cm | 2017-08-25 | 11:56 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 2.00 cm | 2017-08-25 | 11:57 | 1
16:46 | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 2.50 cm | 2017-08-25 | 11:58 | 1
5.00 cm | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 3.00 cm | 2017-08-25 | 12:00 | 1
2017-08-28 |  | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 4.00 cm | 2017-08-25 | 12:01 | 1
16:47 | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 4.50 cm | 2017-08-25 | 12:02 | 1
5.50 cm | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 5.00 cm | 2017-08-25 | 12:04 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 12 MeV | 25x25 | 25.00 cm | 5.50 cm | 2017-08-25 | 12:05 | 1
16:47 | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 0.50 cm | 2017-08-28 | 16:31 | 1
4.00 cm |  | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 1.00 cm | 2017-08-28 | 16:31 | 1
0.50 cm | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 2.00 cm | 2017-08-28 | 16:32 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 2.50 cm | 2017-08-28 | 16:33 | 1
15:55 | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 3.00 cm | 2017-08-28 | 16:33 | 1
1.00 cm | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 4.00 cm | 2017-08-28 | 16:34 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 4.50 cm | 2017-08-28 | 16:34 | 1
15:55 | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 5.00 cm | 2017-08-28 | 16:35 | 1
1.40 cm | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 5.50 cm | 2017-08-28 | 16:36 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 6.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 6.00 cm | 2017-08-28 | 16:36 | 1
15:56 | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 6.50 cm | 2017-08-28 | 16:37 | 1
2.00 cm | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 3.00 cm | 7.00 cm | 2017-08-28 | 16:38 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 0.50 cm | 2017-08-28 | 15:43 | 1
15:57 | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 1.00 cm | 2017-08-28 | 15:43 | 1
2.50 cm | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 2.00 cm | 2017-08-28 | 15:44 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 2.50 cm | 2017-08-28 | 15:45 | 1
15:57 | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 3.00 cm | 2017-08-28 | 15:45 | 1
3.00 cm | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 4.00 cm | 2017-08-28 | 15:46 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 4.50 cm | 2017-08-28 | 15:46 | 1
15:58 | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 5.00 cm | 2017-08-28 | 15:47 | 1
3.50 cm | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 5.50 cm | 2017-08-28 | 15:48 | 1
2017-09-18 |  | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 6.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 6.00 cm | 2017-08-28 | 15:48 | 1
15:25 | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 6.00 cm | 2017-09-18 | 15:26 | 2
4.00 cm | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 6.50 cm | 2017-08-28 | 15:49 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 7.00 cm | 2017-08-28 | 15:50 | 1
15:58 | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 7.00 cm | 2017-09-18 | 15:28 | 2
4.50 cm | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 0.50 cm | 2017-08-28 | 14:56 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 1.00 cm | 2017-08-28 | 14:56 | 1
15:59 | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 2.00 cm | 2017-08-28 | 14:57 | 1
5.00 cm | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 2.50 cm | 2017-08-28 | 14:58 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 3.00 cm | 2017-08-28 | 14:58 | 1
16:00 | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 4.00 cm | 2017-08-28 | 14:59 | 1
5.50 cm | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 4.50 cm | 2017-08-28 | 15:00 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 6.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 5.00 cm | 2017-08-28 | 15:00 | 1
16:00 | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 5.50 cm | 2017-08-28 | 15:01 | 1
6.00 cm |  | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 6.00 cm | 2017-08-28 | 15:02 | 1
0.50 cm | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 6.00 cm | 2017-09-18 | 15:44 | 2
2017-08-28 |  | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 6.50 cm | 2017-08-28 | 15:02 | 1
15:06 | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 7.00 cm | 2017-08-28 | 15:03 | 1
1.00 cm | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 7.00 cm | 2017-09-18 | 15:46 | 2
2017-08-28 |  | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 0.50 cm | 2017-08-28 | 11:44 | 1
15:07 | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 1.00 cm | 2017-08-28 | 11:44 | 1
1.40 cm | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 2.00 cm | 2017-08-28 | 11:45 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 2.50 cm | 2017-08-28 | 11:45 | 1
15:08 | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 3.00 cm | 2017-08-28 | 11:46 | 1
2.00 cm | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 6.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 4.00 cm | 2017-08-28 | 11:47 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 4.50 cm | 2017-08-28 | 11:47 | 1
15:08 | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 5.00 cm | 2017-08-28 | 11:48 | 1
2.50 cm | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 5.50 cm | 2017-08-28 | 11:48 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 6.00 cm | 2017-08-28 | 11:49 | 1
15:09 | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 6.50 cm | 2017-08-28 | 11:50 | 1
3.00 cm | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 3.00 cm | 7.00 cm | 2017-08-28 | 11:50 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 0.50 cm | 2017-08-25 | 15:26 | 1
15:10 | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 1.00 cm | 2017-08-25 | 15:27 | 1
3.50 cm | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 2.00 cm | 2017-08-25 | 15:27 | 1
2017-09-18 |  | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 2.50 cm | 2017-08-25 | 15:28 | 1
15:43 | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 3.00 cm | 2017-08-25 | 15:29 | 1
4.00 cm | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 6.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 4.00 cm | 2017-08-25 | 15:30 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 4.50 cm | 2017-08-25 | 15:30 | 1
15:10 | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 0.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 5.00 cm | 2017-08-25 | 15:31 | 1
4.50 cm | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 5.50 cm | 2017-08-25 | 15:32 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 6.00 cm | 2017-08-25 | 15:33 | 1
15:11 | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 1.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 6.50 cm | 2017-08-25 | 15:34 | 1
5.00 cm | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 7.00 cm | 7.00 cm | 2017-08-25 | 15:34 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 0.00 cm | 2017-08-24 | 16:56 | 1
15:12 | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-08-24 | 16:57 | 1
5.50 cm | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 1.00 cm | 2017-08-24 | 16:58 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 1.50 cm | 2017-08-24 | 16:59 | 1
15:13 | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 2.00 cm | 2017-08-24 | 17:00 | 1
16 MeV |  | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 2.50 cm | 2017-08-24 | 17:01 | 1
10x10 |  | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 3.00 cm | 2017-08-24 | 17:02 | 1
10.00 cm |  | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 3.50 cm | 2017-08-24 | 17:02 | 1
0.00 cm | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 6.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 4.00 cm | 2017-08-24 | 17:03 | 1
2017-08-24 |  | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 4.50 cm | 2017-08-24 | 17:04 | 1
16:56 | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 5.00 cm | 2017-08-24 | 17:05 | 1
0.50 cm | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 5.50 cm | 2017-08-24 | 17:06 | 1
2017-08-24 |  | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 6.00 cm | 2017-08-24 | 17:07 | 1
16:57 | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 6.00 cm | 2017-09-18 | 14:26 | 2
1.00 cm | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 6.50 cm | 2017-08-24 | 17:08 | 1
2017-08-24 |  | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 7.00 cm | 2017-08-24 | 17:08 | 1
16:58 | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 7.00 cm | 2017-09-18 | 14:28 | 2
1.50 cm | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 0.50 cm | 2017-08-28 | 13:49 | 1
2017-08-24 |  | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 1.00 cm | 2017-08-28 | 13:50 | 1
16:59 | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 2.00 cm | 2017-08-28 | 13:51 | 1
2.00 cm | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 6.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 2.50 cm | 2017-08-28 | 13:52 | 1
2017-08-24 |  | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 3.00 cm | 2017-08-28 | 13:53 | 1
17:00 | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 4.00 cm | 2017-08-28 | 13:54 | 1
2.50 cm | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 4.50 cm | 2017-08-28 | 13:55 | 1
2017-08-24 |  | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 5.00 cm | 2017-08-28 | 13:56 | 1
17:01 | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 5.50 cm | 2017-08-28 | 13:57 | 1
3.00 cm | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 6.00 cm | 2017-08-28 | 13:58 | 1
2017-08-24 |  | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 6.50 cm | 2017-08-28 | 13:59 | 1
17:02 | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 11.00 cm | 7.00 cm | 2017-08-28 | 14:00 | 1
3.50 cm | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 0.50 cm | 2017-08-28 | 12:33 | 1
2017-08-24 |  | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 1.00 cm | 2017-08-28 | 12:35 | 1
17:02 | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 2.00 cm | 2017-08-28 | 12:36 | 1
4.00 cm | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 6.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 2.50 cm | 2017-08-28 | 12:37 | 1
2017-08-24 |  | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 3.00 cm | 2017-08-28 | 12:38 | 1
17:03 | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 4.00 cm | 2017-08-28 | 12:39 | 1
4.50 cm | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 4.50 cm | 2017-08-28 | 12:40 | 1
2017-08-24 |  | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 5.00 cm | 2017-08-28 | 12:41 | 1
17:04 | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 5.50 cm | 2017-08-28 | 12:42 | 1
5.00 cm | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 6.00 cm | 2017-08-28 | 12:43 | 1
2017-08-24 |  | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 6.00 cm | 2017-09-18 | 12:18 | 2
17:05 | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 6.50 cm | 2017-08-28 | 12:44 | 1
5.50 cm | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 7.00 cm | 2017-08-28 | 12:45 | 1
2017-08-24 |  | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 7.00 cm | 2017-09-18 | 12:20 | 2
17:06 | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 0.50 cm | 2017-08-28 | 10:27 | 1
6.00 cm | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 6.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 1.00 cm | 2017-08-28 | 10:28 | 1
2017-09-18 |  | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 2.00 cm | 2017-08-28 | 10:30 | 1
14:26 | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 2.50 cm | 2017-08-28 | 10:31 | 1
6.50 cm | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 3.00 cm | 2017-08-28 | 10:32 | 1
2017-08-24 |  | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 4.00 cm | 2017-08-28 | 10:33 | 1
17:08 | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 4.50 cm | 2017-08-28 | 10:34 | 1
7.00 cm | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 5.00 cm | 2017-08-28 | 10:35 | 1
2017-09-18 |  | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 5.50 cm | 2017-08-28 | 10:36 | 1
14:28 | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 6.00 cm | 2017-08-28 | 10:37 | 1
3.00 cm |  | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 6.50 cm | 2017-08-28 | 10:39 | 1
0.50 cm | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 20x20 | 20.00 cm | 7.00 cm | 2017-08-28 | 10:40 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 0.50 cm | 2017-08-25 | 13:01 | 1
11:44 | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 6.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 1.00 cm | 2017-08-25 | 13:02 | 1
1.00 cm | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 2.00 cm | 2017-08-25 | 13:03 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 2.50 cm | 2017-08-25 | 13:04 | 1
11:44 | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 3.00 cm | 2017-08-25 | 13:06 | 1
2.00 cm | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 4.00 cm | 2017-08-25 | 13:07 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 4.50 cm | 2017-08-25 | 13:08 | 1
11:45 | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 5.00 cm | 2017-08-25 | 13:09 | 1
2.50 cm | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 5.50 cm | 2017-08-25 | 13:10 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 6.00 cm | 2017-08-25 | 13:11 | 1
11:45 | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 6.50 cm | 2017-08-25 | 13:13 | 1
3.00 cm | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 21.00 cm | 7.00 cm | 2017-08-25 | 13:14 | 1
2017-08-28 |  | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 0.50 cm | 2017-08-25 | 11:32 | 1
11:46 | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 6.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 1.00 cm | 2017-08-25 | 11:33 | 1
4.00 cm | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 2.00 cm | 2017-08-25 | 11:34 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 2.50 cm | 2017-08-25 | 11:36 | 1
11:47 | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 3.00 cm | 2017-08-25 | 11:37 | 1
4.50 cm | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 4.00 cm | 2017-08-25 | 11:38 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 4.50 cm | 2017-08-25 | 11:40 | 1
11:47 | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 5.00 cm | 2017-08-25 | 11:41 | 1
5.00 cm | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 5.50 cm | 2017-08-25 | 11:42 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 6.00 cm | 2017-08-25 | 11:44 | 1
11:48 | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 6.50 cm | 2017-08-25 | 11:45 | 1
5.50 cm | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 16 MeV | 25x25 | 25.00 cm | 7.00 cm | 2017-08-25 | 11:46 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 8.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 0.50 cm | 2017-08-28 | 16:20 | 1
11:48 | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 9.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 1.00 cm | 2017-08-28 | 16:21 | 1
6.00 cm | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 2.00 cm | 2017-08-28 | 16:22 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 2.50 cm | 2017-08-28 | 16:22 | 1
11:49 | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 3.00 cm | 2017-08-28 | 16:23 | 1
6.50 cm | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 4.00 cm | 2017-08-28 | 16:23 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 5.00 cm | 2017-08-28 | 16:24 | 1
11:50 | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 6.00 cm | 2017-08-28 | 16:25 | 1
7.00 cm | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 7.00 cm | 2017-08-28 | 16:25 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 8.00 cm | 2017-08-28 | 16:26 | 1
11:50 | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 3.00 cm | 9.00 cm | 2017-08-28 | 16:27 | 1
7.00 cm |  | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 8.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 0.50 cm | 2017-08-28 | 15:33 | 1
0.50 cm | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 9.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 1.00 cm | 2017-08-28 | 15:33 | 1
2017-08-25 |  | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 2.00 cm | 2017-08-28 | 15:34 | 1
15:26 | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 2.50 cm | 2017-08-28 | 15:35 | 1
1.00 cm | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 3.00 cm | 2017-08-28 | 15:35 | 1
2017-08-25 |  | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 4.00 cm | 2017-08-28 | 15:36 | 1
15:27 | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 5.00 cm | 2017-08-28 | 15:36 | 1
2.00 cm | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 6.00 cm | 2017-08-28 | 15:37 | 1
2017-08-25 |  | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 7.00 cm | 2017-08-28 | 15:38 | 1
15:27 | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 8.00 cm | 2017-08-28 | 15:38 | 1
2.50 cm | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 4.00 cm | 9.00 cm | 2017-08-28 | 15:39 | 1
2017-08-25 |  | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 8.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 0.50 cm | 2017-08-28 | 14:46 | 1
15:28 | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 9.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 1.00 cm | 2017-08-28 | 14:46 | 1
3.00 cm | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 2.00 cm | 2017-08-28 | 14:47 | 1
2017-08-25 |  | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 2.50 cm | 2017-08-28 | 14:48 | 1
15:29 | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 3.00 cm | 2017-08-28 | 14:48 | 1
4.00 cm | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 4.00 cm | 2017-08-28 | 14:49 | 1
2017-08-25 |  | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 5.00 cm | 2017-08-28 | 14:49 | 1
15:30 | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 6.00 cm | 2017-08-28 | 14:50 | 1
4.50 cm | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 7.00 cm | 2017-08-28 | 14:51 | 1
2017-08-25 |  | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 8.00 cm | 2017-08-28 | 14:51 | 1
15:30 | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 6x6 | 6.00 cm | 9.00 cm | 2017-08-28 | 14:52 | 1
5.00 cm | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 8.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 0.50 cm | 2017-08-28 | 11:31 | 1
2017-08-25 |  | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 9.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 1.00 cm | 2017-08-28 | 11:31 | 1
15:31 | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 2.00 cm | 2017-08-28 | 11:32 | 1
5.50 cm | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 2.50 cm | 2017-08-28 | 11:33 | 1
2017-08-25 |  | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 3.00 cm | 2017-08-28 | 11:33 | 1
15:32 | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 4.00 cm | 2017-08-28 | 11:34 | 1
6.00 cm | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 5.00 cm | 2017-08-28 | 11:34 | 1
2017-08-25 |  | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 6.00 cm | 2017-08-28 | 11:35 | 1
15:33 | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 7.00 cm | 2017-08-28 | 11:36 | 1
6.50 cm | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 8.00 cm | 2017-08-28 | 11:36 | 1
2017-08-25 |  | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 3.00 cm | 9.00 cm | 2017-08-28 | 11:37 | 1
15:34 | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 8.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 0.50 cm | 2017-08-25 | 15:04 | 1
7.00 cm | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 9.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 1.00 cm | 2017-08-25 | 15:04 | 1
2017-08-25 |  | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 0.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 2.00 cm | 2017-08-25 | 15:05 | 1
15:34 | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 2.50 cm | 2017-08-25 | 15:06 | 1
15x15 |  | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 3.00 cm | 2017-08-25 | 15:07 | 1
11.00 cm |  | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 1.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 4.00 cm | 2017-08-25 | 15:08 | 1
0.50 cm | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 5.00 cm | 2017-08-25 | 15:08 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 6.00 cm | 2017-08-25 | 15:09 | 1
13:49 | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 7.00 cm | 2017-08-25 | 15:10 | 1
1.00 cm | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 3.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 8.00 cm | 2017-08-25 | 15:11 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 7.00 cm | 9.00 cm | 2017-08-25 | 15:12 | 1
13:50 | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 4.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 0.00 cm | 2017-08-25 | 14:24 | 1
2.00 cm | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-08-25 | 14:25 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 5.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 1.00 cm | 2017-08-25 | 14:26 | 1
13:51 | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 1.50 cm | 2017-08-25 | 14:27 | 1
2.50 cm | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 6.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 2.00 cm | 2017-08-25 | 14:28 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 2.50 cm | 2017-08-25 | 14:28 | 1
13:52 | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 8.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 3.00 cm | 2017-08-25 | 14:29 | 1
3.00 cm | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 9.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 3.50 cm | 2017-08-25 | 14:30 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 4.00 cm | 2017-08-25 | 14:31 | 1
13:53 | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 4.50 cm | 2017-08-25 | 14:32 | 1
4.00 cm | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 5.00 cm | 2017-08-25 | 14:33 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 5.50 cm | 2017-08-25 | 14:34 | 1
13:54 | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 6.00 cm | 2017-08-25 | 14:34 | 1
4.50 cm | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 6.50 cm | 2017-08-25 | 14:35 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 7.00 cm | 2017-08-25 | 14:36 | 1
13:55 | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 8.00 cm | 2017-08-25 | 14:37 | 1
5.00 cm | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 10x10 | 10.00 cm | 9.00 cm | 2017-08-25 | 14:38 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 8.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 0.50 cm | 2017-08-28 | 13:35 | 1
13:56 | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 9.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 1.00 cm | 2017-08-28 | 13:36 | 1
5.50 cm | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 2.00 cm | 2017-08-28 | 13:37 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 2.50 cm | 2017-08-28 | 13:38 | 1
13:57 | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 3.00 cm | 2017-08-28 | 13:39 | 1
6.00 cm | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 4.00 cm | 2017-08-28 | 13:40 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 5.00 cm | 2017-08-28 | 13:41 | 1
13:58 | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 6.00 cm | 2017-08-28 | 13:42 | 1
6.50 cm | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 7.00 cm | 2017-08-28 | 13:43 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 8.00 cm | 2017-08-28 | 13:44 | 1
13:59 | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 11.00 cm | 9.00 cm | 2017-08-28 | 13:45 | 1
7.00 cm | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 8.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 0.50 cm | 2017-08-28 | 12:19 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 9.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 1.00 cm | 2017-08-28 | 12:20 | 1
14:00 | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 2.00 cm | 2017-08-28 | 12:21 | 1
15.00 cm |  | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 2.50 cm | 2017-08-28 | 12:22 | 1
0.50 cm | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 3.00 cm | 2017-08-28 | 12:24 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 4.00 cm | 2017-08-28 | 12:25 | 1
12:33 | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 5.00 cm | 2017-08-28 | 12:26 | 1
1.00 cm | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 6.00 cm | 2017-08-28 | 12:27 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 7.00 cm | 2017-08-28 | 12:28 | 1
12:35 | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 8.00 cm | 2017-08-28 | 12:29 | 1
2.00 cm | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 15x15 | 15.00 cm | 9.00 cm | 2017-08-28 | 12:30 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 8.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 0.50 cm | 2017-08-28 | 10:12 | 1
12:36 | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 9.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 1.00 cm | 2017-08-28 | 10:13 | 1
2.50 cm | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 2.00 cm | 2017-08-28 | 10:14 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 2.50 cm | 2017-08-28 | 10:15 | 1
12:37 | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 3.00 cm | 2017-08-28 | 10:16 | 1
3.00 cm | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 4.00 cm | 2017-08-28 | 10:17 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 5.00 cm | 2017-08-28 | 10:19 | 1
12:38 | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 6.00 cm | 2017-08-28 | 10:20 | 1
4.00 cm | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 7.00 cm | 2017-08-28 | 10:21 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 8.00 cm | 2017-08-28 | 10:22 | 1
12:39 | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 20x20 | 20.00 cm | 9.00 cm | 2017-08-28 | 10:23 | 1
4.50 cm | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 8.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 0.50 cm | 2017-08-25 | 12:45 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 9.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 1.00 cm | 2017-08-25 | 12:47 | 1
12:40 | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 2.00 cm | 2017-08-25 | 12:48 | 1
5.00 cm | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 2.50 cm | 2017-08-25 | 12:49 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 3.00 cm | 2017-08-25 | 12:50 | 1
12:41 | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 4.00 cm | 2017-08-25 | 12:51 | 1
5.50 cm | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 3.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 5.00 cm | 2017-08-25 | 12:52 | 1
2017-08-28 |  | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 4.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 6.00 cm | 2017-08-25 | 12:54 | 1
12:42 | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 5.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 7.00 cm | 2017-08-25 | 12:55 | 1
6.00 cm | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 6.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 8.00 cm | 2017-08-25 | 12:56 | 1
2017-09-18 |  | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 7.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 21.00 cm | 9.00 cm | 2017-08-25 | 12:57 | 1
12:18 | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 8.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 0.50 cm | 2017-08-25 | 11:12 | 1
6.50 cm | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 9.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 1.00 cm | 2017-08-25 | 11:14 | 1
2017-08-28 |  | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 0.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 2.00 cm | 2017-08-25 | 11:15 | 1
12:44 | 1 | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 0.80 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 2.50 cm | 2017-08-25 | 11:16 | 1
7.00 cm | 1 | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 1.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 3.00 cm | 2017-08-25 | 11:17 | 1
2017-09-18 |  | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 1.40 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 4.00 cm | 2017-08-25 | 11:19 | 1
12:20 | 1 | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 1.80 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 5.00 cm | 2017-08-25 | 11:20 | 1
20x20 |  | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.00 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 6.00 cm | 2017-08-25 | 11:21 | 1
20.00 cm |  | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.20 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 7.00 cm | 2017-08-25 | 11:23 | 1
0.50 cm | 1 | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.50 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 8.00 cm | 2017-08-25 | 11:24 | 1
2017-08-28 |  | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.70 cm |  |  | 1 | 100.00 cm | 20 MeV | 25x25 | 25.00 cm | 9.00 cm | 2017-08-25 | 11:25 | 1
10:27 | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 0.50 cm | 2017-09-19 | 14:55 | 1
1.00 cm | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 0.80 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 0.80 cm | 2017-09-19 | 14:55 | 1
2017-08-28 |  | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 1.00 cm | 2017-09-19 | 14:56 | 1
10:28 | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 1.40 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 1.40 cm | 2017-09-19 | 14:57 | 1
2.00 cm | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 1.80 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 1.80 cm | 2017-09-19 | 14:57 | 1
2017-08-28 |  | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.00 cm | 2017-09-19 | 14:58 | 1
10:30 | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.20 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.20 cm | 2017-09-19 | 14:59 | 1
2.50 cm | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.50 cm | 2017-09-19 | 14:59 | 1
2017-08-28 |  | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.70 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 4.00 cm | 2.70 cm | 2017-09-19 | 15:00 | 1
10:31 | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 0.50 cm | 2017-09-19 | 13:54 | 1
3.00 cm | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 0.80 cm | 2017-08-29 | 16:55 | 1
2017-08-28 |  | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 0.80 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 1.00 cm | 2017-08-29 | 16:56 | 1
10:32 | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 1.40 cm | 2017-08-29 | 16:57 | 1
4.00 cm | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 1.40 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 1.80 cm | 2017-09-19 | 13:54 | 1
2017-08-28 |  | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 1.80 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.00 cm | 2017-08-29 | 16:58 | 1
10:33 | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.20 cm | 2017-09-19 | 13:55 | 1
4.50 cm | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.20 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.50 cm | 2017-08-29 | 16:58 | 1
2017-08-28 |  | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 2.70 cm | 2017-09-19 | 13:56 | 1
10:34 | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.70 cm |  |  | 1 | 110.00 cm | 6 MeV | 6x6 | 6.00 cm | 3.00 cm | 2017-08-29 | 17:00 | 1
5.00 cm | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-09-19 | 12:03 | 1
2017-08-28 |  | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 0.80 cm | 2017-08-29 | 15:48 | 1
10:35 | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 0.80 cm |  |  | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 1.00 cm | 2017-08-29 | 15:49 | 1
5.50 cm | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 1.40 cm | 2017-08-29 | 15:50 | 1
2017-08-28 |  | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 1.40 cm |  |  | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 1.80 cm | 2017-09-19 | 12:04 | 1
10:36 | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 1.80 cm |  |  | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.00 cm | 2017-08-29 | 15:51 | 1
6.00 cm | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.20 cm | 2017-09-19 | 12:05 | 1
2017-08-28 |  | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.20 cm |  |  | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.50 cm | 2017-08-29 | 15:52 | 1
10:37 | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 2.70 cm | 2017-09-19 | 12:06 | 1
6.50 cm | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.70 cm |  |  | 1 | 110.00 cm | 6 MeV | 10x10 | 10.00 cm | 3.00 cm | 2017-08-29 | 15:53 | 1
2017-08-28 |  | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 0.50 cm | 2017-09-19 | 11:14 | 1
10:39 | 1 | 110.00 cm | 6 MeV | 20x20 | 20.00 cm | 0.80 cm |  |  | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 0.80 cm | 2017-08-29 | 14:27 | 1
7.00 cm | 1 | 110.00 cm | 6 MeV | 20x20 | 20.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 1.00 cm | 2017-08-29 | 14:28 | 1
2017-08-28 |  | 110.00 cm | 6 MeV | 20x20 | 20.00 cm | 1.40 cm |  |  | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 1.40 cm | 2017-08-29 | 14:29 | 1
10:40 | 1 | 110.00 cm | 6 MeV | 20x20 | 20.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 1.80 cm | 2017-09-19 | 11:16 | 1
25x25 |  | 110.00 cm | 6 MeV | 20x20 | 20.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.00 cm | 2017-08-29 | 14:31 | 1
21.00 cm |  | 110.00 cm | 6 MeV | 20x20 | 20.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.20 cm | 2017-09-19 | 11:17 | 1
0.50 cm | 1 | 110.00 cm | 6 MeV | 25x25 | 25.00 cm | 0.80 cm |  |  | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.50 cm | 2017-08-29 | 14:32 | 1
2017-08-25 |  | 110.00 cm | 6 MeV | 25x25 | 25.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 2.70 cm | 2017-09-19 | 11:18 | 1
13:01 | 1 | 110.00 cm | 6 MeV | 25x25 | 25.00 cm | 1.40 cm |  |  | 1 | 110.00 cm | 6 MeV | 15x15 | 15.00 cm | 3.00 cm | 2017-08-29 | 14:33 | 1
1.00 cm | 1 | 110.00 cm | 6 MeV | 25x25 | 25.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 20x20 | 20.00 cm | 0.80 cm | 2017-08-29 | 13:10 | 1
2017-08-25 |  | 110.00 cm | 6 MeV | 25x25 | 25.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 6 MeV | 20x20 | 20.00 cm | 1.00 cm | 2017-08-29 | 13:11 | 1
13:02 | 1 | 110.00 cm | 6 MeV | 25x25 | 25.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 20x20 | 20.00 cm | 1.40 cm | 2017-08-29 | 13:13 | 1
2.00 cm | 1 | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 6 MeV | 20x20 | 20.00 cm | 2.00 cm | 2017-08-29 | 13:14 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 20x20 | 20.00 cm | 2.50 cm | 2017-08-29 | 13:15 | 1
13:03 | 1 | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 1.50 cm |  |  | 1 | 110.00 cm | 6 MeV | 20x20 | 20.00 cm | 3.00 cm | 2017-08-29 | 13:17 | 1
2.50 cm | 1 | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 25x25 | 25.00 cm | 0.80 cm | 2017-08-29 | 11:33 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 6 MeV | 25x25 | 25.00 cm | 1.00 cm | 2017-08-29 | 11:34 | 1
13:04 | 1 | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 25x25 | 25.00 cm | 1.40 cm | 2017-08-29 | 11:36 | 1
3.00 cm | 1 | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 3.50 cm |  |  | 1 | 110.00 cm | 6 MeV | 25x25 | 25.00 cm | 2.00 cm | 2017-08-29 | 11:37 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 6 MeV | 25x25 | 25.00 cm | 2.50 cm | 2017-08-29 | 11:39 | 1
13:06 | 1 | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 4.50 cm |  |  | 1 | 110.00 cm | 6 MeV | 25x25 | 25.00 cm | 3.00 cm | 2017-08-29 | 11:40 | 1
4.00 cm | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 0.50 cm | 2017-09-19 | 15:01 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 0.80 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 1.00 cm | 2017-09-19 | 15:03 | 1
13:07 | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 1.50 cm | 2017-09-19 | 15:03 | 1
4.50 cm | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 1.40 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 2.00 cm | 2017-09-19 | 15:04 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 2.50 cm | 2017-09-19 | 15:05 | 1
13:08 | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 3.00 cm | 2017-09-19 | 15:05 | 1
5.00 cm | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 3.50 cm | 2017-09-19 | 15:06 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 3.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 4.00 cm | 2017-09-19 | 15:07 | 1
13:09 | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 4.00 cm | 4.50 cm | 2017-09-19 | 15:07 | 1
5.50 cm | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 4.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 0.50 cm | 2017-09-19 | 13:57 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 0.80 cm | 2017-08-29 | 16:43 | 1
13:10 | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 0.80 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 1.00 cm | 2017-08-29 | 16:44 | 1
6.00 cm | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 1.40 cm | 2017-08-29 | 16:45 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 1.40 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 2.00 cm | 2017-08-29 | 16:46 | 1
13:11 | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 2.50 cm | 2017-08-29 | 16:47 | 1
6.50 cm | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 3.00 cm | 2017-08-29 | 16:48 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 3.50 cm | 2017-08-29 | 16:49 | 1
13:13 | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 3.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 4.00 cm | 2017-08-29 | 16:50 | 1
7.00 cm | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 6x6 | 6.00 cm | 4.50 cm | 2017-09-19 | 13:58 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 4.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-09-19 | 12:07 | 1
13:14 | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 0.80 cm | 2017-08-29 | 15:33 | 1
25.00 cm |  | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 0.80 cm |  |  | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 1.00 cm | 2017-08-29 | 15:35 | 1
0.50 cm | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 1.40 cm | 2017-08-29 | 15:36 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 1.40 cm |  |  | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 2.00 cm | 2017-08-29 | 15:37 | 1
11:32 | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 2.50 cm | 2017-08-29 | 15:38 | 1
1.00 cm | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 3.00 cm | 2017-08-29 | 15:39 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 3.50 cm | 2017-08-29 | 15:40 | 1
11:33 | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 3.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 4.00 cm | 2017-08-29 | 15:41 | 1
2.00 cm | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 10x10 | 10.00 cm | 4.50 cm | 2017-09-19 | 12:09 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 4.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 0.50 cm | 2017-09-19 | 11:19 | 1
11:34 | 1 | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 0.80 cm |  |  | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 0.80 cm | 2017-08-29 | 14:15 | 1
2.50 cm | 1 | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 1.00 cm | 2017-08-29 | 14:17 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 1.40 cm |  |  | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 1.40 cm | 2017-08-29 | 14:18 | 1
11:36 | 1 | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 2.00 cm | 2017-08-29 | 14:19 | 1
3.00 cm | 1 | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 2.50 cm | 2017-08-29 | 14:20 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 3.00 cm | 2017-08-29 | 14:21 | 1
11:37 | 1 | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 3.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 3.50 cm | 2017-08-29 | 14:22 | 1
4.00 cm | 1 | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 4.00 cm | 2017-08-29 | 14:24 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 0.80 cm |  |  | 1 | 110.00 cm | 9 MeV | 15x15 | 15.00 cm | 4.50 cm | 2017-09-19 | 11:20 | 1
11:38 | 1 | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 0.80 cm | 2017-08-29 | 12:58 | 1
4.50 cm | 1 | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 1.40 cm |  |  | 1 | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 1.00 cm | 2017-08-29 | 12:59 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 1.40 cm | 2017-08-29 | 13:00 | 1
11:40 | 1 | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 2.00 cm | 2017-08-29 | 13:02 | 1
5.00 cm | 1 | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 2.50 cm | 2017-08-29 | 13:03 | 1
2017-08-25 |  | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 3.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 3.00 cm | 2017-08-29 | 13:04 | 1
11:41 | 1 | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 3.50 cm | 2017-08-29 | 13:05 | 1
5.50 cm | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 20x20 | 20.00 cm | 4.00 cm | 2017-08-29 | 13:07 | 1
2017-08-25 |  | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 0.80 cm | 2017-08-29 | 11:20 | 1
11:42 | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 1.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 1.00 cm | 2017-08-29 | 11:21 | 1
6.00 cm | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 1.40 cm | 2017-08-29 | 11:23 | 1
2017-08-25 |  | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 2.00 cm | 2017-08-29 | 11:24 | 1
11:44 | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 2.50 cm | 2017-08-29 | 11:25 | 1
6.50 cm | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 3.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 3.00 cm | 2017-08-29 | 11:27 | 1
2017-08-25 |  | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 3.50 cm | 2017-08-29 | 11:28 | 1
11:45 | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 4.50 cm |  |  | 1 | 110.00 cm | 9 MeV | 25x25 | 25.00 cm | 4.00 cm | 2017-08-29 | 11:30 | 1
7.00 cm | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 0.50 cm | 2017-09-19 | 15:08 | 1
2017-08-25 |  | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 5.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 1.00 cm | 2017-09-19 | 15:10 | 1
11:46 | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 1.50 cm | 2017-09-19 | 15:11 | 1
6x6 |  | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 2.00 cm | 2017-09-19 | 15:11 | 1
3.00 cm |  | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 1.40 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 2.50 cm | 2017-09-19 | 15:12 | 1
0.50 cm | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 3.00 cm | 2017-09-19 | 15:12 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 3.50 cm | 2017-09-19 | 15:13 | 1
16:31 | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 4.00 cm | 2017-09-19 | 15:14 | 1
1.00 cm | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 3.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 4.50 cm | 2017-09-19 | 15:14 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 5.00 cm | 2017-09-19 | 15:15 | 1
16:31 | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 4.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 4.00 cm | 5.50 cm | 2017-09-19 | 15:16 | 1
2.00 cm | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 0.50 cm | 2017-08-29 | 16:31 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 5.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 1.00 cm | 2017-08-29 | 16:32 | 1
16:32 | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 1.40 cm | 2017-08-29 | 16:33 | 1
2.50 cm | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 2.00 cm | 2017-08-29 | 16:34 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 1.40 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 2.50 cm | 2017-08-29 | 16:35 | 1
16:33 | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 3.00 cm | 2017-08-29 | 16:35 | 1
3.00 cm | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 3.50 cm | 2017-09-19 | 13:59 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 4.00 cm | 2017-08-29 | 16:36 | 1
16:33 | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 3.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 4.50 cm | 2017-08-29 | 16:37 | 1
4.00 cm | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 5.00 cm | 2017-08-29 | 16:38 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 4.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 6x6 | 6.00 cm | 5.50 cm | 2017-08-29 | 16:39 | 1
16:34 | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-08-29 | 15:15 | 1
4.50 cm | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 5.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 1.00 cm | 2017-08-29 | 15:16 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 1.40 cm | 2017-08-29 | 15:17 | 1
16:34 | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 2.00 cm | 2017-08-29 | 15:18 | 1
5.00 cm | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 1.40 cm |  |  | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 2.50 cm | 2017-08-29 | 15:19 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 3.00 cm | 2017-08-29 | 15:21 | 1
16:35 | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 3.50 cm | 2017-09-19 | 12:10 | 1
5.50 cm | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 4.00 cm | 2017-08-29 | 15:22 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 3.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 4.50 cm | 2017-08-29 | 15:23 | 1
16:36 | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 5.00 cm | 2017-08-29 | 15:24 | 1
6.00 cm | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 4.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 10x10 | 10.00 cm | 5.50 cm | 2017-08-29 | 15:25 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 0.50 cm | 2017-08-29 | 14:00 | 1
16:36 | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 5.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 1.00 cm | 2017-08-29 | 14:01 | 1
6.50 cm | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 1.40 cm | 2017-08-29 | 14:03 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 2.00 cm | 2017-08-29 | 14:04 | 1
16:37 | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 1.40 cm |  |  | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 2.50 cm | 2017-08-29 | 14:06 | 1
7.00 cm | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 3.00 cm | 2017-08-29 | 14:07 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 3.50 cm | 2017-09-19 | 11:21 | 1
16:38 | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 4.00 cm | 2017-08-29 | 14:08 | 1
4.00 cm |  | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 4.50 cm | 2017-08-29 | 14:09 | 1
0.50 cm | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 4.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 5.00 cm | 2017-08-29 | 14:10 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 15x15 | 15.00 cm | 5.50 cm | 2017-08-29 | 14:12 | 1
15:43 | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 5.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 0.50 cm | 2017-08-29 | 12:43 | 1
1.00 cm | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 1.00 cm | 2017-08-29 | 12:44 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 1.40 cm | 2017-08-29 | 12:45 | 1
15:43 | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 1.40 cm |  |  | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 2.00 cm | 2017-08-29 | 12:47 | 1
2.00 cm | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 2.50 cm | 2017-08-29 | 12:48 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 3.00 cm | 2017-08-29 | 12:49 | 1
15:44 | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 4.00 cm | 2017-08-29 | 12:50 | 1
2.50 cm | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 4.50 cm | 2017-08-29 | 12:52 | 1
2017-08-28 |  | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 4.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 5.00 cm | 2017-08-29 | 12:53 | 1
15:45 | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 20x20 | 20.00 cm | 5.50 cm | 2017-08-29 | 12:54 | 1
3.00 cm | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 5.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 0.50 cm | 2017-08-29 | 11:03 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 1.00 cm | 2017-08-29 | 11:05 | 1
15:45 | 1 | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 1.40 cm | 2017-08-29 | 11:06 | 1
4.00 cm | 1 | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 2.00 cm | 2017-08-29 | 11:08 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 2.50 cm | 2017-08-29 | 11:09 | 1
15:46 | 1 | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 3.00 cm | 2017-08-29 | 11:10 | 1
4.50 cm | 1 | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 4.00 cm | 2017-08-29 | 11:12 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 6.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 4.50 cm | 2017-08-29 | 11:13 | 1
15:46 | 1 | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 7.00 cm |  |  | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 5.00 cm | 2017-08-29 | 11:15 | 1
5.00 cm | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 12 MeV | 25x25 | 25.00 cm | 5.50 cm | 2017-08-29 | 11:16 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 0.50 cm | 2017-09-19 | 15:16 | 2
15:47 | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 1.00 cm | 2017-09-19 | 15:19 | 2
5.50 cm | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 2.00 cm | 2017-09-19 | 15:19 | 2
2017-08-28 |  | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 3.00 cm | 2017-09-19 | 15:20 | 2
15:48 | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 4.00 cm | 2017-09-19 | 15:21 | 2
6.00 cm | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 4.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 5.00 cm | 2017-09-19 | 15:22 | 2
2017-09-18 |  | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 6.00 cm | 2017-09-19 | 15:22 | 2
15:26 | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 5.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 7.00 cm | 2017-09-19 | 15:23 | 2
6.50 cm | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 6.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 0.50 cm | 2017-08-29 | 16:14 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 6.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 1.00 cm | 2017-08-29 | 16:15 | 1
15:49 | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 7.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 2.00 cm | 2017-08-29 | 16:16 | 1
7.00 cm | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 2.50 cm | 2017-08-29 | 16:17 | 1
2017-09-18 |  | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 3.00 cm | 2017-08-29 | 16:18 | 1
15:28 | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 4.00 cm | 2017-08-29 | 16:19 | 1
6.00 cm |  | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 4.50 cm | 2017-08-29 | 16:20 | 1
0.50 cm | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 5.00 cm | 2017-08-29 | 16:21 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 5.50 cm | 2017-08-29 | 16:22 | 1
14:56 | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 4.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 6.00 cm | 2017-08-29 | 16:23 | 1
1.00 cm | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 6.50 cm | 2017-08-29 | 16:24 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 5.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 6x6 | 6.00 cm | 7.00 cm | 2017-08-29 | 16:25 | 1
14:56 | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 6.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-08-29 | 14:57 | 1
2.00 cm | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 6.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 1.00 cm | 2017-08-29 | 14:58 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 7.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 2.00 cm | 2017-08-29 | 14:59 | 1
14:57 | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 2.50 cm | 2017-08-29 | 15:00 | 1
2.50 cm | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 3.00 cm | 2017-08-29 | 15:01 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 4.00 cm | 2017-08-29 | 15:02 | 1
14:58 | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 4.50 cm | 2017-08-29 | 15:04 | 1
3.00 cm | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 5.00 cm | 2017-08-29 | 15:05 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 5.50 cm | 2017-08-29 | 15:06 | 1
14:58 | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 4.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 6.00 cm | 2017-08-29 | 15:07 | 1
4.00 cm | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 6.50 cm | 2017-08-29 | 15:08 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 5.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 10x10 | 10.00 cm | 7.00 cm | 2017-08-29 | 15:09 | 1
14:59 | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 6.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 0.50 cm | 2017-08-29 | 13:42 | 1
4.50 cm | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 6.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 1.00 cm | 2017-08-29 | 13:43 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 7.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 2.00 cm | 2017-08-29 | 13:44 | 1
15:00 | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 2.50 cm | 2017-08-29 | 13:45 | 1
5.00 cm | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 3.00 cm | 2017-08-29 | 13:46 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 4.00 cm | 2017-08-29 | 13:47 | 1
15:00 | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 4.50 cm | 2017-08-29 | 13:49 | 1
5.50 cm | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 5.00 cm | 2017-08-29 | 13:50 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 5.50 cm | 2017-08-29 | 13:51 | 1
15:01 | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 4.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 6.00 cm | 2017-08-29 | 13:52 | 1
6.00 cm | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 6.50 cm | 2017-08-29 | 13:53 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 5.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 15x15 | 15.00 cm | 7.00 cm | 2017-08-29 | 13:55 | 1
15:02 | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 6.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 0.50 cm | 2017-08-29 | 12:25 | 1
6.50 cm | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 6.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 1.00 cm | 2017-08-29 | 12:26 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 7.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 2.00 cm | 2017-08-29 | 12:27 | 1
15:02 | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 2.50 cm | 2017-08-29 | 12:29 | 1
7.00 cm | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 3.00 cm | 2017-08-29 | 12:30 | 1
2017-08-28 |  | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 4.00 cm | 2017-08-29 | 12:31 | 1
15:03 | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 4.50 cm | 2017-08-29 | 12:32 | 1
20 MeV |  | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 5.00 cm | 2017-08-29 | 12:34 | 1
10x10 |  | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 5.50 cm | 2017-08-29 | 12:35 | 1
10.00 cm |  | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 4.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 6.00 cm | 2017-08-29 | 12:36 | 1
0.00 cm | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 6.50 cm | 2017-08-29 | 12:38 | 1
2017-08-25 |  | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 5.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 20x20 | 20.00 cm | 7.00 cm | 2017-08-29 | 12:39 | 1
14:24 | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 6.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 0.50 cm | 2017-08-29 | 10:44 | 1
0.50 cm | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 6.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 1.00 cm | 2017-08-29 | 10:45 | 1
2017-08-25 |  | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 7.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 2.00 cm | 2017-08-29 | 10:47 | 1
14:25 | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 2.50 cm | 2017-08-29 | 10:48 | 1
1.00 cm | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 3.00 cm | 2017-08-29 | 10:49 | 1
2017-08-25 |  | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 4.00 cm | 2017-08-29 | 10:51 | 1
14:26 | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 4.50 cm | 2017-08-29 | 10:52 | 1
1.50 cm | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 5.00 cm | 2017-08-29 | 10:54 | 1
2017-08-25 |  | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 5.50 cm | 2017-08-29 | 10:55 | 1
14:27 | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 6.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 6.00 cm | 2017-08-29 | 10:56 | 1
2.00 cm | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 7.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 6.50 cm | 2017-08-29 | 10:58 | 1
2017-08-25 |  | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 8.00 cm |  |  | 1 | 110.00 cm | 16 MeV | 25x25 | 25.00 cm | 7.00 cm | 2017-08-29 | 10:59 | 1
14:28 | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 9.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 0.50 cm | 2017-09-19 | 15:24 | 1
2.50 cm | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 1.00 cm | 2017-09-19 | 15:25 | 1
2017-08-25 |  | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 2.00 cm | 2017-09-19 | 15:26 | 1
14:28 | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 3.00 cm | 2017-09-19 | 15:26 | 1
3.00 cm | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 4.00 cm | 2017-09-19 | 15:27 | 1
2017-08-25 |  | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 5.00 cm | 2017-09-19 | 15:28 | 1
14:29 | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 6.00 cm | 2017-09-19 | 15:28 | 1
3.50 cm | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 7.00 cm | 2017-09-19 | 15:29 | 1
2017-08-25 |  | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 6.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 8.00 cm | 2017-09-19 | 15:30 | 1
14:30 | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 7.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 4.00 cm | 9.00 cm | 2017-09-19 | 15:30 | 1
4.00 cm | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 8.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 0.50 cm | 2017-08-29 | 16:02 | 1
2017-08-25 |  | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 9.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 1.00 cm | 2017-08-29 | 16:03 | 1
14:31 | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 2.00 cm | 2017-08-29 | 16:04 | 1
4.50 cm | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 2.50 cm | 2017-08-29 | 16:05 | 1
2017-08-25 |  | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 3.00 cm | 2017-08-29 | 16:05 | 1
14:32 | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 4.00 cm | 2017-08-29 | 16:06 | 1
5.00 cm | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 5.00 cm | 2017-08-29 | 16:07 | 1
2017-08-25 |  | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 6.00 cm | 2017-08-29 | 16:08 | 1
14:33 | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 7.00 cm | 2017-08-29 | 16:09 | 1
5.50 cm | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 6.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 8.00 cm | 2017-08-29 | 16:10 | 1
2017-08-25 |  | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 7.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 6x6 | 6.00 cm | 9.00 cm | 2017-08-29 | 16:11 | 1
14:34 | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 8.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-08-29 | 14:42 | 1
6.00 cm | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 9.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 1.00 cm | 2017-08-29 | 14:43 | 1
2017-08-25 |  | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 2.00 cm | 2017-08-29 | 14:44 | 1
14:34 | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 2.50 cm | 2017-08-29 | 14:45 | 1
6.50 cm | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 3.00 cm | 2017-08-29 | 14:46 | 1
2017-08-25 |  | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 4.00 cm | 2017-08-29 | 14:47 | 1
14:35 | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 5.00 cm | 2017-08-29 | 14:48 | 1
7.00 cm | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 6.00 cm | 2017-08-29 | 14:50 | 1
2017-08-25 |  | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 7.00 cm | 2017-08-29 | 14:51 | 1
14:36 | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 6.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 8.00 cm | 2017-08-29 | 14:52 | 1
8.00 cm | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 7.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 10x10 | 10.00 cm | 9.00 cm | 2017-08-29 | 14:53 | 1
2017-08-25 |  | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 8.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 0.50 cm | 2017-08-29 | 13:25 | 1
14:37 | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 9.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 1.00 cm | 2017-08-29 | 13:26 | 1
9.00 cm | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 2.00 cm | 2017-08-29 | 13:27 | 1
2017-08-25 |  | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 2.50 cm | 2017-08-29 | 13:28 | 1
14:38 | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 3.00 cm | 2017-08-29 | 13:29 | 1
3.00 cm |  | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 4.00 cm | 2017-08-29 | 13:30 | 1
0.50 cm | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 5.00 cm | 2017-08-29 | 13:32 | 1
2017-08-28 |  | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 6.00 cm | 2017-08-29 | 13:33 | 1
11:31 | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 7.00 cm | 2017-08-29 | 13:34 | 1
1.00 cm | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 6.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 8.00 cm | 2017-08-29 | 13:35 | 1
2017-08-28 |  | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 7.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 15x15 | 15.00 cm | 9.00 cm | 2017-08-29 | 13:36 | 1
11:31 | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 8.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 0.50 cm | 2017-08-29 | 12:08 | 1
2.00 cm | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 9.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 1.00 cm | 2017-08-29 | 12:09 | 1
2017-08-28 |  | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 0.50 cm |  |  | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 2.00 cm | 2017-08-29 | 12:10 | 1
11:32 | 1 | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 1.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 2.50 cm | 2017-08-29 | 12:12 | 1
2.50 cm | 1 | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 2.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 3.00 cm | 2017-08-29 | 12:13 | 1
2017-08-28 |  | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 2.50 cm |  |  | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 4.00 cm | 2017-08-29 | 12:14 | 1
11:33 | 1 | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 3.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 5.00 cm | 2017-08-29 | 12:15 | 1
3.00 cm | 1 | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 4.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 6.00 cm | 2017-08-29 | 12:16 | 1
2017-08-28 |  | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 5.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 7.00 cm | 2017-08-29 | 12:18 | 1
11:33 | 1 | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 6.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 8.00 cm | 2017-08-29 | 12:19 | 1
4.00 cm | 1 | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 7.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 20x20 | 20.00 cm | 9.00 cm | 2017-08-29 | 12:20 | 1
2017-08-28 |  | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 8.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 0.50 cm | 2017-08-29 | 10:26 | 1
11:34 | 1 | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 9.00 cm |  |  | 1 | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 1.00 cm | 2017-08-29 | 10:27 | 1
5.00 cm | 1 | [8 empty] | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 2.00 cm | 2017-08-29 | 10:29 | 1
2017-08-28 | [9 empty] | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 2.50 cm | 2017-08-29 | 10:30 | 1
11:34 | 1 | [8 empty] | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 3.00 cm | 2017-08-29 | 10:31 | 1
6.00 cm | 1 | [8 empty] | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 4.00 cm | 2017-08-29 | 10:33 | 1
2017-08-28 | [9 empty] | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 5.00 cm | 2017-08-29 | 10:34 | 1
11:35 | 1 | [8 empty] | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 6.00 cm | 2017-08-29 | 10:36 | 1
7.00 cm | 1 | [8 empty] | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 7.00 cm | 2017-08-29 | 10:37 | 1
2017-08-28 | [9 empty] | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 8.00 cm | 2017-08-29 | 10:38 | 1
11:36 | 1 | [8 empty] | 110.00 cm | 20 MeV | 25x25 | 25.00 cm | 9.00 cm | 2017-08-29 | 10:40 | 1
8.00 cm | 1
2017-08-28
11:36 | 1
9.00 cm | 1
2017-08-28
11:37 | 1
7.00 cm
0.50 cm | 1
2017-08-25
15:04 | 1
1.00 cm | 1
2017-08-25
15:04 | 1
2.00 cm | 1
2017-08-25
15:05 | 1
2.50 cm | 1
2017-08-25
15:06 | 1
3.00 cm | 1
2017-08-25
15:07 | 1
4.00 cm | 1
2017-08-25
15:08 | 1
5.00 cm | 1
2017-08-25
15:08 | 1
6.00 cm | 1
2017-08-25
15:09 | 1
7.00 cm | 1
2017-08-25
15:10 | 1
8.00 cm | 1
2017-08-25
15:11 | 1
9.00 cm | 1
2017-08-25
15:12 | 1
15x15
11.00 cm
0.50 cm | 1
2017-08-28
13:35 | 1
1.00 cm | 1
2017-08-28
13:36 | 1
2.00 cm | 1
2017-08-28
13:37 | 1
2.50 cm | 1
2017-08-28
13:38 | 1
3.00 cm | 1
2017-08-28
13:39 | 1
4.00 cm | 1
2017-08-28
13:40 | 1
5.00 cm | 1
2017-08-28
13:41 | 1
6.00 cm | 1
2017-08-28
13:42 | 1
7.00 cm | 1
2017-08-28
13:43 | 1
8.00 cm | 1
2017-08-28
13:44 | 1
9.00 cm | 1
2017-08-28
13:45 | 1
15.00 cm
0.50 cm | 1
2017-08-28
12:19 | 1
1.00 cm | 1
2017-08-28
12:20 | 1
2.00 cm | 1
2017-08-28
12:21 | 1
2.50 cm | 1
2017-08-28
12:22 | 1
3.00 cm | 1
2017-08-28
12:24 | 1
4.00 cm | 1
2017-08-28
12:25 | 1
5.00 cm | 1
2017-08-28
12:26 | 1
6.00 cm | 1
2017-08-28
12:27 | 1
7.00 cm | 1
2017-08-28
12:28 | 1
8.00 cm | 1
2017-08-28
12:29 | 1
9.00 cm | 1
2017-08-28
12:30 | 1
20x20
20.00 cm
0.50 cm | 1
2017-08-28
10:12 | 1
1.00 cm | 1
2017-08-28
10:13 | 1
2.00 cm | 1
2017-08-28
10:14 | 1
2.50 cm | 1
2017-08-28
10:15 | 1
3.00 cm | 1
2017-08-28
10:16 | 1
4.00 cm | 1
2017-08-28
10:17 | 1
5.00 cm | 1
2017-08-28
10:19 | 1
6.00 cm | 1
2017-08-28
10:20 | 1
7.00 cm | 1
2017-08-28
10:21 | 1
8.00 cm | 1
2017-08-28
10:22 | 1
9.00 cm | 1
2017-08-28
10:23 | 1
25x25
21.00 cm
0.50 cm | 1
2017-08-25
12:45 | 1
1.00 cm | 1
2017-08-25
12:47 | 1
2.00 cm | 1
2017-08-25
12:48 | 1
2.50 cm | 1
2017-08-25
12:49 | 1
3.00 cm | 1
2017-08-25
12:50 | 1
4.00 cm | 1
2017-08-25
12:51 | 1
5.00 cm | 1
2017-08-25
12:52 | 1
6.00 cm | 1
2017-08-25
12:54 | 1
7.00 cm | 1
2017-08-25
12:55 | 1
8.00 cm | 1
2017-08-25
12:56 | 1
9.00 cm | 1
2017-08-25
12:57 | 1
25.00 cm
0.50 cm | 1
2017-08-25
11:12 | 1
1.00 cm | 1
2017-08-25
11:14 | 1
2.00 cm | 1
2017-08-25
11:15 | 1
2.50 cm | 1
2017-08-25
11:16 | 1
3.00 cm | 1
2017-08-25
11:17 | 1
4.00 cm | 1
2017-08-25
11:19 | 1
5.00 cm | 1
2017-08-25
11:20 | 1
6.00 cm | 1
2017-08-25
11:21 | 1
7.00 cm | 1
2017-08-25
11:23 | 1
8.00 cm | 1
2017-08-25
11:24 | 1
9.00 cm | 1
2017-08-25
11:25 | 1
6x6
3.00 cm
0.50 cm | 1
2017-08-28
16:20 | 1
1.00 cm | 1
2017-08-28
16:21 | 1
2.00 cm | 1
2017-08-28
16:22 | 1
2.50 cm | 1
2017-08-28
16:22 | 1
3.00 cm | 1
2017-08-28
16:23 | 1
4.00 cm | 1
2017-08-28
16:23 | 1
5.00 cm | 1
2017-08-28
16:24 | 1
6.00 cm | 1
2017-08-28
16:25 | 1
7.00 cm | 1
2017-08-28
16:25 | 1
8.00 cm | 1
2017-08-28
16:26 | 1
9.00 cm | 1
2017-08-28
16:27 | 1
4.00 cm
0.50 cm | 1
2017-08-28
15:33 | 1
1.00 cm | 1
2017-08-28
15:33 | 1
2.00 cm | 1
2017-08-28
15:34 | 1
2.50 cm | 1
2017-08-28
15:35 | 1
3.00 cm | 1
2017-08-28
15:35 | 1
4.00 cm | 1
2017-08-28
15:36 | 1
5.00 cm | 1
2017-08-28
15:36 | 1
6.00 cm | 1
2017-08-28
15:37 | 1
7.00 cm | 1
2017-08-28
15:38 | 1
8.00 cm | 1
2017-08-28
15:38 | 1
9.00 cm | 1
2017-08-28
15:39 | 1
6.00 cm
0.50 cm | 1
2017-08-28
14:46 | 1
1.00 cm | 1
2017-08-28
14:46 | 1
2.00 cm | 1
2017-08-28
14:47 | 1
2.50 cm | 1
2017-08-28
14:48 | 1
3.00 cm | 1
2017-08-28
14:48 | 1
4.00 cm | 1
2017-08-28
14:49 | 1
5.00 cm | 1
2017-08-28
14:49 | 1
6.00 cm | 1
2017-08-28
14:50 | 1
7.00 cm | 1
2017-08-28
14:51 | 1
8.00 cm | 1
2017-08-28
14:51 | 1
9.00 cm | 1
2017-08-28
14:52 | 1
6 MeV
10x10
10.00 cm
0.00 cm | 1
2017-08-24
15:16 | 1
0.50 cm | 1
2017-09-18
14:18 | 1
0.80 cm | 1
2017-08-24
15:17 | 1
1.00 cm | 1
2017-08-24
15:18 | 1
1.40 cm | 1
2017-08-24
15:19 | 1
1.80 cm | 1
2017-09-18
14:19 | 1
2.00 cm | 1
2017-08-24
15:20 | 1
2.20 cm | 1
2017-09-18
14:20 | 1
2.50 cm | 1
2017-08-24
15:20 | 1
2.70 cm | 1
2017-09-18
14:21 | 1
3.00 cm | 1
2017-08-24
15:21 | 1
3.00 cm
0.80 cm | 1
2017-08-28
12:10 | 1
1.00 cm | 1
2017-08-28
12:11 | 1
1.40 cm | 1
2017-08-28
12:11 | 1
2.00 cm | 1
2017-08-28
12:12 | 1
2.50 cm | 1
2017-08-28
12:13 | 1
3.00 cm | 1
2017-08-28
12:13 | 1
7.00 cm
0.80 cm | 1
2017-08-25
16:20 | 1
1.00 cm | 1
2017-08-25
16:21 | 1
1.40 cm | 1
2017-08-25
16:22 | 1
2.00 cm | 1
2017-08-25
16:23 | 1
2.50 cm | 1
2017-08-25
16:24 | 1
3.00 cm | 1
2017-08-25
16:24 | 1
15x15
11.00 cm
0.80 cm | 1
2017-08-28
14:29 | 1
1.00 cm | 1
2017-08-28
14:30 | 1
1.40 cm | 1
2017-08-28
14:31 | 1
2.00 cm | 1
2017-08-28
14:32 | 1
2.50 cm | 1
2017-08-28
14:33 | 1
3.00 cm | 1
2017-08-28
14:34 | 1
15.00 cm
0.50 cm | 1
2017-09-18
12:08 | 1
0.80 cm | 1
2017-08-28
13:15 | 1
1.00 cm | 1
2017-08-28
13:17 | 1
1.40 cm | 1
2017-08-28
13:18 | 1
1.80 cm | 1
2017-09-18
12:09 | 1
2.00 cm | 1
2017-08-28
13:19 | 1
2.20 cm | 1
2017-09-18
12:10 | 1
2.50 cm | 1
2017-08-28
13:20 | 1
2.70 cm | 1
2017-09-18
12:12 | 1
3.00 cm | 1
2017-08-28
13:21 | 1
20x20
20.00 cm
0.80 cm | 1
2017-08-28
11:14 | 1
1.00 cm | 1
2017-08-28
11:15 | 1
1.40 cm | 1
2017-08-28
11:17 | 1
2.00 cm | 1
2017-08-28
11:18 | 1
2.50 cm | 1
2017-08-28
11:19 | 1
3.00 cm | 1
2017-08-28
11:20 | 1
25x25
21.00 cm
0.80 cm | 1
2017-08-25
14:05 | 1
1.00 cm | 1
2017-08-25
14:06 | 1
1.40 cm | 1
2017-08-25
14:07 | 1
2.00 cm | 1
2017-08-25
14:08 | 1
2.50 cm | 1
2017-08-25
14:10 | 1
3.00 cm | 1
2017-08-25
14:11 | 1
25.00 cm
0.80 cm | 1
2017-08-25
12:24 | 1
1.00 cm | 1
2017-08-25
12:25 | 1
1.40 cm | 1
2017-08-25
12:27 | 1
2.00 cm | 1
2017-08-25
12:28 | 1
2.50 cm | 1
2017-08-25
12:29 | 1
3.00 cm | 1
2017-08-25
12:31 | 1
6x6
3.00 cm
0.80 cm | 1
2017-08-28
16:58 | 1
1.00 cm | 1
2017-08-28
16:59 | 1
1.40 cm | 1
2017-08-28
17:00 | 1
2.00 cm | 1
2017-08-28
17:00 | 1
2.50 cm | 1
2017-08-28
17:01 | 1
3.00 cm | 1
2017-08-28
17:01 | 1
4.00 cm
0.50 cm | 1
2017-09-18
15:19 | 1
0.80 cm | 1
2017-08-28
16:12 | 1
1.00 cm | 1
2017-08-28
16:12 | 1
1.40 cm | 1
2017-08-28
16:13 | 1
1.80 cm | 1
2017-09-18
15:20 | 1
2.00 cm | 1
2017-08-28
16:14 | 1
2.20 cm | 1
2017-09-18
15:20 | 1
2.50 cm | 1
2017-08-28
16:14 | 1
2.70 cm | 1
2017-09-18
15:21 | 1
3.00 cm | 1
2017-08-28
16:15 | 1
6.00 cm
0.50 cm | 1
2017-09-18
15:38 | 1
0.80 cm | 1
2017-08-28
15:23 | 1
1.00 cm | 1
2017-08-28
15:23 | 1
1.40 cm | 1
2017-08-28
15:24 | 1
1.80 cm | 1
2017-09-18
15:38 | 1
2.00 cm | 1
2017-08-28
15:25 | 1
2.20 cm | 1
2017-09-18
15:39 | 1
2.50 cm | 1
2017-08-28
15:25 | 1
2.70 cm | 1
2017-09-18
15:40 | 1
3.00 cm | 1
2017-08-28
15:26 | 1
9 MeV
10x10
10.00 cm
0.00 cm | 1
2017-08-24
15:47 | 1
0.50 cm | 1
2017-09-18
14:22 | 1
0.80 cm | 1
2017-08-24
15:49 | 1
1.00 cm | 1
2017-08-24
15:50 | 1
1.40 cm | 1
2017-08-24
15:51 | 1
2.00 cm | 1
2017-08-24
15:52 | 1
2.50 cm | 1
2017-08-24
15:53 | 1
3.00 cm | 1
2017-08-24
15:54 | 1
3.50 cm | 1
2017-08-24
15:55 | 1
4.00 cm | 1
2017-08-24
15:56 | 1
4.50 cm | 1
2017-09-18
14:23 | 1
3.00 cm
0.80 cm | 1
2017-08-28
12:02 | 1
1.00 cm | 1
2017-08-28
12:03 | 1
1.40 cm | 1
2017-08-28
12:03 | 1
2.00 cm | 1
2017-08-28
12:04 | 1
2.50 cm | 1
2017-08-28
12:05 | 1
3.00 cm | 1
2017-08-28
12:05 | 1
3.50 cm | 1
2017-08-28
12:06 | 1
4.00 cm | 1
2017-08-28
12:07 | 1
7.00 cm
0.80 cm | 1
2017-08-25
16:03 | 1
1.00 cm | 1
2017-08-25
16:03 | 1
1.40 cm | 1
2017-08-25
16:04 | 1
2.00 cm | 1
2017-08-25
16:05 | 1
2.50 cm | 1
2017-08-25
16:06 | 1
3.00 cm | 1
2017-08-25
16:07 | 1
3.50 cm | 1
2017-08-25
16:07 | 1
4.00 cm | 1
2017-08-25
16:08 | 1
15x15
11.00 cm
0.80 cm | 1
2017-08-28
14:17 | 1
1.00 cm | 1
2017-08-28
14:18 | 1
1.40 cm | 1
2017-08-28
14:19 | 1
2.00 cm | 1
2017-08-28
14:20 | 1
2.50 cm | 1
2017-08-28
14:21 | 1
3.00 cm | 1
2017-08-28
14:22 | 1
3.50 cm | 1
2017-08-28
14:23 | 1
4.00 cm | 1
2017-08-28
14:24 | 1
15.00 cm
0.50 cm | 1
2017-09-18
12:13 | 1
0.80 cm | 1
2017-08-28
13:02 | 1
1.00 cm | 1
2017-08-28
13:03 | 1
1.40 cm | 1
2017-08-28
13:04 | 1
2.00 cm | 1
2017-08-28
13:05 | 1
2.50 cm | 1
2017-08-28
13:06 | 1
3.00 cm | 1
2017-08-28
13:08 | 1
3.50 cm | 1
2017-08-28
13:09 | 1
4.00 cm | 1
2017-08-28
13:10 | 1
4.50 cm | 1
2017-09-18
12:15 | 1
20x20
20.00 cm
0.80 cm | 1
2017-08-28
10:58 | 1
1.00 cm | 1
2017-08-28
10:59 | 1
1.40 cm | 1
2017-08-28
11:00 | 1
2.00 cm | 1
2017-08-28
11:01 | 1
2.50 cm | 1
2017-08-28
11:02 | 1
3.00 cm | 1
2017-08-28
11:04 | 1
3.50 cm | 1
2017-08-28
11:05 | 1
4.00 cm | 1
2017-08-28
11:06 | 1
25x25
21.00 cm
0.80 cm | 1
2017-08-25
13:53 | 1
1.00 cm | 1
2017-08-25
13:54 | 1
1.40 cm | 1
2017-08-25
13:55 | 1
2.00 cm | 1
2017-08-25
13:57 | 1
2.50 cm | 1
2017-08-25
13:58 | 1
3.00 cm | 1
2017-08-25
13:59 | 1
3.50 cm | 1
2017-08-25
14:00 | 1
4.00 cm | 1
2017-08-25
14:01 | 1
25.00 cm
0.80 cm | 1
2017-08-25
12:10 | 1
1.00 cm | 1
2017-08-25
12:11 | 1
1.40 cm | 1
2017-08-25
12:13 | 1
2.00 cm | 1
2017-08-25
12:14 | 1
2.50 cm | 1
2017-08-25
12:15 | 1
3.00 cm | 1
2017-08-25
12:17 | 1
3.50 cm | 1
2017-08-25
12:18 | 1
4.00 cm | 1
2017-08-25
12:20 | 1
6x6
3.00 cm
0.80 cm | 1
2017-08-28
16:51 | 1
1.00 cm | 1
2017-08-28
16:52 | 1
1.40 cm | 1
2017-08-28
16:52 | 1
2.00 cm | 1
2017-08-28
16:53 | 1
2.50 cm | 1
2017-08-28
16:54 | 1
3.00 cm | 1
2017-08-28
16:54 | 1
3.50 cm | 1
2017-08-28
16:55 | 1
4.00 cm | 1
2017-08-28
16:56 | 1
4.00 cm
0.50 cm | 1
2017-09-18
15:22 | 1
0.80 cm | 1
2017-08-28
16:04 | 1
1.00 cm | 1
2017-08-28
16:05 | 1
1.40 cm | 1
2017-08-28
16:05 | 1
2.00 cm | 1
2017-08-28
16:06 | 1
2.50 cm | 1
2017-08-28
16:07 | 1
3.00 cm | 1
2017-08-28
16:07 | 1
3.50 cm | 1
2017-08-28
16:08 | 1
4.00 cm | 1
2017-08-28
16:09 | 1
4.50 cm | 1
2017-09-18
15:24 | 1
6.00 cm
0.50 cm | 1
2017-09-18
15:41 | 1
0.80 cm | 1
2017-08-28
15:15 | 1
1.00 cm | 1
2017-08-28
15:16 | 1
1.40 cm | 1
2017-08-28
15:16 | 1
2.00 cm | 1
2017-08-28
15:17 | 1
2.50 cm | 1
2017-08-28
15:18 | 1
3.00 cm | 1
2017-08-28
15:18 | 1
3.50 cm | 1
2017-08-28
15:19 | 1
4.00 cm | 1
2017-08-28
15:20 | 1
4.50 cm | 1
2017-09-18
15:42 | 1
110.00 cm
12 MeV
10x10
10.00 cm
0.50 cm | 1
2017-08-29
15:15 | 1
1.00 cm | 1
2017-08-29
15:16 | 1
1.40 cm | 1
2017-08-29
15:17 | 1
2.00 cm | 1
2017-08-29
15:18 | 1
2.50 cm | 1
2017-08-29
15:19 | 1
3.00 cm | 1
2017-08-29
15:21 | 1
3.50 cm | 1
2017-09-19
12:10 | 1
4.00 cm | 1
2017-08-29
15:22 | 1
4.50 cm | 1
2017-08-29
15:23 | 1
5.00 cm | 1
2017-08-29
15:24 | 1
5.50 cm | 1
2017-08-29
15:25 | 1
15x15
15.00 cm
0.50 cm | 1
2017-08-29
14:00 | 1
1.00 cm | 1
2017-08-29
14:01 | 1
1.40 cm | 1
2017-08-29
14:03 | 1
2.00 cm | 1
2017-08-29
14:04 | 1
2.50 cm | 1
2017-08-29
14:06 | 1
3.00 cm | 1
2017-08-29
14:07 | 1
3.50 cm | 1
2017-09-19
11:21 | 1
4.00 cm | 1
2017-08-29
14:08 | 1
4.50 cm | 1
2017-08-29
14:09 | 1
5.00 cm | 1
2017-08-29
14:10 | 1
5.50 cm | 1
2017-08-29
14:12 | 1
20x20
20.00 cm
0.50 cm | 1
2017-08-29
12:43 | 1
1.00 cm | 1
2017-08-29
12:44 | 1
1.40 cm | 1
2017-08-29
12:45 | 1
2.00 cm | 1
2017-08-29
12:47 | 1
2.50 cm | 1
2017-08-29
12:48 | 1
3.00 cm | 1
2017-08-29
12:49 | 1
4.00 cm | 1
2017-08-29
12:50 | 1
4.50 cm | 1
2017-08-29
12:52 | 1
5.00 cm | 1
2017-08-29
12:53 | 1
5.50 cm | 1
2017-08-29
12:54 | 1
25x25
25.00 cm
0.50 cm | 1
2017-08-29
11:03 | 1
1.00 cm | 1
2017-08-29
11:05 | 1
1.40 cm | 1
2017-08-29
11:06 | 1
2.00 cm | 1
2017-08-29
11:08 | 1
2.50 cm | 1
2017-08-29
11:09 | 1
3.00 cm | 1
2017-08-29
11:10 | 1
4.00 cm | 1
2017-08-29
11:12 | 1
4.50 cm | 1
2017-08-29
11:13 | 1
5.00 cm | 1
2017-08-29
11:15 | 1
5.50 cm | 1
2017-08-29
11:16 | 1
6x6
4.00 cm
0.50 cm | 1
2017-09-19
15:08 | 1
1.00 cm | 1
2017-09-19
15:10 | 1
1.50 cm | 1
2017-09-19
15:11 | 1
2.00 cm | 1
2017-09-19
15:11 | 1
2.50 cm | 1
2017-09-19
15:12 | 1
3.00 cm | 1
2017-09-19
15:12 | 1
3.50 cm | 1
2017-09-19
15:13 | 1
4.00 cm | 1
2017-09-19
15:14 | 1
4.50 cm | 1
2017-09-19
15:14 | 1
5.00 cm | 1
2017-09-19
15:15 | 1
5.50 cm | 1
2017-09-19
15:16 | 1
6.00 cm
0.50 cm | 1
2017-08-29
16:31 | 1
1.00 cm | 1
2017-08-29
16:32 | 1
1.40 cm | 1
2017-08-29
16:33 | 1
2.00 cm | 1
2017-08-29
16:34 | 1
2.50 cm | 1
2017-08-29
16:35 | 1
3.00 cm | 1
2017-08-29
16:35 | 1
3.50 cm | 1
2017-09-19
13:59 | 1
4.00 cm | 1
2017-08-29
16:36 | 1
4.50 cm | 1
2017-08-29
16:37 | 1
5.00 cm | 1
2017-08-29
16:38 | 1
5.50 cm | 1
2017-08-29
16:39 | 1
16 MeV
10x10
10.00 cm
0.50 cm | 1
2017-08-29
14:57 | 1
1.00 cm | 1
2017-08-29
14:58 | 1
2.00 cm | 1
2017-08-29
14:59 | 1
2.50 cm | 1
2017-08-29
15:00 | 1
3.00 cm | 1
2017-08-29
15:01 | 1
4.00 cm | 1
2017-08-29
15:02 | 1
4.50 cm | 1
2017-08-29
15:04 | 1
5.00 cm | 1
2017-08-29
15:05 | 1
5.50 cm | 1
2017-08-29
15:06 | 1
6.00 cm | 1
2017-08-29
15:07 | 1
6.50 cm | 1
2017-08-29
15:08 | 1
7.00 cm | 1
2017-08-29
15:09 | 1
15x15
15.00 cm
0.50 cm | 1
2017-08-29
13:42 | 1
1.00 cm | 1
2017-08-29
13:43 | 1
2.00 cm | 1
2017-08-29
13:44 | 1
2.50 cm | 1
2017-08-29
13:45 | 1
3.00 cm | 1
2017-08-29
13:46 | 1
4.00 cm | 1
2017-08-29
13:47 | 1
4.50 cm | 1
2017-08-29
13:49 | 1
5.00 cm | 1
2017-08-29
13:50 | 1
5.50 cm | 1
2017-08-29
13:51 | 1
6.00 cm | 1
2017-08-29
13:52 | 1
6.50 cm | 1
2017-08-29
13:53 | 1
7.00 cm | 1
2017-08-29
13:55 | 1
20x20
20.00 cm
0.50 cm | 1
2017-08-29
12:25 | 1
1.00 cm | 1
2017-08-29
12:26 | 1
2.00 cm | 1
2017-08-29
12:27 | 1
2.50 cm | 1
2017-08-29
12:29 | 1
3.00 cm | 1
2017-08-29
12:30 | 1
4.00 cm | 1
2017-08-29
12:31 | 1
4.50 cm | 1
2017-08-29
12:32 | 1
5.00 cm | 1
2017-08-29
12:34 | 1
5.50 cm | 1
2017-08-29
12:35 | 1
6.00 cm | 1
2017-08-29
12:36 | 1
6.50 cm | 1
2017-08-29
12:38 | 1
7.00 cm | 1
2017-08-29
12:39 | 1
25x25
25.00 cm
0.50 cm | 1
2017-08-29
10:44 | 1
1.00 cm | 1
2017-08-29
10:45 | 1
2.00 cm | 1
2017-08-29
10:47 | 1
2.50 cm | 1
2017-08-29
10:48 | 1
3.00 cm | 1
2017-08-29
10:49 | 1
4.00 cm | 1
2017-08-29
10:51 | 1
4.50 cm | 1
2017-08-29
10:52 | 1
5.00 cm | 1
2017-08-29
10:54 | 1
5.50 cm | 1
2017-08-29
10:55 | 1
6.00 cm | 1
2017-08-29
10:56 | 1
6.50 cm | 1
2017-08-29
10:58 | 1
7.00 cm | 1
2017-08-29
10:59 | 1
6x6
4.00 cm
0.50 cm | 1
2017-09-19
15:16 | 1
1.00 cm | 1
2017-09-19
15:19 | 1
2.00 cm | 1
2017-09-19
15:19 | 1
3.00 cm | 1
2017-09-19
15:20 | 1
4.00 cm | 1
2017-09-19
15:21 | 1
5.00 cm | 1
2017-09-19
15:22 | 1
6.00 cm | 1
2017-09-19
15:22 | 1
7.00 cm | 1
2017-09-19
15:23 | 1
6.00 cm
0.50 cm | 1
2017-08-29
16:14 | 1
1.00 cm | 1
2017-08-29
16:15 | 1
2.00 cm | 1
2017-08-29
16:16 | 1
2.50 cm | 1
2017-08-29
16:17 | 1
3.00 cm | 1
2017-08-29
16:18 | 1
4.00 cm | 1
2017-08-29
16:19 | 1
4.50 cm | 1
2017-08-29
16:20 | 1
5.00 cm | 1
2017-08-29
16:21 | 1
5.50 cm | 1
2017-08-29
16:22 | 1
6.00 cm | 1
2017-08-29
16:23 | 1
6.50 cm | 1
2017-08-29
16:24 | 1
7.00 cm | 1
2017-08-29
16:25 | 1
20 MeV
10x10
10.00 cm
0.50 cm | 1
2017-08-29
14:42 | 1
1.00 cm | 1
2017-08-29
14:43 | 1
2.00 cm | 1
2017-08-29
14:44 | 1
2.50 cm | 1
2017-08-29
14:45 | 1
3.00 cm | 1
2017-08-29
14:46 | 1
4.00 cm | 1
2017-08-29
14:47 | 1
5.00 cm | 1
2017-08-29
14:48 | 1
6.00 cm | 1
2017-08-29
14:50 | 1
7.00 cm | 1
2017-08-29
14:51 | 1
8.00 cm | 1
2017-08-29
14:52 | 1
9.00 cm | 1
2017-08-29
14:53 | 1
15x15
15.00 cm
0.50 cm | 1
2017-08-29
13:25 | 1
1.00 cm | 1
2017-08-29
13:26 | 1
2.00 cm | 1
2017-08-29
13:27 | 1
2.50 cm | 1
2017-08-29
13:28 | 1
3.00 cm | 1
2017-08-29
13:29 | 1
4.00 cm | 1
2017-08-29
13:30 | 1
5.00 cm | 1
2017-08-29
13:32 | 1
6.00 cm | 1
2017-08-29
13:33 | 1
7.00 cm | 1
2017-08-29
13:34 | 1
8.00 cm | 1
2017-08-29
13:35 | 1
9.00 cm | 1
2017-08-29
13:36 | 1
20x20
20.00 cm
0.50 cm | 1
2017-08-29
12:08 | 1
1.00 cm | 1
2017-08-29
12:09 | 1
2.00 cm | 1
2017-08-29
12:10 | 1
2.50 cm | 1
2017-08-29
12:12 | 1
3.00 cm | 1
2017-08-29
12:13 | 1
4.00 cm | 1
2017-08-29
12:14 | 1
5.00 cm | 1
2017-08-29
12:15 | 1
6.00 cm | 1
2017-08-29
12:16 | 1
7.00 cm | 1
2017-08-29
12:18 | 1
8.00 cm | 1
2017-08-29
12:19 | 1
9.00 cm | 1
2017-08-29
12:20 | 1
25x25
25.00 cm
0.50 cm | 1
2017-08-29
10:26 | 1
1.00 cm | 1
2017-08-29
10:27 | 1
2.00 cm | 1
2017-08-29
10:29 | 1
2.50 cm | 1
2017-08-29
10:30 | 1
3.00 cm | 1
2017-08-29
10:31 | 1
4.00 cm | 1
2017-08-29
10:33 | 1
5.00 cm | 1
2017-08-29
10:34 | 1
6.00 cm | 1
2017-08-29
10:36 | 1
7.00 cm | 1
2017-08-29
10:37 | 1
8.00 cm | 1
2017-08-29
10:38 | 1
9.00 cm | 1
2017-08-29
10:40 | 1
6x6
4.00 cm
0.50 cm | 1
2017-09-19
15:24 | 1
1.00 cm | 1
2017-09-19
15:25 | 1
2.00 cm | 1
2017-09-19
15:26 | 1
3.00 cm | 1
2017-09-19
15:26 | 1
4.00 cm | 1
2017-09-19
15:27 | 1
5.00 cm | 1
2017-09-19
15:28 | 1
6.00 cm | 1
2017-09-19
15:28 | 1
7.00 cm | 1
2017-09-19
15:29 | 1
8.00 cm | 1
2017-09-19
15:30 | 1
9.00 cm | 1
2017-09-19
15:30 | 1
6.00 cm
0.50 cm | 1
2017-08-29
16:02 | 1
1.00 cm | 1
2017-08-29
16:03 | 1
2.00 cm | 1
2017-08-29
16:04 | 1
2.50 cm | 1
2017-08-29
16:05 | 1
3.00 cm | 1
2017-08-29
16:05 | 1
4.00 cm | 1
2017-08-29
16:06 | 1
5.00 cm | 1
2017-08-29
16:07 | 1
6.00 cm | 1
2017-08-29
16:08 | 1
7.00 cm | 1
2017-08-29
16:09 | 1
8.00 cm | 1
2017-08-29
16:10 | 1
9.00 cm | 1
2017-08-29
16:11 | 1
6 MeV
10x10
10.00 cm
0.50 cm | 1
2017-09-19
12:03 | 1
0.80 cm | 1
2017-08-29
15:48 | 1
1.00 cm | 1
2017-08-29
15:49 | 1
1.40 cm | 1
2017-08-29
15:50 | 1
1.80 cm | 1
2017-09-19
12:04 | 1
2.00 cm | 1
2017-08-29
15:51 | 1
2.20 cm | 1
2017-09-19
12:05 | 1
2.50 cm | 1
2017-08-29
15:52 | 1
2.70 cm | 1
2017-09-19
12:06 | 1
3.00 cm | 1
2017-08-29
15:53 | 1
15x15
15.00 cm
0.50 cm | 1
2017-09-19
11:14 | 1
0.80 cm | 1
2017-08-29
14:27 | 1
1.00 cm | 1
2017-08-29
14:28 | 1
1.40 cm | 1
2017-08-29
14:29 | 1
1.80 cm | 1
2017-09-19
11:16 | 1
2.00 cm | 1
2017-08-29
14:31 | 1
2.20 cm | 1
2017-09-19
11:17 | 1
2.50 cm | 1
2017-08-29
14:32 | 1
2.70 cm | 1
2017-09-19
11:18 | 1
3.00 cm | 1
2017-08-29
14:33 | 1
20x20
20.00 cm
0.80 cm | 1
2017-08-29
13:10 | 1
1.00 cm | 1
2017-08-29
13:11 | 1
1.40 cm | 1
2017-08-29
13:13 | 1
2.00 cm | 1
2017-08-29
13:14 | 1
2.50 cm | 1
2017-08-29
13:15 | 1
3.00 cm | 1
2017-08-29
13:17 | 1
25x25
25.00 cm
0.80 cm | 1
2017-08-29
11:33 | 1
1.00 cm | 1
2017-08-29
11:34 | 1
1.40 cm | 1
2017-08-29
11:36 | 1
2.00 cm | 1
2017-08-29
11:37 | 1
2.50 cm | 1
2017-08-29
11:39 | 1
3.00 cm | 1
2017-08-29
11:40 | 1
6x6
4.00 cm
0.50 cm | 1
2017-09-19
14:55 | 1
0.80 cm | 1
2017-09-19
14:55 | 1
1.00 cm | 1
2017-09-19
14:56 | 1
1.40 cm | 1
2017-09-19
14:57 | 1
1.80 cm | 1
2017-09-19
14:57 | 1
2.00 cm | 1
2017-09-19
14:58 | 1
2.20 cm | 1
2017-09-19
14:59 | 1
2.50 cm | 1
2017-09-19
14:59 | 1
2.70 cm | 1
2017-09-19
15:00 | 1
6.00 cm
0.50 cm | 1
2017-09-19
13:54 | 1
0.80 cm | 1
2017-08-29
16:55 | 1
1.00 cm | 1
2017-08-29
16:56 | 1
1.40 cm | 1
2017-08-29
16:57 | 1
1.80 cm | 1
2017-09-19
13:54 | 1
2.00 cm | 1
2017-08-29
16:58 | 1
2.20 cm | 1
2017-09-19
13:55 | 1
2.50 cm | 1
2017-08-29
16:58 | 1
2.70 cm | 1
2017-09-19
13:56 | 1
3.00 cm | 1
2017-08-29
17:00 | 1
9 MeV
10x10
10.00 cm
0.50 cm | 1
2017-09-19
12:07 | 1
0.80 cm | 1
2017-08-29
15:33 | 1
1.00 cm | 1
2017-08-29
15:35 | 1
1.40 cm | 1
2017-08-29
15:36 | 1
2.00 cm | 1
2017-08-29
15:37 | 1
2.50 cm | 1
2017-08-29
15:38 | 1
3.00 cm | 1
2017-08-29
15:39 | 1
3.50 cm | 1
2017-08-29
15:40 | 1
4.00 cm | 1
2017-08-29
15:41 | 1
4.50 cm | 1
2017-09-19
12:09 | 1
15x15
15.00 cm
0.50 cm | 1
2017-09-19
11:19 | 1
0.80 cm | 1
2017-08-29
14:15 | 1
1.00 cm | 1
2017-08-29
14:17 | 1
1.40 cm | 1
2017-08-29
14:18 | 1
2.00 cm | 1
2017-08-29
14:19 | 1
2.50 cm | 1
2017-08-29
14:20 | 1
3.00 cm | 1
2017-08-29
14:21 | 1
3.50 cm | 1
2017-08-29
14:22 | 1
4.00 cm | 1
2017-08-29
14:24 | 1
4.50 cm | 1
2017-09-19
11:20 | 1
20x20
20.00 cm
0.80 cm | 1
2017-08-29
12:58 | 1
1.00 cm | 1
2017-08-29
12:59 | 1
1.40 cm | 1
2017-08-29
13:00 | 1
2.00 cm | 1
2017-08-29
13:02 | 1
2.50 cm | 1
2017-08-29
13:03 | 1
3.00 cm | 1
2017-08-29
13:04 | 1
3.50 cm | 1
2017-08-29
13:05 | 1
4.00 cm | 1
2017-08-29
13:07 | 1
25x25
25.00 cm
0.80 cm | 1
2017-08-29
11:20 | 1
1.00 cm | 1
2017-08-29
11:21 | 1
1.40 cm | 1
2017-08-29
11:23 | 1
2.00 cm | 1
2017-08-29
11:24 | 1
2.50 cm | 1
2017-08-29
11:25 | 1
3.00 cm | 1
2017-08-29
11:27 | 1
3.50 cm | 1
2017-08-29
11:28 | 1
4.00 cm | 1
2017-08-29
11:30 | 1
6x6
4.00 cm
0.50 cm | 1
2017-09-19
15:01 | 1
1.00 cm | 1
2017-09-19
15:03 | 1
1.50 cm | 1
2017-09-19
15:03 | 1
2.00 cm | 1
2017-09-19
15:04 | 1
2.50 cm | 1
2017-09-19
15:05 | 1
3.00 cm | 1
2017-09-19
15:05 | 1
3.50 cm | 1
2017-09-19
15:06 | 1
4.00 cm | 1
2017-09-19
15:07 | 1
4.50 cm | 1
2017-09-19
15:07 | 1
6.00 cm
0.50 cm | 1
2017-09-19
13:57 | 1
0.80 cm | 1
2017-08-29
16:43 | 1
1.00 cm | 1
2017-08-29
16:44 | 1
1.40 cm | 1
2017-08-29
16:45 | 1
2.00 cm | 1
2017-08-29
16:46 | 1
2.50 cm | 1
2017-08-29
16:47 | 1
3.00 cm | 1
2017-08-29
16:48 | 1
3.50 cm | 1
2017-08-29
16:49 | 1
4.00 cm | 1
2017-08-29
16:50 | 1
4.50 cm | 1
2017-09-19
13:58 | 1

## Duplicates
SSD | Energy | Applicator | Field size inline | Start position [De] | Count | SSD | Energy | Applicator | Field size inline | Start position [De] | Date | Time | Count
100.00 cm | 6 MeV | 10x10 | 10.00 cm | 0.50 cm | 2 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-08-24 | 15:17 | 1
100.00 cm | 9 MeV | 10x10 | 10.00 cm | 0.50 cm | 2 | 100.00 cm | 6 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-09-18 | 14:18 | 1
100.00 cm | 9 MeV | 10x10 | 10.00 cm | 4.50 cm | 2
100.00 cm | 12 MeV | 10x10 | 10.00 cm | 3.50 cm | 2 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-08-24 | 15:48 | 1
100.00 cm | 16 MeV | 6x6 | 4.00 cm | 6.00 cm | 3 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 0.50 cm | 2017-09-18 | 14:22 | 1
100.00 cm | 16 MeV | 6x6 | 4.00 cm | 7.00 cm | 3
100.00 cm | 16 MeV | 6x6 | 6.00 cm | 6.00 cm | 3 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 4.50 cm | 2017-08-24 | 15:57 | 1
100.00 cm | 16 MeV | 6x6 | 6.00 cm | 7.00 cm | 3 | 100.00 cm | 9 MeV | 10x10 | 10.00 cm | 4.50 cm | 2017-09-18 | 14:23 | 1
100.00 cm | 16 MeV | 10x10 | 10.00 cm | 6.00 cm | 3
100.00 cm | 16 MeV | 10x10 | 10.00 cm | 7.00 cm | 3 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 3.50 cm | 2017-08-24 | 16:28 | 1
100.00 cm | 16 MeV | 15x15 | 15.00 cm | 6.00 cm | 3 | 100.00 cm | 12 MeV | 10x10 | 10.00 cm | 3.50 cm | 2017-09-18 | 14:24 | 1
100.00 cm | 16 MeV | 15x15 | 15.00 cm | 7.00 cm | 3
110.00 cm | 16 MeV | 6x6 | 4.00 cm | 0.50 cm | 2 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 6.00 cm | 2017-08-28 | 15:48 | 1
110.00 cm | 16 MeV | 6x6 | 4.00 cm | 1.00 cm | 2 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 6.00 cm | 2017-09-18 | 15:26 | 2
110.00 cm | 16 MeV | 6x6 | 4.00 cm | 2.00 cm | 2
110.00 cm | 16 MeV | 6x6 | 4.00 cm | 3.00 cm | 2 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 7.00 cm | 2017-08-28 | 15:50 | 1
110.00 cm | 16 MeV | 6x6 | 4.00 cm | 4.00 cm | 2 | 100.00 cm | 16 MeV | 6x6 | 4.00 cm | 7.00 cm | 2017-09-18 | 15:28 | 2
110.00 cm | 16 MeV | 6x6 | 4.00 cm | 5.00 cm | 2
110.00 cm | 16 MeV | 6x6 | 4.00 cm | 6.00 cm | 2 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 6.00 cm | 2017-08-28 | 15:02 | 1
110.00 cm | 16 MeV | 6x6 | 4.00 cm | 7.00 cm | 2 | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 6.00 cm | 2017-09-18 | 15:44 | 2
[6 empty] | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 7.00 cm | 2017-08-28 | 15:03 | 1
[6 empty] | 100.00 cm | 16 MeV | 6x6 | 6.00 cm | 7.00 cm | 2017-09-18 | 15:46 | 2
[6 empty] | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 6.00 cm | 2017-08-24 | 17:07 | 1
[6 empty] | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 6.00 cm | 2017-09-18 | 14:26 | 2
[6 empty] | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 7.00 cm | 2017-08-24 | 17:08 | 1
[6 empty] | 100.00 cm | 16 MeV | 10x10 | 10.00 cm | 7.00 cm | 2017-09-18 | 14:28 | 2
[6 empty] | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 6.00 cm | 2017-08-28 | 12:43 | 1
[6 empty] | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 6.00 cm | 2017-09-18 | 12:18 | 2
[6 empty] | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 7.00 cm | 2017-08-28 | 12:45 | 1
[6 empty] | 100.00 cm | 16 MeV | 15x15 | 15.00 cm | 7.00 cm | 2017-09-18 | 12:20 | 1
[6 empty] | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 0.50 cm | 2017-09-19 | 15:16 | 2
[6 empty] | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 1.00 cm | 2017-09-19 | 15:19 | 2
[6 empty] | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 2.00 cm | 2017-09-19 | 15:19 | 2
[6 empty] | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 3.00 cm | 2017-09-19 | 15:20 | 2
[6 empty] | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 4.00 cm | 2017-09-19 | 15:21 | 2
[6 empty] | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 5.00 cm | 2017-09-19 | 15:22 | 2
[6 empty] | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 6.00 cm | 2017-09-19 | 15:22 | 2
[6 empty] | 110.00 cm | 16 MeV | 6x6 | 4.00 cm | 7.00 cm | 2017-09-19 | 15:23 | 2