## Scan Files
Q:\Treatment machines including photos\LINACS\Truebeam 2\Commissioning\Omnipro data\TR2 commissioning\Electron scans | [16 empty] | Scan Files
| Energy | Applicator | Field Size | SSD | Scan Date | File Name | Modified Date | Modified Time | File Size | Content | Extension | Energy | Applicator | Field Size | SSD | Scan Date | Energy | Applicator | Field Size | SSD | Scan Date | File Name | Modified Date
Volume in drive Q is 158 | 12 | 10 | 10x10 | 110 | 29/08/2017 | 12MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017 | 4298500% | 17:37 | 335987 | PDD&Profiles | opab | 12MeV | 10x10app | 10x10field | 110SSD | 29082017 | 12 | 10 | 10x10 | 110 | 29/08/2017 | 12MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017 | 4298500%
Volume Serial Number is 0000-0023 | 12 | 15 | 15x15 | 110 | 29/08/2017 | 12MeV_15x15app_15x15field_PDD&Profiles_SSD110_29082017 | 4297600% | 14:13 | 390793 | PDD&Profiles | opab | 12MeV | 15x15app | 15x15field | 110SSD | 29082017 | 12 | 15 | 15x15 | 110 | 29/08/2017 | 12MeV_15x15app_15x15field_PDD&Profiles_SSD110_29082017 | 4297600%
| 12 | 20 | 20x20 | 110 | 29/08/2017 | 12MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017 | 4297600% | 12:56 | 446253 | PDD&Profiles | opab | 12MeV | 20x20app | 20x20field | 110SSD | 29082017 | 12 | 20 | 20x20 | 110 | 29/08/2017 | 12MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017 | 4297600%
Directory of Q:\Treatment machines including photos\LINACS\Truebeam 2\Commissioning\Omnipro data\TR2 commissioning\Electron scans\6MeV | 12 | 25 | 25x25 | 110 | 29/08/2017 | 12MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017 | 4297600% | 11:18 | 501477 | PDD&Profiles | opab | 12MeV | 25x25app | 25x25field | 110SSD | 29082017 | 12 | 25 | 25x25 | 110 | 29/08/2017 | 12MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017 | 4297600%
| 12 | 6 | 6x6 | 110 | 29/08/2017 | 12MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017 | 4297600% | 16:41 | 291835 | PDD&Profiles | opab | 12MeV | 6x6app | 6x6field | 110SSD | 29082017 | 12 | 6 | 6x6 | 110 | 29/08/2017 | 12MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017 | 4297600%
10/10/2017 01:56 PM <DIR> . | 9 | 10 | 10x10 | 110 | 29/08/2017 | 9MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017 | 4298500% | 17:38 | 269864 | PDD&Profiles | opab | 9MeV | 10x10app | 10x10field | 110SSD | 29082017 | 9 | 10 | 10x10 | 110 | 29/08/2017 | 9MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017 | 4298500%
26/09/2017 08:59 AM <DIR> .. | 9 | 15 | 15x15 | 110 | 29/08/2017 | 9MeV_15x15app_15x15field_PDD&Profiles_SSD110_29082017 | 4297600% | 14:25 | 313813 | PDD&Profiles | opab | 9MeV | 15x15app | 15x15field | 110SSD | 29082017 | 9 | 15 | 15x15 | 110 | 29/08/2017 | 9MeV_15x15app_15x15field_PDD&Profiles_SSD110_29082017 | 4297600%
07/09/2017 05:38 PM 231,895 6MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017.opab | 9 | 20 | 20x20 | 110 | 29/08/2017 | 9MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017 | 4297600% | 13:09 | 357473 | PDD&Profiles | opab | 9MeV | 20x20app | 20x20field | 110SSD | 29082017 | 9 | 20 | 20x20 | 110 | 29/08/2017 | 9MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017 | 4297600%
25/08/2017 10:12 AM 885,864 6MeV_10x10app_10x10field_PDD&Profiles_24082017.opab | 9 | 25 | 25x25 | 110 | 29/08/2017 | 9MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017 | 4298500% | 17:38 | 401552 | PDD&Profiles | opab | 9MeV | 25x25app | 25x25field | 110SSD | 29082017 | 9 | 25 | 25x25 | 110 | 29/08/2017 | 9MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017 | 4298500%
28/08/2017 12:15 PM 158,583 6MeV_10x10app_3x3field_PDD&Profiles_28082017.opab | 9 | 6 | 6x6 | 110 | 29/08/2017 | 9MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017 | 4297600% | 16:53 | 234691 | PDD&Profiles | opab | 9MeV | 6x6app | 6x6field | 110SSD | 29082017 | 9 | 6 | 6x6 | 110 | 29/08/2017 | 9MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017 | 4297600%
25/08/2017 04:26 PM 182,419 6MeV_10x10app_7x7field_PDD&Profiles_25082017.opab | 16 | 10 | 10x10 | 110 | 29/08/2017 | 16MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017 | 4297600% | 15:13 | 401809 | PDD&Profiles | opab | 16MeV | 10x10app | 10x10field | 110SSD | 29082017 | 16 | 10 | 10x10 | 110 | 29/08/2017 | 16MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017 | 4297600%
28/08/2017 02:36 PM 206,377 6MeV_15x15app_11x11field_PDD&Profiles_28082017.opab | 16 | 15 | 15x15 | 110 | 29/08/2017 | 16MeV_15x15app_15x15field_PDD&Profiles_110SSD_29082017 | 4297600% | 13:58 | 468597 | PDD&Profiles | opab | 16MeV | 15x15app | 15x15field | 110SSD | 29082017 | 16 | 15 | 15x15 | 110 | 29/08/2017 | 16MeV_15x15app_15x15field_PDD&Profiles_110SSD_29082017 | 4297600%
29/08/2017 02:35 PM 238,723 6MeV_15x15app_15x15field_PDD&Profiles_110SSD_29082017.opab | 16 | 20 | 20x20 | 110 | 29/08/2017 | 16MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017 | 4297600% | 12:41 | 535385 | PDD&Profiles | opab | 16MeV | 20x20app | 20x20field | 110SSD | 29082017 | 16 | 20 | 20x20 | 110 | 29/08/2017 | 16MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017 | 4297600%
28/08/2017 01:24 PM 229,977 6MeV_15x15app_15x15field_PDD&Profiles_28082017.opab | 16 | 25 | 25x25 | 110 | 29/08/2017 | 16MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017 | 4297600% | 11:01 | 602409 | PDD&Profiles | opab | 16MeV | 25x25app | 25x25field | 110SSD | 29082017 | 16 | 25 | 25x25 | 110 | 29/08/2017 | 16MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017 | 4297600%
28/08/2017 11:23 AM 259,831 6MeV_20x20app_20x20field_PDD&Profiles_28082017.opab | 16 | 6 | 6x6 | 110 | 29/08/2017 | 16MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017 | 4297600% | 16:27 | 347797 | PDD&Profiles | opab | 16MeV | 6x6app | 6x6field | 110SSD | 29082017 | 16 | 6 | 6x6 | 110 | 29/08/2017 | 16MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017 | 4297600%
29/08/2017 01:18 PM 271,527 6MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017.opab | 20 | 10 | 10x10 | 110 | 29/08/2017 | 20MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017 | 4297600% | 14:54 | 372642 | PDD&Profiles | opab | 20MeV | 10x10app | 10x10field | 110SSD | 29082017 | 20 | 10 | 10x10 | 110 | 29/08/2017 | 20MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017 | 4297600%
25/08/2017 02:12 PM 265,731 6MeV_25x25app_21x21field_PDD&Profiles_25082017.opab | 20 | 15 | 15x15 | 110 | 29/08/2017 | 20MeV_15x15app_15x15field_PDD&Profiles_SSD110_29082017 | 4297600% | 13:38 | 434238 | PDD&Profiles | opab | 20MeV | 15x15app | 15x15field | 110SSD | 29082017 | 20 | 15 | 15x15 | 110 | 29/08/2017 | 20MeV_15x15app_15x15field_PDD&Profiles_SSD110_29082017 | 4297600%
13/09/2017 10:17 AM 4,075,192 All 6MeV Data.opab | 20 | 20 | 20x20 | 110 | 29/08/2017 | 20MeV_20x20app_20x20field_PDD&Profiles_110SSD_29082017 | 4297600% | 12:22 | 495834 | PDD&Profiles | opab | 20MeV | 20x20app | 20x20field | 110SSD | 29082017 | 20 | 20 | 20x20 | 110 | 29/08/2017 | 20MeV_20x20app_20x20field_PDD&Profiles_110SSD_29082017 | 4297600%
07/09/2017 05:38 PM 289,986 6MeV_25x25app_25x25field_PDD&Profiles_25082017.opab | 20 | 25 | 25x25 | 110 | 29/08/2017 | 20MeV_25x25app_25x25field_PDD&Profiles_110SSD_29082017 | 4297600% | 10:42 | 557312 | PDD&Profiles | opab | 20MeV | 25x25app | 25x25field | 110SSD | 29082017 | 20 | 25 | 25x25 | 110 | 29/08/2017 | 20MeV_25x25app_25x25field_PDD&Profiles_110SSD_29082017 | 4297600%
29/08/2017 11:42 AM 304,567 6MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017.opab | 20 | 6 | 6x6 | 110 | 29/08/2017 | 20MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017 | 4297600% | 16:13 | 323356 | PDD&Profiles | opab | 20MeV | 6x6app | 6x6field | 110SSD | 29082017 | 20 | 6 | 6x6 | 110 | 29/08/2017 | 20MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017 | 4297600%
07/09/2017 05:38 PM 158,602 6MeV_6x6app_3x3field_PDD&Profiles_28082017.opab | 6 | 10 | 10x10 | 110 | 29/08/2017 | 6MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017 | 4298500% | 17:38 | 231895 | PDD&Profiles | opab | 6MeV | 10x10app | 10x10field | 110SSD | 29082017 | 6 | 10 | 10x10 | 110 | 29/08/2017 | 6MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017 | 4298500%
11/09/2017 05:41 PM 164,266 6MeV_6x6app_4x4field_PDD&Profiles_28082017.opab | 6 | 15 | 15x15 | 110 | 29/08/2017 | 6MeV_15x15app_15x15field_PDD&Profiles_110SSD_29082017 | 4297600% | 14:35 | 238723 | PDD&Profiles | opab | 6MeV | 15x15app | 15x15field | 110SSD | 29082017 | 6 | 15 | 15x15 | 110 | 29/08/2017 | 6MeV_15x15app_15x15field_PDD&Profiles_110SSD_29082017 | 4297600%
29/08/2017 05:02 PM 179,909 6MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017.opab | 6 | 20 | 20x20 | 110 | 29/08/2017 | 6MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017 | 4297600% | 13:18 | 271527 | PDD&Profiles | opab | 6MeV | 20x20app | 20x20field | 110SSD | 29082017 | 6 | 20 | 20x20 | 110 | 29/08/2017 | 6MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017 | 4297600%
07/09/2017 05:38 PM 176,538 6MeV_6x6app_6x6field_PDD&Profiles_28082017.opab | 6 | 25 | 25x25 | 110 | 29/08/2017 | 6MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017 | 4297600% | 11:42 | 304567 | PDD&Profiles | opab | 6MeV | 25x25app | 25x25field | 110SSD | 29082017 | 6 | 25 | 25x25 | 110 | 29/08/2017 | 6MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017 | 4297600%
10/10/2017 01:56 PM 204 files.txt | 6 | 6 | 6x6 | 110 | 29/08/2017 | 6MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017 | 4297600% | 17:02 | 179909 | PDD&Profiles | opab | 6MeV | 6x6app | 6x6field | 110SSD | 29082017 | 6 | 6 | 6x6 | 110 | 29/08/2017 | 6MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017 | 4297600%
18 File(s) 8,280,191 bytes | 20 | 10 | 10x10 | 100 | 25/08/2017 | 20MeV_10x10app_10x10field_PDD&Profiles_25082017 | 4297200% | 14:53 | 1002020 | PDD&Profiles | opab | 20MeV | 10x10app | 10x10field | 100SSD | 25082017 | 20 | 10 | 10x10 | 100 | 25/08/2017 | 20MeV_10x10app_10x10field_PDD&Profiles_25082017 | 4297200%
2 Dir(s) 834,279,194,624 bytes free | 12 | 10 | 10x10 | 100 | 24/08/2017 | 12MeV_10x10app_10x10field_PDD&Profiles_24082017 | 4297100% | 16:44 | 781848 | PDD&Profiles | opab | 12MeV | 10x10app | 10x10field | 100SSD | 24082017 | 12 | 10 | 10x10 | 100 | 24/08/2017 | 12MeV_10x10app_10x10field_PDD&Profiles_24082017 | 4297100%
| 12 | 10 | 3x3 | 100 | 28/08/2017 | 12MeV_10x10app_3x3field_PDD&Profiles_28082017 | 4297500% | 12:01 | 255665 | PDD&Profiles | opab | 12MeV | 10x10app | 3x3field | 100SSD | 28082017 | 12 | 10 | 3x3 | 100 | 28/08/2017 | 12MeV_10x10app_3x3field_PDD&Profiles_28082017 | 4297500%
| 12 | 10 | 7x7 | 100 | 25/08/2017 | 12MeV_10x10app_7x7field_PDD&Profiles_25082017 | 4297200% | 15:55 | 295903 | PDD&Profiles | opab | 12MeV | 10x10app | 7x7field | 100SSD | 25082017 | 12 | 10 | 7x7 | 100 | 25/08/2017 | 12MeV_10x10app_7x7field_PDD&Profiles_25082017 | 4297200%
Volume Serial Number is 0000-0023 | 12 | 15 | 11x11 | 100 | 28/08/2017 | 12MeV_15x15app_11x11field_PDD&Profiles_28082017 | 4297500% | 14:15 | 336381 | PDD&Profiles | opab | 12MeV | 15x15app | 11x11field | 100SSD | 28082017 | 12 | 15 | 11x11 | 100 | 28/08/2017 | 12MeV_15x15app_11x11field_PDD&Profiles_28082017 | 4297500%
| 12 | 15 | 15x15 | 100 | 28/08/2017 | 12MeV_15x15app_15x15field_PDD&Profiles_28082017 | 4297500% | 13:01 | 376383 | PDD&Profiles | opab | 12MeV | 15x15app | 15x15field | 100SSD | 28082017 | 12 | 15 | 15x15 | 100 | 28/08/2017 | 12MeV_15x15app_15x15field_PDD&Profiles_28082017 | 4297500%
Directory of Q:\Treatment machines including photos\LINACS\Truebeam 2\Commissioning\Omnipro data\TR2 commissioning\Electron scans\12MeV | 12 | 20 | 20x20 | 100 | 28/08/2017 | 12MeV_20x20app_20x20field_PDD&Profiles_28082017 | 4297500% | 10:56 | 426890 | PDD&Profiles | opab | 12MeV | 20x20app | 20x20field | 100SSD | 28082017 | 12 | 20 | 20x20 | 100 | 28/08/2017 | 12MeV_20x20app_20x20field_PDD&Profiles_28082017 | 4297500%
| 12 | 25 | 21x21 | 100 | 25/08/2017 | 12MeV_25x25app_21x21field_PDD&Profiles_25082017 | 4297200% | 13:49 | 436681 | PDD&Profiles | opab | 12MeV | 25x25app | 21x21field | 100SSD | 25082017 | 12 | 25 | 21x21 | 100 | 25/08/2017 | 12MeV_25x25app_21x21field_PDD&Profiles_25082017 | 4297200%
11/10/2017 09:17 AM <DIR> . | 12 | 25 | 25x25 | 100 | 25/08/2017 | 12MeV_25x25app_25x25field_PDD&Profiles_25082017 | 4297200% | 12:08 | 477155 | PDD&Profiles | opab | 12MeV | 25x25app | 25x25field | 100SSD | 25082017 | 12 | 25 | 25x25 | 100 | 25/08/2017 | 12MeV_25x25app_25x25field_PDD&Profiles_25082017 | 4297200%
26/09/2017 08:59 AM <DIR> .. | 12 | 6 | 3x3 | 100 | 28/08/2017 | 12MeV_6x6app_3x3field_PDD&Profiles_28082017 | 4298500% | 17:38 | 255778 | PDD&Profiles | opab | 12MeV | 6x6app | 3x3field | 100SSD | 28082017 | 12 | 6 | 3x3 | 100 | 28/08/2017 | 12MeV_6x6app_3x3field_PDD&Profiles_28082017 | 4298500%
07/09/2017 05:37 PM 335,987 12MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017.opab | 12 | 6 | 4x4 | 100 | 28/08/2017 | 12MeV_6x6app_4x4field_PDD&Profiles_28082017 | 4297500% | 16:02 | 265271 | PDD&Profiles | opab | 12MeV | 6x6app | 4x4field | 100SSD | 28082017 | 12 | 6 | 4x4 | 100 | 28/08/2017 | 12MeV_6x6app_4x4field_PDD&Profiles_28082017 | 4297500%
24/08/2017 04:44 PM 781,848 12MeV_10x10app_10x10field_PDD&Profiles_24082017.opab | 12 | 6 | 6x6 | 100 | 28/08/2017 | 12MeV_6x6app_6x6field_PDD&Profiles_28082017 | 4297500% | 15:14 | 285921 | PDD&Profiles | opab | 12MeV | 6x6app | 6x6field | 100SSD | 28082017 | 12 | 6 | 6x6 | 100 | 28/08/2017 | 12MeV_6x6app_6x6field_PDD&Profiles_28082017 | 4297500%
28/08/2017 12:01 PM 255,665 12MeV_10x10app_3x3field_PDD&Profiles_28082017.opab | 9 | 10 | 10x10 | 100 | 24/08/2017 | 9MeV_10x10app_10x10field_PDD&Profiles_24082017 | 4297100% | 16:08 | 627678 | PDD&Profiles | opab | 9MeV | 10x10app | 10x10field | 100SSD | 24082017 | 9 | 10 | 10x10 | 100 | 24/08/2017 | 9MeV_10x10app_10x10field_PDD&Profiles_24082017 | 4297100%
25/08/2017 03:55 PM 295,903 12MeV_10x10app_7x7field_PDD&Profiles_25082017.opab | 9 | 10 | 3x3 | 100 | 28/08/2017 | 9MeV_10x10app_3x3field_PDD&Profiles_28082017 | 4297500% | 12:08 | 205943 | PDD&Profiles | opab | 9MeV | 10x10app | 3x3field | 100SSD | 28082017 | 9 | 10 | 3x3 | 100 | 28/08/2017 | 9MeV_10x10app_3x3field_PDD&Profiles_28082017 | 4297500%
28/08/2017 02:15 PM 336,381 12MeV_15x15app_11x11field_PDD&Profiles_28082017.opab | 9 | 10 | 7x7 | 100 | 25/08/2017 | 9MeV_10x10app_7x7field_PDD&Profiles_25082017 | 4297200% | 16:10 | 238039 | PDD&Profiles | opab | 9MeV | 10x10app | 7x7field | 100SSD | 25082017 | 9 | 10 | 7x7 | 100 | 25/08/2017 | 9MeV_10x10app_7x7field_PDD&Profiles_25082017 | 4297200%
28/08/2017 01:01 PM 376,383 12MeV_15x15app_15x15field_PDD&Profiles_28082017.opab | 9 | 15 | 11x11 | 100 | 28/08/2017 | 9MeV_15x15app_11x11field_PDD&Profiles_28082017 | 4297500% | 14:26 | 270021 | PDD&Profiles | opab | 9MeV | 15x15app | 11x11field | 100SSD | 28082017 | 9 | 15 | 11x11 | 100 | 28/08/2017 | 9MeV_15x15app_11x11field_PDD&Profiles_28082017 | 4297500%
29/08/2017 02:13 PM 390,793 12MeV_15x15app_15x15field_PDD&Profiles_SSD110_29082017.opab | 9 | 15 | 15x15 | 100 | 28/08/2017 | 9MeV_15x15app_15x15field_PDD&Profiles_28082017 | 4297500% | 13:12 | 301881 | PDD&Profiles | opab | 9MeV | 15x15app | 15x15field | 100SSD | 28082017 | 9 | 15 | 15x15 | 100 | 28/08/2017 | 9MeV_15x15app_15x15field_PDD&Profiles_28082017 | 4297500%
28/08/2017 10:56 AM 426,890 12MeV_20x20app_20x20field_PDD&Profiles_28082017.opab | 9 | 20 | 20x20 | 100 | 28/08/2017 | 9MeV_20x20app_20x20field_PDD&Profiles_28082017 | 4297500% | 11:09 | 342121 | PDD&Profiles | opab | 9MeV | 20x20app | 20x20field | 100SSD | 28082017 | 9 | 20 | 20x20 | 100 | 28/08/2017 | 9MeV_20x20app_20x20field_PDD&Profiles_28082017 | 4297500%
29/08/2017 12:56 PM 446,253 12MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017.opab | 9 | 25 | 21x21 | 100 | 25/08/2017 | 9MeV_25x25app_21x21field_PDD&Profiles_25082017 | 4297200% | 14:03 | 350969 | PDD&Profiles | opab | 9MeV | 25x25app | 21x21field | 100SSD | 25082017 | 9 | 25 | 21x21 | 100 | 25/08/2017 | 9MeV_25x25app_21x21field_PDD&Profiles_25082017 | 4297200%
25/08/2017 01:49 PM 436,681 12MeV_25x25app_21x21field_PDD&Profiles_25082017.opab | 9 | 25 | 25x25 | 100 | 25/08/2017 | 9MeV_25x25app_25x25field_PDD&Profiles_25082017 | 4297200% | 12:22 | 382121 | PDD&Profiles | opab | 9MeV | 25x25app | 25x25field | 100SSD | 25082017 | 9 | 25 | 25x25 | 100 | 25/08/2017 | 9MeV_25x25app_25x25field_PDD&Profiles_25082017 | 4297200%
11/10/2017 09:17 AM 205 files.txt | 9 | 6 | 3x3 | 100 | 28/08/2017 | 9MeV_6x6app_3x3field_PDD&Profiles_28082017 | 4298500% | 17:38 | 205950 | PDD&Profiles | opab | 9MeV | 6x6app | 3x3field | 100SSD | 28082017 | 9 | 6 | 3x3 | 100 | 28/08/2017 | 9MeV_6x6app_3x3field_PDD&Profiles_28082017 | 4298500%
25/08/2017 12:08 PM 477,155 12MeV_25x25app_25x25field_PDD&Profiles_25082017.opab | 9 | 6 | 4x4 | 100 | 28/08/2017 | 9MeV_6x6app_4x4field_PDD&Profiles_28082017 | 4297500% | 16:10 | 213909 | PDD&Profiles | opab | 9MeV | 6x6app | 4x4field | 100SSD | 28082017 | 9 | 6 | 4x4 | 100 | 28/08/2017 | 9MeV_6x6app_4x4field_PDD&Profiles_28082017 | 4297500%
29/08/2017 11:18 AM 501,477 12MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017.opab | 9 | 6 | 6x6 | 100 | 28/08/2017 | 9MeV_6x6app_6x6field_PDD&Profiles_28082017 | 4297500% | 15:21 | 229839 | PDD&Profiles | opab | 9MeV | 6x6app | 6x6field | 100SSD | 28082017 | 9 | 6 | 6x6 | 100 | 28/08/2017 | 9MeV_6x6app_6x6field_PDD&Profiles_28082017 | 4297500%
07/09/2017 05:38 PM 255,778 12MeV_6x6app_3x3field_PDD&Profiles_28082017.opab | 16 | 10 | 10x10 | 100 | 24/08/2017 | 16MeV_10x10app_10x10field_PDD&Profiles_24082017 | 4297100% | 17:26 | 884138 | PDD&Profiles | opab | 16MeV | 10x10app | 10x10field | 100SSD | 24082017 | 16 | 10 | 10x10 | 100 | 24/08/2017 | 16MeV_10x10app_10x10field_PDD&Profiles_24082017 | 4297100%
28/08/2017 04:02 PM 265,271 12MeV_6x6app_4x4field_PDD&Profiles_28082017.opab | 16 | 10 | 3x3 | 100 | 28/08/2017 | 16MeV_10x10app_3x3field_PDD&Profiles_28082017 | 4297500% | 11:52 | 304795 | PDD&Profiles | opab | 16MeV | 10x10app | 3x3field | 100SSD | 28082017 | 16 | 10 | 3x3 | 100 | 28/08/2017 | 16MeV_10x10app_3x3field_PDD&Profiles_28082017 | 4297500%
29/08/2017 04:41 PM 291,835 12MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017.opab | 16 | 10 | 7x7 | 100 | 25/08/2017 | 16MeV_10x10app_7x7field_PDD&Profiles_25082017 | 4297200% | 15:41 | 353647 | PDD&Profiles | opab | 16MeV | 10x10app | 7x7field | 100SSD | 25082017 | 16 | 10 | 7x7 | 100 | 25/08/2017 | 16MeV_10x10app_7x7field_PDD&Profiles_25082017 | 4297200%
28/08/2017 03:14 PM 285,921 12MeV_6x6app_6x6field_PDD&Profiles_28082017.opab | 16 | 15 | 11x11 | 100 | 28/08/2017 | 16MeV_15x15app_11x11field_PDD&Profiles_28082017 | 4297500% | 14:02 | 402267 | PDD&Profiles | opab | 16MeV | 15x15app | 11x11field | 100SSD | 28082017 | 16 | 15 | 11x11 | 100 | 28/08/2017 | 16MeV_15x15app_11x11field_PDD&Profiles_28082017 | 4297500%
17 File(s) 6,160,426 bytes | 16 | 15 | 15x15 | 100 | 28/08/2017 | 16MeV_15x15app_15x15field_PDD&Profiles_28082017 | 4297500% | 12:47 | 451119 | PDD&Profiles | opab | 16MeV | 15x15app | 15x15field | 100SSD | 28082017 | 16 | 15 | 15x15 | 100 | 28/08/2017 | 16MeV_15x15app_15x15field_PDD&Profiles_28082017 | 4297500%
2 Dir(s) 833,814,085,632 bytes free | 16 | 20 | 20x20 | 100 | 28/08/2017 | 16MeV_20x20app_20x20field_PDD&Profiles_28082017 | 4297500% | 10:42 | 511774 | PDD&Profiles | opab | 16MeV | 20x20app | 20x20field | 100SSD | 28082017 | 16 | 20 | 20x20 | 100 | 28/08/2017 | 16MeV_20x20app_20x20field_PDD&Profiles_28082017 | 4297500%
| 16 | 25 | 21x21 | 100 | 25/08/2017 | 16MeV_25x25app_21x21field_PDD&Profiles_25082017 | 4297200% | 13:16 | 524043 | PDD&Profiles | opab | 16MeV | 25x25app | 21x21field | 100SSD | 25082017 | 16 | 25 | 21x21 | 100 | 25/08/2017 | 16MeV_25x25app_21x21field_PDD&Profiles_25082017 | 4297200%
| 16 | 25 | 25x25 | 100 | 25/08/2017 | 16MeV_25x25app_25x25field_PDD&Profiles_25082017 | 4297200% | 11:49 | 572659 | PDD&Profiles | opab | 16MeV | 25x25app | 25x25field | 100SSD | 25082017 | 16 | 25 | 25x25 | 100 | 25/08/2017 | 16MeV_25x25app_25x25field_PDD&Profiles_25082017 | 4297200%
| 16 | 6 | 3x3 | 100 | 28/08/2017 | 16MeV_6x6app_3x3field_PDD&Profiles_28082017 | 4297500% | 16:39 | 304831 | PDD&Profiles | opab | 16MeV | 6x6app | 3x3field | 100SSD | 28082017 | 16 | 6 | 3x3 | 100 | 28/08/2017 | 16MeV_6x6app_3x3field_PDD&Profiles_28082017 | 4297500%
| 16 | 6 | 4x4 | 100 | 28/08/2017 | 16MeV_6x6app_4x4field_PDD&Profiles_28082017 | 4297500% | 15:51 | 316395 | PDD&Profiles | opab | 16MeV | 6x6app | 4x4field | 100SSD | 28082017 | 16 | 6 | 4x4 | 100 | 28/08/2017 | 16MeV_6x6app_4x4field_PDD&Profiles_28082017 | 4297500%
Volume in drive Q is 158 | 16 | 6 | 6x6 | 100 | 28/08/2017 | 16MeV_6x6app_6x6field_PDD&Profiles_28082017 | 4297500% | 15:05 | 340939 | PDD&Profiles | opab | 16MeV | 6x6app | 6x6field | 100SSD | 28082017 | 16 | 6 | 6x6 | 100 | 28/08/2017 | 16MeV_6x6app_6x6field_PDD&Profiles_28082017 | 4297500%
Volume Serial Number is 0000-0023 | 20 | 10 | 3x3 | 100 | 28/08/2017 | 20MeV_10x10app_3x3field_PDD&Profiles_28082017 | 4297500% | 11:39 | 283416 | PDD&Profiles | opab | 20MeV | 10x10app | 3x3field | 100SSD | 28082017 | 20 | 10 | 3x3 | 100 | 28/08/2017 | 20MeV_10x10app_3x3field_PDD&Profiles_28082017 | 4297500%
| 20 | 10 | 7x7 | 100 | 25/08/2017 | 20MeV_10x10app_7x7field_PDD&Profiles_25082017 | 4297200% | 15:13 | 328256 | PDD&Profiles | opab | 20MeV | 10x10app | 7x7field | 100SSD | 25082017 | 20 | 10 | 7x7 | 100 | 25/08/2017 | 20MeV_10x10app_7x7field_PDD&Profiles_25082017 | 4297200%
Directory of Q:\Treatment machines including photos\LINACS\Truebeam 2\Commissioning\Omnipro data\TR2 commissioning\Electron scans\9MeV | 20 | 15 | 11x11 | 100 | 28/08/2017 | 20MeV_15x15app_11x11field_PDD&Profiles_28082017 | 4297500% | 13:47 | 372982 | PDD&Profiles | opab | 20MeV | 15x15app | 11x11field | 100SSD | 28082017 | 20 | 15 | 11x11 | 100 | 28/08/2017 | 20MeV_15x15app_11x11field_PDD&Profiles_28082017 | 4297500%
| 20 | 15 | 15x15 | 100 | 28/08/2017 | 20MeV_15x15app_15x15field_PDD&Profiles_28082017 | 4297500% | 12:32 | 417704 | PDD&Profiles | opab | 20MeV | 15x15app | 15x15field | 100SSD | 28082017 | 20 | 15 | 15x15 | 100 | 28/08/2017 | 20MeV_15x15app_15x15field_PDD&Profiles_28082017 | 4297500%
11/10/2017 09:18 AM <DIR> . | 20 | 20 | 20x20 | 100 | 28/08/2017 | 20MeV_20x20app_20x20field_PDD&Profiles_28082017 | 4297500% | 10:25 | 473993 | PDD&Profiles | opab | 20MeV | 20x20app | 20x20field | 100SSD | 28082017 | 20 | 20 | 20x20 | 100 | 28/08/2017 | 20MeV_20x20app_20x20field_PDD&Profiles_28082017 | 4297500%
26/09/2017 08:59 AM <DIR> .. | 20 | 25 | 21x21 | 100 | 25/08/2017 | 20MeV_25x25app_21x21field_PDD&Profiles_25082017 | 4297200% | 12:59 | 485436 | PDD&Profiles | opab | 20MeV | 25x25app | 21x21field | 100SSD | 25082017 | 20 | 25 | 21x21 | 100 | 25/08/2017 | 20MeV_25x25app_21x21field_PDD&Profiles_25082017 | 4297200%
07/09/2017 05:38 PM 269,864 9MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017.opab | 20 | 25 | 25x25 | 100 | 25/08/2017 | 20MeV_25x25app_25x25field_PDD&Profiles_25082017 | 4297200% | 11:28 | 530158 | PDD&Profiles | opab | 20MeV | 25x25app | 25x25field | 100SSD | 25082017 | 20 | 25 | 25x25 | 100 | 25/08/2017 | 20MeV_25x25app_25x25field_PDD&Profiles_25082017 | 4297200%
24/08/2017 04:08 PM 627,678 9MeV_10x10app_10x10field_PDD&Profiles_24082017.opab | 20 | 6 | 3x3 | 100 | 28/08/2017 | 20MeV_6x6app_3x3field_PDD&Profiles_28082017 | 4297500% | 16:28 | 283458 | PDD&Profiles | opab | 20MeV | 6x6app | 3x3field | 100SSD | 28082017 | 20 | 6 | 3x3 | 100 | 28/08/2017 | 20MeV_6x6app_3x3field_PDD&Profiles_28082017 | 4297500%
28/08/2017 12:08 PM 205,943 9MeV_10x10app_3x3field_PDD&Profiles_28082017.opab | 20 | 6 | 4x4 | 100 | 28/08/2017 | 20MeV_6x6app_4x4field_PDD&Profiles_28082017 | 4297500% | 15:40 | 294432 | PDD&Profiles | opab | 20MeV | 6x6app | 4x4field | 100SSD | 28082017 | 20 | 6 | 4x4 | 100 | 28/08/2017 | 20MeV_6x6app_4x4field_PDD&Profiles_28082017 | 4297500%
25/08/2017 04:10 PM 238,039 9MeV_10x10app_7x7field_PDD&Profiles_25082017.opab | 20 | 6 | 6x6 | 100 | 28/08/2017 | 20MeV_6x6app_6x6field_PDD&Profiles_28082017 | 4297500% | 14:54 | 316616 | PDD&Profiles | opab | 20MeV | 6x6app | 6x6field | 100SSD | 28082017 | 20 | 6 | 6x6 | 100 | 28/08/2017 | 20MeV_6x6app_6x6field_PDD&Profiles_28082017 | 4297500%
28/08/2017 02:26 PM 270,021 9MeV_15x15app_11x11field_PDD&Profiles_28082017.opab | 6 | 10 | 10x10 | 100 | 24/08/2017 | 6MeV_10x10app_10x10field_PDD&Profiles_24082017 | 4297200% | 10:12 | 885864 | PDD&Profiles | opab | 6MeV | 10x10app | 10x10field | 100SSD | 24082017 | 6 | 10 | 10x10 | 100 | 24/08/2017 | 6MeV_10x10app_10x10field_PDD&Profiles_24082017 | 4297200%
28/08/2017 01:12 PM 301,881 9MeV_15x15app_15x15field_PDD&Profiles_28082017.opab | 6 | 10 | 3x3 | 100 | 28/08/2017 | 6MeV_10x10app_3x3field_PDD&Profiles_28082017 | 4297500% | 12:15 | 158583 | PDD&Profiles | opab | 6MeV | 10x10app | 3x3field | 100SSD | 28082017 | 6 | 10 | 3x3 | 100 | 28/08/2017 | 6MeV_10x10app_3x3field_PDD&Profiles_28082017 | 4297500%
29/08/2017 02:25 PM 313,813 9MeV_15x15app_15x15field_PDD&Profiles_SSD110_29082017.opab | 6 | 10 | 7x7 | 100 | 25/08/2017 | 6MeV_10x10app_7x7field_PDD&Profiles_25082017 | 4297200% | 16:26 | 182419 | PDD&Profiles | opab | 6MeV | 10x10app | 7x7field | 100SSD | 25082017 | 6 | 10 | 7x7 | 100 | 25/08/2017 | 6MeV_10x10app_7x7field_PDD&Profiles_25082017 | 4297200%
28/08/2017 11:09 AM 342,121 9MeV_20x20app_20x20field_PDD&Profiles_28082017.opab | 6 | 15 | 11x11 | 100 | 28/08/2017 | 6MeV_15x15app_11x11field_PDD&Profiles_28082017 | 4297500% | 14:36 | 206377 | PDD&Profiles | opab | 6MeV | 15x15app | 11x11field | 100SSD | 28082017 | 6 | 15 | 11x11 | 100 | 28/08/2017 | 6MeV_15x15app_11x11field_PDD&Profiles_28082017 | 4297500%
29/08/2017 01:09 PM 357,473 9MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017.opab | 6 | 15 | 15x15 | 100 | 28/08/2017 | 6MeV_15x15app_15x15field_PDD&Profiles_28082017 | 4297500% | 13:24 | 229977 | PDD&Profiles | opab | 6MeV | 15x15app | 15x15field | 100SSD | 28082017 | 6 | 15 | 15x15 | 100 | 28/08/2017 | 6MeV_15x15app_15x15field_PDD&Profiles_28082017 | 4297500%
25/08/2017 02:03 PM 350,969 9MeV_25x25app_21x21field_PDD&Profiles_25082017.opab | 6 | 20 | 20x20 | 100 | 28/08/2017 | 6MeV_20x20app_20x20field_PDD&Profiles_28082017 | 4297500% | 11:23 | 259831 | PDD&Profiles | opab | 6MeV | 20x20app | 20x20field | 100SSD | 28082017 | 6 | 20 | 20x20 | 100 | 28/08/2017 | 6MeV_20x20app_20x20field_PDD&Profiles_28082017 | 4297500%
11/10/2017 09:18 AM 204 files.txt | 6 | 25 | 21x21 | 100 | 25/08/2017 | 6MeV_25x25app_21x21field_PDD&Profiles_25082017 | 4297200% | 14:12 | 265731 | PDD&Profiles | opab | 6MeV | 25x25app | 21x21field | 100SSD | 25082017 | 6 | 25 | 21x21 | 100 | 25/08/2017 | 6MeV_25x25app_21x21field_PDD&Profiles_25082017 | 4297200%
25/08/2017 12:22 PM 382,121 9MeV_25x25app_25x25field_PDD&Profiles_25082017.opab | 6 | 25 | 25x25 | 100 | 25/08/2017 | 6MeV_25x25app_25x25field_PDD&Profiles_25082017 | 4298500% | 17:38 | 289986 | PDD&Profiles | opab | 6MeV | 25x25app | 25x25field | 100SSD | 25082017 | 6 | 25 | 25x25 | 100 | 25/08/2017 | 6MeV_25x25app_25x25field_PDD&Profiles_25082017 | 4298500%
07/09/2017 05:38 PM 401,552 9MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017.opab | 6 | 6 | 3x3 | 100 | 28/08/2017 | 6MeV_6x6app_3x3field_PDD&Profiles_28082017 | 4298500% | 17:38 | 158602 | PDD&Profiles | opab | 6MeV | 6x6app | 3x3field | 100SSD | 28082017 | 6 | 6 | 3x3 | 100 | 28/08/2017 | 6MeV_6x6app_3x3field_PDD&Profiles_28082017 | 4298500%
07/09/2017 05:38 PM 205,950 9MeV_6x6app_3x3field_PDD&Profiles_28082017.opab | 6 | 6 | 4x4 | 100 | 28/08/2017 | 6MeV_6x6app_4x4field_PDD&Profiles_28082017 | 4298900% | 17:41 | 164266 | PDD&Profiles | opab | 6MeV | 6x6app | 4x4field | 100SSD | 28082017 | 6 | 6 | 4x4 | 100 | 28/08/2017 | 6MeV_6x6app_4x4field_PDD&Profiles_28082017 | 4298900%
28/08/2017 04:10 PM 213,909 9MeV_6x6app_4x4field_PDD&Profiles_28082017.opab | 6 | 6 | 6x6 | 100 | 28/08/2017 | 6MeV_6x6app_6x6field_PDD&Profiles_28082017 | 4298500% | 17:38 | 176538 | PDD&Profiles | opab | 6MeV | 6x6app | 6x6field | 100SSD | 28082017 | 6 | 6 | 6x6 | 100 | 28/08/2017 | 6MeV_6x6app_6x6field_PDD&Profiles_28082017 | 4298500%
29/08/2017 04:53 PM 234,691 9MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017.opab
28/08/2017 03:21 PM 229,839 9MeV_6x6app_6x6field_PDD&Profiles_28082017.opab
17 File(s) 4,946,068 bytes
2 Dir(s) 833,814,077,440 bytes free
Volume in drive Q is 158
Volume Serial Number is 0000-0023
Directory of Q:\Treatment machines including photos\LINACS\Truebeam 2\Commissioning\Omnipro data\TR2 commissioning\Electron scans\16MeV
11/10/2017 09:24 AM <DIR> .
26/09/2017 08:59 AM <DIR> ..
29/08/2017 03:13 PM 401,809 16MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017.opab
24/08/2017 05:26 PM 884,138 16MeV_10x10app_10x10field_PDD&Profiles_24082017.opab
28/08/2017 11:52 AM 304,795 16MeV_10x10app_3x3field_PDD&Profiles_28082017.opab
25/08/2017 03:41 PM 353,647 16MeV_10x10app_7x7field_PDD&Profiles_25082017.opab
28/08/2017 02:02 PM 402,267 16MeV_15x15app_11x11field_PDD&Profiles_28082017.opab
29/08/2017 01:58 PM 468,597 16MeV_15x15app_15x15field_PDD&Profiles_110SSD_29082017.opab
28/08/2017 12:47 PM 451,119 16MeV_15x15app_15x15field_PDD&Profiles_28082017.opab
28/08/2017 10:42 AM 511,774 16MeV_20x20app_20x20field_PDD&Profiles_28082017.opab
29/08/2017 12:41 PM 535,385 16MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017.opab
25/08/2017 01:16 PM 524,043 16MeV_25x25app_21x21field_PDD&Profiles_25082017.opab
11/10/2017 09:24 AM 205 files.txt
25/08/2017 11:49 AM 572,659 16MeV_25x25app_25x25field_PDD&Profiles_25082017.opab
29/08/2017 11:01 AM 602,409 16MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017.opab
28/08/2017 04:39 PM 304,831 16MeV_6x6app_3x3field_PDD&Profiles_28082017.opab
28/08/2017 03:51 PM 316,395 16MeV_6x6app_4x4field_PDD&Profiles_28082017.opab
29/08/2017 04:27 PM 347,797 16MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017.opab
28/08/2017 03:05 PM 340,939 16MeV_6x6app_6x6field_PDD&Profiles_28082017.opab
17 File(s) 7,322,809 bytes
2 Dir(s) 833,770,405,888 bytes free
Volume in drive Q is 158
Volume Serial Number is 0000-0023
Directory of Q:\Treatment machines including photos\LINACS\Truebeam 2\Commissioning\Omnipro data\TR2 commissioning\Electron scans\20MeV
11/10/2017 09:24 AM <DIR> .
26/09/2017 08:59 AM <DIR> ..
29/08/2017 02:54 PM 372,642 20MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017.opab
25/08/2017 02:53 PM 1,002,020 20MeV_10x10app_10x10field_PDD&Profiles_25082017.opab
28/08/2017 11:39 AM 283,416 20MeV_10x10app_3x3field_PDD&Profiles_28082017.opab
25/08/2017 03:13 PM 328,256 20MeV_10x10app_7x7field_PDD&Profiles_25082017.opab
28/08/2017 01:47 PM 372,982 20MeV_15x15app_11x11field_PDD&Profiles_28082017.opab
28/08/2017 12:32 PM 417,704 20MeV_15x15app_15x15field_PDD&Profiles_28082017.opab
29/08/2017 01:38 PM 434,238 20MeV_15x15app_15x15field_PDD&Profiles_SSD110_29082017.opab
29/08/2017 12:22 PM 495,834 20MeV_20x20app_20x20field_PDD&Profiles_110SSD_29082017.opab
28/08/2017 10:25 AM 473,993 20MeV_20x20app_20x20field_PDD&Profiles_28082017.opab
25/08/2017 12:59 PM 485,436 20MeV_25x25app_21x21field_PDD&Profiles_25082017.opab
11/10/2017 09:24 AM 205 files.txt
29/08/2017 10:42 AM 557,312 20MeV_25x25app_25x25field_PDD&Profiles_110SSD_29082017.opab
25/08/2017 11:28 AM 530,158 20MeV_25x25app_25x25field_PDD&Profiles_25082017.opab
28/08/2017 04:28 PM 283,458 20MeV_6x6app_3x3field_PDD&Profiles_28082017.opab
28/08/2017 03:40 PM 294,432 20MeV_6x6app_4x4field_PDD&Profiles_28082017.opab
29/08/2017 04:13 PM 323,356 20MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017.opab
28/08/2017 02:54 PM 316,616 20MeV_6x6app_6x6field_PDD&Profiles_28082017.opab
17 File(s) 6,972,058 bytes
2 Dir(s) 833,758,576,640 bytes free

## PDD Scan Analysis
ParameterGroup
Electrons depth analysis (current protocol: 'AAPM TG-25 Extended')
| Radiation device | Energy | Applicator | SSD | Scan type | Start position [De] | Field detector | Field size inline | Field size crossline | R100 | R50 | R95 | R90 | R80 | Rtt | Rp | Ds | G | Ep0 | Field size inline | Field size crossline | SSD | Energy | Applicator | Field size | R100 | R50 | R95 | R90 | R80 | Rtt | Rp | Ds | G | Ep0
| TR2 | 9 MeV | 6 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 3.00 cm | 3.00 cm | 1.65 cm | 3.53 cm | 2.29 cm | 2.51 cm | 2.84 cm | 0.87 cm | 4.50 cm | 86.8% | 2.29 | 9.17 MeV | 3 | 3 | 100 | 9 | 6 | 3x3 | 1.65 | 3.53 | 2.29 | 2.51 | 2.84 | 0.87 | 4.5 | 86.8% | 2.29 | 9.17
| TR2 | 9 MeV | 25 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 25.00 cm | 25.00 cm | 2.03 cm | 3.54 cm | 2.55 cm | 2.75 cm | 3.01 cm | 1.32 cm | 4.31 cm | 82.1% | 2.75 | 8.81 MeV | 25 | 25 | 110 | 9 | 25 | 25x25 | 2.03 | 3.54 | 2.55 | 2.75 | 3.01 | 1.32 | 4.31 | 82.1% | 2.75 | 8.81
| TR2 | 9 MeV | 25 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 25.00 cm | 25.00 cm | 2.04 cm | 3.67 cm | 2.66 cm | 2.85 cm | 3.12 cm | 1.43 cm | 4.44 cm | 90.2% | 2.85 | 9.06 MeV | 25 | 25 | 100 | 9 | 25 | 25x25 | 2.04 | 3.67 | 2.66 | 2.85 | 3.12 | 1.43 | 4.44 | 90.2% | 2.85 | 9.06
| TR2 | 9 MeV | 25 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 21.00 cm | 21.00 cm | 2.19 cm | 3.70 cm | 2.69 cm | 2.89 cm | 3.15 cm | 1.45 cm | 4.48 cm | 86.5% | 2.82 | 9.14 MeV | 21 | 21 | 100 | 9 | 25 | 21x21 | 2.19 | 3.7 | 2.69 | 2.89 | 3.15 | 1.45 | 4.48 | 86.5% | 2.82 | 9.14
| TR2 | 9 MeV | 20 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 20.00 cm | 20.00 cm | 2.14 cm | 3.57 cm | 2.59 cm | 2.78 cm | 3.03 cm | 1.37 cm | 4.34 cm | 111.6% | 2.78 | 8.86 MeV | 20 | 20 | 110 | 9 | 20 | 20x20 | 2.14 | 3.57 | 2.59 | 2.78 | 3.03 | 1.37 | 4.34 | 111.6% | 2.78 | 8.86
| TR2 | 9 MeV | 20 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 20.00 cm | 20.00 cm | 2.02 cm | 3.58 cm | 2.57 cm | 2.76 cm | 3.03 cm | 1.29 cm | 4.35 cm | 92.2% | 2.77 | 8.88 MeV | 20 | 20 | 100 | 9 | 20 | 20x20 | 2.02 | 3.58 | 2.57 | 2.76 | 3.03 | 1.29 | 4.35 | 92.2% | 2.77 | 8.88
| TR2 | 9 MeV | 15 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 15.00 cm | 15.00 cm | 2.15 cm | 3.58 cm | 2.64 cm | 2.81 cm | 3.06 cm | 1.41 cm | 4.36 cm | 111.3% | 2.79 | 8.90 MeV | 15 | 15 | 110 | 9 | 15 | 15x15 | 2.15 | 3.58 | 2.64 | 2.81 | 3.06 | 1.41 | 4.36 | 111.3% | 2.79 | 8.9
| TR2 | 9 MeV | 15 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 15.00 cm | 15.00 cm | 2.14 cm | 3.60 cm | 2.59 cm | 2.80 cm | 3.06 cm |  | 4.38 cm | 88.4% | 2.78 | 8.94 MeV | 15 | 15 | 100 | 9 | 15 | 15x15 | 2.14 | 3.6 | 2.59 | 2.8 | 3.06 |  | 4.38 | 88.4% | 2.78 | 8.94
| TR2 | 9 MeV | 15 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 11.00 cm | 11.00 cm | 2.16 cm | 3.61 cm | 2.62 cm | 2.81 cm | 3.07 cm | 1.39 cm | 4.40 cm | 88.6% | 2.77 | 8.98 MeV | 11 | 11 | 100 | 9 | 15 | 11x11 | 2.16 | 3.61 | 2.62 | 2.81 | 3.07 | 1.39 | 4.4 | 88.6% | 2.77 | 8.98
| TR2 | 9 MeV | 10 | 100.00 cm | Depth Dose | -0.02 cm | EFD3G | 10.00 cm | 10.00 cm | 2.17 cm | 3.64 cm | 2.68 cm | 2.84 cm | 3.11 cm | 1.40 cm | 4.42 cm | 88.9% | 2.78 | 9.03 MeV | 10 | 10 | 100 | 9 | 10 | 10x10 | 2.17 | 3.64 | 2.68 | 2.84 | 3.11 | 1.4 | 4.42 | 88.9% | 2.78 | 9.03
| TR2 | 9 MeV | 10 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 10.00 cm | 10.00 cm | 2.04 cm | 3.60 cm | 2.65 cm | 2.83 cm | 3.07 cm |  | 4.38 cm | 80% | 2.8 | 8.93 MeV | 10 | 10 | 110 | 9 | 10 | 10x10 | 2.04 | 3.6 | 2.65 | 2.83 | 3.07 |  | 4.38 | 80% | 2.8 | 8.93
| TR2 | 9 MeV | 10 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 7.00 cm | 7.00 cm | 2.26 cm | 3.74 cm | 2.74 cm | 2.94 cm | 3.20 cm | 1.48 cm | 4.51 cm | 76.7% | 2.87 | 9.20 MeV | 7 | 7 | 100 | 9 | 10 | 7x7 | 2.26 | 3.74 | 2.74 | 2.94 | 3.2 | 1.48 | 4.51 | 76.7% | 2.87 | 9.2
| TR2 | 9 MeV | 10 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 3.00 cm | 3.00 cm | 1.56 cm | 3.45 cm | 2.19 cm | 2.42 cm | 2.74 cm | 0.65 cm | 4.43 cm | 90.6% | 2.24 | 9.04 MeV | 3 | 3 | 100 | 9 | 10 | 3x3 | 1.56 | 3.45 | 2.19 | 2.42 | 2.74 | 0.65 | 4.43 | 90.6% | 2.24 | 9.04
| TR2 | 9 MeV | 6 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 6.00 cm | 6.00 cm | 2.04 cm | 3.62 cm | 2.64 cm | 2.82 cm | 3.09 cm | 1.37 cm | 4.42 cm | 92.6% | 2.76 | 9.02 MeV | 6 | 6 | 100 | 9 | 6 | 6x6 | 2.04 | 3.62 | 2.64 | 2.82 | 3.09 | 1.37 | 4.42 | 92.6% | 2.76 | 9.02
| TR2 | 9 MeV | 6 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 6.00 cm | 6.00 cm | 2.16 cm | 3.62 cm | 2.64 cm | 2.82 cm | 3.09 cm | 1.46 cm | 4.40 cm | 103.9% | 2.78 | 8.98 MeV | 6 | 6 | 110 | 9 | 6 | 6x6 | 2.16 | 3.62 | 2.64 | 2.82 | 3.09 | 1.46 | 4.4 | 103.9% | 2.78 | 8.98
| TR2 | 9 MeV | 6 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 4.00 cm | 4.00 cm | 1.92 cm | 3.61 cm | 2.52 cm | 2.72 cm | 3.02 cm |  | 4.44 cm | 93.5% | 2.63 | 9.07 MeV | 4 | 4 | 100 | 9 | 6 | 4x4 | 1.92 | 3.61 | 2.52 | 2.72 | 3.02 |  | 4.44 | 93.5% | 2.63 | 9.07
| TR2 | 6 MeV | 25 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 25.00 cm | 25.00 cm | 1.26 cm | 2.34 cm | 1.64 cm | 1.77 cm | 1.95 cm | 0.86 cm | 2.91 cm | 91% | 2.53 | 6.00 MeV | 25 | 25 | 110 | 6 | 25 | 25x25 | 1.26 | 2.34 | 1.64 | 1.77 | 1.95 | 0.86 | 2.91 | 91% | 2.53 | 6
| TR2 | 6 MeV | 25 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 25.00 cm | 25.00 cm | 1.40 cm | 2.45 cm | 1.71 cm | 1.85 cm | 2.04 cm | 0.96 cm | 3.03 cm | 79.2% | 2.55 | 6.25 MeV | 25 | 25 | 100 | 6 | 25 | 25x25 | 1.4 | 2.45 | 1.71 | 1.85 | 2.04 | 0.96 | 3.03 | 79.2% | 2.55 | 6.25
| TR2 | 6 MeV | 25 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 21.00 cm | 21.00 cm | 1.41 cm | 2.48 cm | 1.75 cm | 1.88 cm | 2.07 cm | 0.99 cm | 3.06 cm | 88.9% | 2.59 | 6.30 MeV | 21 | 21 | 100 | 6 | 25 | 21x21 | 1.41 | 2.48 | 1.75 | 1.88 | 2.07 | 0.99 | 3.06 | 88.9% | 2.59 | 6.3
| TR2 | 6 MeV | 20 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 20.00 cm | 20.00 cm | 1.29 cm | 2.36 cm | 1.66 cm | 1.78 cm | 1.97 cm | 0.93 cm | 2.93 cm | 117.8% | 2.56 | 6.04 MeV | 20 | 20 | 110 | 6 | 20 | 20x20 | 1.29 | 2.36 | 1.66 | 1.78 | 1.97 | 0.93 | 2.93 | 117.8% | 2.56 | 6.04
| TR2 | 6 MeV | 20 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 20.00 cm | 20.00 cm | 1.29 cm | 2.36 cm | 1.63 cm | 1.76 cm | 1.95 cm |  | 2.95 cm | 96.9% | 2.44 | 6.09 MeV | 20 | 20 | 100 | 6 | 20 | 20x20 | 1.29 | 2.36 | 1.63 | 1.76 | 1.95 |  | 2.95 | 96.9% | 2.44 | 6.09
| TR2 | 6 MeV | 15 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 15.00 cm | 15.00 cm | 1.40 cm | 2.38 cm | 1.66 cm | 1.79 cm | 1.98 cm | 0.95 cm | 2.98 cm | 90.8% | 2.49 | 6.14 MeV | 15 | 15 | 100 | 6 | 15 | 15x15 | 1.4 | 2.38 | 1.66 | 1.79 | 1.98 | 0.95 | 2.98 | 90.8% | 2.49 | 6.14
| TR2 | 6 MeV | 15 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 15.00 cm | 15.00 cm | 1.39 cm | 2.39 cm | 1.69 cm | 1.82 cm | 2.00 cm | 0.96 cm | 2.95 cm | 116.3% | 2.61 | 6.08 MeV | 15 | 15 | 110 | 6 | 15 | 15x15 | 1.39 | 2.39 | 1.69 | 1.82 | 2 | 0.96 | 2.95 | 116.3% | 2.61 | 6.08
| TR2 | 6 MeV | 15 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 11.00 cm | 11.00 cm | 1.38 cm | 2.40 cm | 1.67 cm | 1.80 cm | 2.00 cm | 0.96 cm | 2.99 cm | 89.4% | 2.5 | 6.16 MeV | 11 | 11 | 100 | 6 | 15 | 11x11 | 1.38 | 2.4 | 1.67 | 1.8 | 2 | 0.96 | 2.99 | 89.4% | 2.5 | 6.16
| TR2 | 6 MeV | 10 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 10.00 cm | 10.00 cm | 1.38 cm | 2.42 cm | 1.70 cm | 1.83 cm | 2.02 cm | 0.96 cm | 2.99 cm | 90.3% | 2.57 | 6.17 MeV | 10 | 10 | 100 | 6 | 10 | 10x10 | 1.38 | 2.42 | 1.7 | 1.83 | 2.02 | 0.96 | 2.99 | 90.3% | 2.57 | 6.17
| TR2 | 6 MeV | 10 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 10.00 cm | 10.00 cm | 1.38 cm | 2.40 cm | 1.71 cm | 1.84 cm | 2.02 cm | 0.98 cm | 2.97 cm | 76.2% | 2.61 | 6.12 MeV | 10 | 10 | 110 | 6 | 10 | 10x10 | 1.38 | 2.4 | 1.71 | 1.84 | 2.02 | 0.98 | 2.97 | 76.2% | 2.61 | 6.12
| TR2 | 6 MeV | 10 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 7.00 cm | 7.00 cm | 1.50 cm | 2.52 cm | 1.80 cm | 1.94 cm | 2.12 cm | 1.08 cm | 3.11 cm | 76.6% | 2.61 | 6.41 MeV | 7 | 7 | 100 | 6 | 10 | 7x7 | 1.5 | 2.52 | 1.8 | 1.94 | 2.12 | 1.08 | 3.11 | 76.6% | 2.61 | 6.41
| TR2 | 6 MeV | 10 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 3.00 cm | 3.00 cm | 1.17 cm | 2.35 cm | 1.56 cm | 1.71 cm | 1.91 cm |  | 2.99 cm | 87.7% | 2.33 | 6.15 MeV | 3 | 3 | 100 | 6 | 10 | 3x3 | 1.17 | 2.35 | 1.56 | 1.71 | 1.91 |  | 2.99 | 87.7% | 2.33 | 6.15
| TR2 | 6 MeV | 6 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 6.00 cm | 6.00 cm | 1.39 cm | 2.42 cm | 1.70 cm | 1.82 cm | 2.02 cm | 0.97 cm | 3.00 cm | 77% | 2.56 | 6.17 MeV | 6 | 6 | 100 | 6 | 6 | 6x6 | 1.39 | 2.42 | 1.7 | 1.82 | 2.02 | 0.97 | 3 | 77% | 2.56 | 6.17
| TR2 | 6 MeV | 6 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 6.00 cm | 6.00 cm | 1.40 cm | 2.42 cm | 1.75 cm | 1.86 cm | 2.04 cm | 1.00 cm | 3.00 cm | 99.6% | 2.59 | 6.18 MeV | 6 | 6 | 110 | 6 | 6 | 6x6 | 1.4 | 2.42 | 1.75 | 1.86 | 2.04 | 1 | 3 | 99.6% | 2.59 | 6.18
| TR2 | 6 MeV | 6 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 4.00 cm | 4.00 cm | 1.40 cm | 2.43 cm | 1.70 cm | 1.84 cm | 2.03 cm | 0.95 cm | 3.02 cm | 77.5% | 2.51 | 6.23 MeV | 4 | 4 | 100 | 6 | 6 | 4x4 | 1.4 | 2.43 | 1.7 | 1.84 | 2.03 | 0.95 | 3.02 | 77.5% | 2.51 | 6.23
| TR2 | 6 MeV | 6 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 3.00 cm | 3.00 cm | 1.27 cm | 2.42 cm | 1.65 cm | 1.79 cm | 2.00 cm | 0.87 cm | 3.04 cm | 80.2% | 2.41 | 6.27 MeV | 3 | 3 | 100 | 6 | 6 | 3x3 | 1.27 | 2.42 | 1.65 | 1.79 | 2 | 0.87 | 3.04 | 80.2% | 2.41 | 6.27
| TR2 | 12 MeV | 25 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 25.00 cm | 25.00 cm | 2.85 cm | 4.94 cm | 3.57 cm | 3.83 cm | 4.21 cm |  | 5.96 cm | 88.4% | 2.84 | 12.11 MeV | 25 | 25 | 110 | 12 | 25 | 25x25 | 2.85 | 4.94 | 3.57 | 3.83 | 4.21 |  | 5.96 | 88.4% | 2.84 | 12.11
| TR2 | 12 MeV | 25 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 25.00 cm | 25.00 cm | 2.73 cm | 5.05 cm | 3.67 cm | 3.96 cm | 4.33 cm | 1.19 cm | 6.09 cm | 90.5% | 2.88 | 12.36 MeV | 25 | 25 | 100 | 12 | 25 | 25x25 | 2.73 | 5.05 | 3.67 | 3.96 | 4.33 | 1.19 | 6.09 | 90.5% | 2.88 | 12.36
| TR2 | 12 MeV | 25 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 21.00 cm | 21.00 cm | 2.85 cm | 5.09 cm | 3.71 cm | 4.00 cm | 4.35 cm | 1.13 cm | 6.11 cm | 87.1% | 2.91 | 12.42 MeV | 21 | 21 | 100 | 12 | 25 | 21x21 | 2.85 | 5.09 | 3.71 | 4 | 4.35 | 1.13 | 6.11 | 87.1% | 2.91 | 12.42
| TR2 | 12 MeV | 20 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 20.00 cm | 20.00 cm | 2.97 cm | 4.95 cm | 3.58 cm | 3.85 cm | 4.22 cm | 1.40 cm | 5.98 cm | 109.7% | 2.84 | 12.15 MeV | 20 | 20 | 110 | 12 | 20 | 20x20 | 2.97 | 4.95 | 3.58 | 3.85 | 4.22 | 1.4 | 5.98 | 109.7% | 2.84 | 12.15
| TR2 | 12 MeV | 20 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 20.00 cm | 20.00 cm | 2.84 cm | 4.96 cm | 3.60 cm | 3.87 cm | 4.23 cm | 1.05 cm | 6.00 cm | 91.6% | 2.83 | 12.18 MeV | 20 | 20 | 100 | 12 | 20 | 20x20 | 2.84 | 4.96 | 3.6 | 3.87 | 4.23 | 1.05 | 6 | 91.6% | 2.83 | 12.18
| TR2 | 12 MeV | 15 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 15.00 cm | 15.00 cm | 2.95 cm | 4.97 cm | 3.64 cm | 3.90 cm | 4.26 cm | 1.49 cm | 6.01 cm | 109.4% | 2.84 | 12.21 MeV | 15 | 15 | 110 | 12 | 15 | 15x15 | 2.95 | 4.97 | 3.64 | 3.9 | 4.26 | 1.49 | 6.01 | 109.4% | 2.84 | 12.21
| TR2 | 12 MeV | 15 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 15.00 cm | 15.00 cm | 2.75 cm | 4.99 cm | 3.62 cm | 3.89 cm | 4.26 cm | 1.34 cm | 6.02 cm | 88.1% | 2.85 | 12.23 MeV | 15 | 15 | 100 | 12 | 15 | 15x15 | 2.75 | 4.99 | 3.62 | 3.89 | 4.26 | 1.34 | 6.02 | 88.1% | 2.85 | 12.23
| TR2 | 12 MeV | 15 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 11.00 cm | 11.00 cm | 2.88 cm | 5.00 cm | 3.64 cm | 3.91 cm | 4.28 cm | 1.31 cm | 6.04 cm | 89.2% | 2.85 | 12.27 MeV | 11 | 11 | 100 | 12 | 15 | 11x11 | 2.88 | 5 | 3.64 | 3.91 | 4.28 | 1.31 | 6.04 | 89.2% | 2.85 | 12.27
| TR2 | 12 MeV | 10 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 10.00 cm | 10.00 cm | 2.85 cm | 5.03 cm | 3.67 cm | 3.94 cm | 4.31 cm |  | 6.06 cm | 86.3% | 2.89 | 12.31 MeV | 10 | 10 | 100 | 12 | 10 | 10x10 | 2.85 | 5.03 | 3.67 | 3.94 | 4.31 |  | 6.06 | 86.3% | 2.89 | 12.31
| TR2 | 12 MeV | 10 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 10.00 cm | 10.00 cm | 2.85 cm | 4.98 cm | 3.64 cm | 3.92 cm | 4.27 cm |  | 6.02 cm | 83.6% | 2.87 | 12.22 MeV | 10 | 10 | 110 | 12 | 10 | 10x10 | 2.85 | 4.98 | 3.64 | 3.92 | 4.27 |  | 6.02 | 83.6% | 2.87 | 12.22
| TR2 | 12 MeV | 10 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 7.00 cm | 7.00 cm | 2.98 cm | 5.11 cm | 3.74 cm | 4.02 cm | 4.39 cm | 1.46 cm | 6.16 cm | 78.1% | 2.89 | 12.51 MeV | 7 | 7 | 100 | 12 | 10 | 7x7 | 2.98 | 5.11 | 3.74 | 4.02 | 4.39 | 1.46 | 6.16 | 78.1% | 2.89 | 12.51
| TR2 | 12 MeV | 10 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 3.00 cm | 3.00 cm | 1.89 cm | 4.49 cm | 2.68 cm | 2.99 cm | 3.45 cm | 0.37 cm | 6.06 cm | 93.3% | 1.9 | 12.31 MeV | 3 | 3 | 100 | 12 | 10 | 3x3 | 1.89 | 4.49 | 2.68 | 2.99 | 3.45 | 0.37 | 6.06 | 93.3% | 1.9 | 12.31
| TR2 | 12 MeV | 6 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 6.00 cm | 6.00 cm | 2.72 cm | 4.99 cm | 3.60 cm | 3.87 cm | 4.25 cm |  | 6.04 cm | 94.7% | 2.84 | 12.27 MeV | 6 | 6 | 100 | 12 | 6 | 6x6 | 2.72 | 4.99 | 3.6 | 3.87 | 4.25 |  | 6.04 | 94.7% | 2.84 | 12.27
| TR2 | 12 MeV | 6 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 6.00 cm | 6.00 cm | 2.85 cm | 4.99 cm | 3.63 cm | 3.91 cm | 4.26 cm |  | 6.03 cm | 106.5% | 2.85 | 12.26 MeV | 6 | 6 | 110 | 12 | 6 | 6x6 | 2.85 | 4.99 | 3.63 | 3.91 | 4.26 |  | 6.03 | 106.5% | 2.85 | 12.26
| TR2 | 12 MeV | 6 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 4.00 cm | 4.00 cm | 2.25 cm | 4.86 cm | 3.21 cm | 3.54 cm | 3.96 cm | 0.75 cm | 6.12 cm | 96% | 2.38 | 12.44 MeV | 4 | 4 | 100 | 12 | 6 | 4x4 | 2.25 | 4.86 | 3.21 | 3.54 | 3.96 | 0.75 | 6.12 | 96% | 2.38 | 12.44
| TR2 | 12 MeV | 6 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 3.00 cm | 3.00 cm | 2.00 cm | 4.62 cm | 2.81 cm | 3.15 cm | 3.60 cm | 0.51 cm | 6.12 cm | 90.7% | 2 | 12.43 MeV | 3 | 3 | 100 | 12 | 6 | 3x3 | 2 | 4.62 | 2.81 | 3.15 | 3.6 | 0.51 | 6.12 | 90.7% | 2 | 12.43
| TR2 | 16 MeV | 25 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 25.00 cm | 25.00 cm | 3.28 cm | 6.59 cm | 4.57 cm | 4.99 cm | 5.56 cm | 0.65 cm | 7.95 cm | 90% | 2.82 | 16.11 MeV | 25 | 25 | 110 | 16 | 25 | 25x25 | 3.28 | 6.59 | 4.57 | 4.99 | 5.56 | 0.65 | 7.95 | 90% | 2.82 | 16.11
| TR2 | 16 MeV | 25 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 25.00 cm | 25.00 cm | 3.39 cm | 6.66 cm | 4.66 cm | 5.10 cm | 5.66 cm | 0.55 cm | 8.03 cm | 92.1% | 2.84 | 16.27 MeV | 25 | 25 | 100 | 16 | 25 | 25x25 | 3.39 | 6.66 | 4.66 | 5.1 | 5.66 | 0.55 | 8.03 | 92.1% | 2.84 | 16.27
| TR2 | 16 MeV | 25 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 21.00 cm | 21.00 cm | 3.14 cm | 6.69 cm | 4.67 cm | 5.12 cm | 5.67 cm | 0.51 cm | 8.06 cm | 92.3% | 2.83 | 16.35 MeV | 21 | 21 | 100 | 16 | 25 | 21x21 | 3.14 | 6.69 | 4.67 | 5.12 | 5.67 | 0.51 | 8.06 | 92.3% | 2.83 | 16.35
| TR2 | 16 MeV | 20 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 20.00 cm | 20.00 cm | 3.29 cm | 6.61 cm | 4.61 cm | 5.04 cm | 5.57 cm | 0.70 cm | 7.97 cm | 111.9% | 2.82 | 16.16 MeV | 20 | 20 | 110 | 16 | 20 | 20x20 | 3.29 | 6.61 | 4.61 | 5.04 | 5.57 | 0.7 | 7.97 | 111.9% | 2.82 | 16.16
| TR2 | 16 MeV | 20 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 20.00 cm | 20.00 cm | 2.75 cm | 6.56 cm | 4.54 cm | 5.01 cm | 5.56 cm | 0.41 cm | 7.95 cm | 92.6% | 2.77 | 16.12 MeV | 20 | 20 | 100 | 16 | 20 | 20x20 | 2.75 | 6.56 | 4.54 | 5.01 | 5.56 | 0.41 | 7.95 | 92.6% | 2.77 | 16.12
| TR2 | 16 MeV | 15 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 15.00 cm | 15.00 cm | 2.66 cm | 6.60 cm | 4.61 cm | 5.06 cm | 5.59 cm | 0.50 cm | 7.96 cm | 89.7% | 2.81 | 16.15 MeV | 15 | 15 | 100 | 16 | 15 | 15x15 | 2.66 | 6.6 | 4.61 | 5.06 | 5.59 | 0.5 | 7.96 | 89.7% | 2.81 | 16.15
| TR2 | 16 MeV | 15 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 15.00 cm | 15.00 cm | 2.92 cm | 6.63 cm | 4.68 cm | 5.09 cm | 5.61 cm | 0.84 cm | 8.00 cm | 111.7% | 2.81 | 16.23 MeV | 15 | 15 | 110 | 16 | 15 | 15x15 | 2.92 | 6.63 | 4.68 | 5.09 | 5.61 | 0.84 | 8 | 111.7% | 2.81 | 16.23
| TR2 | 16 MeV | 15 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 11.00 cm | 11.00 cm | 3.03 cm | 6.61 cm | 4.63 cm | 5.07 cm | 5.60 cm | 0.49 cm | 7.98 cm | 91.6% | 2.82 | 16.17 MeV | 11 | 11 | 100 | 16 | 15 | 11x11 | 3.03 | 6.61 | 4.63 | 5.07 | 5.6 | 0.49 | 7.98 | 91.6% | 2.82 | 16.17
| TR2 | 16 MeV | 10 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 10.00 cm | 10.00 cm | 2.92 cm | 6.67 cm | 4.64 cm | 5.11 cm | 5.64 cm | 0.48 cm | 8.06 cm | 87.4% | 2.81 | 16.35 MeV | 10 | 10 | 100 | 16 | 10 | 10x10 | 2.92 | 6.67 | 4.64 | 5.11 | 5.64 | 0.48 | 8.06 | 87.4% | 2.81 | 16.35
| TR2 | 16 MeV | 10 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 10.00 cm | 10.00 cm | 3.52 cm | 6.64 cm | 4.72 cm | 5.11 cm | 5.63 cm | 0.90 cm | 8.01 cm | 112.7% | 2.83 | 16.24 MeV | 10 | 10 | 110 | 16 | 10 | 10x10 | 3.52 | 6.64 | 4.72 | 5.11 | 5.63 | 0.9 | 8.01 | 112.7% | 2.83 | 16.24
| TR2 | 16 MeV | 10 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 7.00 cm | 7.00 cm | 3.18 cm | 6.68 cm | 4.57 cm | 5.05 cm | 5.63 cm | 0.53 cm | 8.13 cm | 79.3% | 2.72 | 16.49 MeV | 7 | 7 | 100 | 16 | 10 | 7x7 | 3.18 | 6.68 | 4.57 | 5.05 | 5.63 | 0.53 | 8.13 | 79.3% | 2.72 | 16.49
| TR2 | 16 MeV | 10 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 3.00 cm | 3.00 cm | 1.58 cm | 5.47 cm | 3.01 cm | 3.48 cm | 4.09 cm | 0.21 cm | 7.71 cm | 95% | 1.69 | 15.63 MeV | 3 | 3 | 100 | 16 | 10 | 3x3 | 1.58 | 5.47 | 3.01 | 3.48 | 4.09 | 0.21 | 7.71 | 95% | 1.69 | 15.63
| TR2 | 16 MeV | 6 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 6.00 cm | 6.00 cm | 2.90 cm | 6.53 cm | 4.34 cm | 4.82 cm | 5.42 cm | 0.44 cm | 8.02 cm | 95.9% | 2.61 | 16.26 MeV | 6 | 6 | 100 | 16 | 6 | 6x6 | 2.9 | 6.53 | 4.34 | 4.82 | 5.42 | 0.44 | 8.02 | 95.9% | 2.61 | 16.26
| TR2 | 16 MeV | 6 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 6.00 cm | 6.00 cm | 3.04 cm | 6.61 cm | 4.52 cm | 4.97 cm | 5.54 cm | 0.85 cm | 8.06 cm | 106.9% | 2.68 | 16.35 MeV | 6 | 6 | 110 | 16 | 6 | 6x6 | 3.04 | 6.61 | 4.52 | 4.97 | 5.54 | 0.85 | 8.06 | 106.9% | 2.68 | 16.35
| TR2 | 16 MeV | 6 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 4.00 cm | 4.00 cm | 1.96 cm | 6.09 cm | 3.64 cm | 4.14 cm | 4.75 cm | 0.36 cm | 8.02 cm | 97.8% | 2.02 | 16.25 MeV | 4 | 4 | 100 | 16 | 6 | 4x4 | 1.96 | 6.09 | 3.64 | 4.14 | 4.75 | 0.36 | 8.02 | 97.8% | 2.02 | 16.25
| TR2 | 16 MeV | 6 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 3.00 cm | 3.00 cm | 1.58 cm | 5.64 cm | 3.19 cm | 3.65 cm | 4.25 cm |  | 7.83 cm | 98.2% | 1.75 | 15.87 MeV | 3 | 3 | 100 | 16 | 6 | 3x3 | 1.58 | 5.64 | 3.19 | 3.65 | 4.25 |  | 7.83 | 98.2% | 1.75 | 15.87
| TR2 | 20 MeV | 25 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 25.00 cm | 25.00 cm | 2.39 cm | 8.23 cm | 5.14 cm | 5.91 cm | 6.79 cm | 0.41 cm | 10.01 cm | 92.6% | 2.65 | 20.30 MeV | 25 | 25 | 100 | 20 | 25 | 25x25 | 2.39 | 8.23 | 5.14 | 5.91 | 6.79 | 0.41 | 10.01 | 92.6% | 2.65 | 20.3
| TR2 | 20 MeV | 25 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 25.00 cm | 25.00 cm | 2.98 cm | 8.14 cm | 5.12 cm | 5.91 cm | 6.75 cm | 0.43 cm | 9.86 cm | 92.1% | 2.71 | 19.99 MeV | 25 | 25 | 110 | 20 | 25 | 25x25 | 2.98 | 8.14 | 5.12 | 5.91 | 6.75 | 0.43 | 9.86 | 92.1% | 2.71 | 19.99
| TR2 | 20 MeV | 25 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 21.00 cm | 21.00 cm | 2.49 cm | 8.26 cm | 5.10 cm | 5.91 cm | 6.80 cm | 0.42 cm | 10.05 cm | 93.2% | 2.65 | 20.36 MeV | 21 | 21 | 100 | 20 | 25 | 21x21 | 2.49 | 8.26 | 5.1 | 5.91 | 6.8 | 0.42 | 10.05 | 93.2% | 2.65 | 20.36
| TR2 | 20 MeV | 20 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 20.00 cm | 20.00 cm | 2.87 cm | 8.14 cm | 4.99 cm | 5.80 cm | 6.68 cm | 0.32 cm | 9.93 cm | 92.9% | 2.62 | 20.13 MeV | 20 | 20 | 100 | 20 | 20 | 20x20 | 2.87 | 8.14 | 4.99 | 5.8 | 6.68 | 0.32 | 9.93 | 92.9% | 2.62 | 20.13
| TR2 | 20 MeV | 20 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 20.00 cm | 20.00 cm | 2.89 cm | 8.15 cm | 5.23 cm | 5.95 cm | 6.78 cm |  | 9.90 cm | 113.2% | 2.7 | 20.06 MeV | 20 | 20 | 110 | 20 | 20 | 20x20 | 2.89 | 8.15 | 5.23 | 5.95 | 6.78 |  | 9.9 | 113.2% | 2.7 | 20.06
| TR2 | 20 MeV | 15 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 15.00 cm | 15.00 cm | 2.62 cm | 8.18 cm | 5.34 cm | 6.03 cm | 6.83 cm | 0.53 cm | 9.90 cm | 113.4% | 2.72 | 20.07 MeV | 15 | 15 | 110 | 20 | 15 | 15x15 | 2.62 | 8.18 | 5.34 | 6.03 | 6.83 | 0.53 | 9.9 | 113.4% | 2.72 | 20.07
| TR2 | 20 MeV | 15 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 15.00 cm | 15.00 cm | 2.61 cm | 8.16 cm | 5.07 cm | 5.86 cm | 6.72 cm | 0.35 cm | 9.96 cm | 90.7% | 2.62 | 20.19 MeV | 15 | 15 | 100 | 20 | 15 | 15x15 | 2.61 | 8.16 | 5.07 | 5.86 | 6.72 | 0.35 | 9.96 | 90.7% | 2.62 | 20.19
| TR2 | 20 MeV | 15 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 11.00 cm | 11.00 cm | 2.26 cm | 8.17 cm | 5.08 cm | 5.83 cm | 6.72 cm | 0.32 cm | 9.98 cm | 93.4% | 2.62 | 20.23 MeV | 11 | 11 | 100 | 20 | 15 | 11x11 | 2.26 | 8.17 | 5.08 | 5.83 | 6.72 | 0.32 | 9.98 | 93.4% | 2.62 | 20.23
| TR2 | 20 MeV | 10 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 10.00 cm | 10.00 cm | 2.25 cm | 8.26 cm | 5.02 cm | 5.85 cm | 6.77 cm | 0.38 cm | 10.05 cm | 80% | 2.65 | 20.37 MeV | 10 | 10 | 100 | 20 | 10 | 10x10 | 2.25 | 8.26 | 5.02 | 5.85 | 6.77 | 0.38 | 10.05 | 80% | 2.65 | 20.37
| TR2 | 20 MeV | 10 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 10.00 cm | 10.00 cm | 3.12 cm | 8.18 cm | 5.29 cm | 5.98 cm | 6.81 cm | 0.62 cm | 9.88 cm | 114.5% | 2.77 | 20.02 MeV | 10 | 10 | 110 | 20 | 10 | 10x10 | 3.12 | 8.18 | 5.29 | 5.98 | 6.81 | 0.62 | 9.88 | 114.5% | 2.77 | 20.02
| TR2 | 20 MeV | 10 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 7.00 cm | 7.00 cm | 2.02 cm | 8.14 cm | 4.74 cm | 5.61 cm | 6.52 cm | 0.39 cm | 10.13 cm | 81% | 2.41 | 20.54 MeV | 7 | 7 | 100 | 20 | 10 | 7x7 | 2.02 | 8.14 | 4.74 | 5.61 | 6.52 | 0.39 | 10.13 | 81% | 2.41 | 20.54
| TR2 | 20 MeV | 10 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 3.00 cm | 3.00 cm | 1.54 cm | 6.35 cm | 3.32 cm | 3.88 cm | 4.64 cm | 0.10 cm | 9.16 cm | 98% | 1.59 | 18.57 MeV | 3 | 3 | 100 | 20 | 10 | 3x3 | 1.54 | 6.35 | 3.32 | 3.88 | 4.64 | 0.1 | 9.16 | 98% | 1.59 | 18.57
| TR2 | 20 MeV | 6 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 6.00 cm | 6.00 cm | 1.91 cm | 7.91 cm | 4.55 cm | 5.34 cm | 6.25 cm |  | 10.04 cm | 96.4% | 2.25 | 20.36 MeV | 6 | 6 | 100 | 20 | 6 | 6x6 | 1.91 | 7.91 | 4.55 | 5.34 | 6.25 |  | 10.04 | 96.4% | 2.25 | 20.36
| TR2 | 20 MeV | 6 | 110.00 cm | Depth Dose | -0.05 cm | EFD3G | 6.00 cm | 6.00 cm | 3.25 cm | 8.05 cm | 4.97 cm | 5.70 cm | 6.56 cm | 0.60 cm | 9.95 cm | 107.3% | 2.48 | 20.18 MeV | 6 | 6 | 110 | 20 | 6 | 6x6 | 3.25 | 8.05 | 4.97 | 5.7 | 6.56 | 0.6 | 9.95 | 107.3% | 2.48 | 20.18
| TR2 | 20 MeV | 6 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 4.00 cm | 4.00 cm | 1.78 cm | 7.13 cm | 3.85 cm | 4.56 cm | 5.36 cm | 0.24 cm | 9.80 cm | 98.6% | 1.77 | 19.87 MeV | 4 | 4 | 100 | 20 | 6 | 4x4 | 1.78 | 7.13 | 3.85 | 4.56 | 5.36 | 0.24 | 9.8 | 98.6% | 1.77 | 19.87
| TR2 | 20 MeV | 6 | 100.00 cm | Depth Dose | -0.05 cm | EFD3G | 3.00 cm | 3.00 cm | 1.65 cm | 6.51 cm | 3.42 cm | 4.04 cm | 4.79 cm | 0.22 cm | 9.32 cm | 99.3% | 1.61 | 18.90 MeV | 3 | 3 | 100 | 20 | 6 | 3x3 | 1.65 | 6.51 | 3.42 | 4.04 | 4.79 | 0.22 | 9.32 | 99.3% | 1.61 | 18.9

## PDD Parameters
PDD Parameters
Scan Date | File Name | Modified Date | Machine | SSD | Energy | Applicator | Field Size | R100 | R50 | R95 | R90 | R80 | Rtt | Rp | Ds | G | Ep0
28/08/2017 | 6MeV_6x6app_3x3field_PDD&Profiles_28082017 | 4298500% | TR2 | 100 | 6 | 6 | 3x3 | 1.27 | 2.42 | 1.65 | 1.79 | 2 | 0.87 | 3.04 | 0.802 | 2.41 | 6.27
28/08/2017 | 6MeV_6x6app_4x4field_PDD&Profiles_28082017 | 4298900% | TR2 | 100 | 6 | 6 | 4x4 | 1.4 | 2.43 | 1.7 | 1.84 | 2.03 | 0.95 | 3.02 | 0.775 | 2.51 | 6.23
28/08/2017 | 6MeV_6x6app_6x6field_PDD&Profiles_28082017 | 4298500% | TR2 | 100 | 6 | 6 | 6x6 | 1.39 | 2.42 | 1.7 | 1.82 | 2.02 | 0.97 | 3 | 0.77 | 2.56 | 6.17
24/08/2017 | 6MeV_10x10app_10x10field_PDD&Profiles_24082017 | 4297200% | TR2 | 100 | 6 | 10 | 10x10 | 1.38 | 2.42 | 1.7 | 1.83 | 2.02 | 0.96 | 2.99 | 0.903 | 2.57 | 6.17
28/08/2017 | 6MeV_10x10app_3x3field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 6 | 10 | 3x3 | 1.17 | 2.35 | 1.56 | 1.71 | 1.91 |  | 2.99 | 0.877 | 2.33 | 6.15
25/08/2017 | 6MeV_10x10app_7x7field_PDD&Profiles_25082017 | 4297200% | TR2 | 100 | 6 | 10 | 7x7 | 1.5 | 2.52 | 1.8 | 1.94 | 2.12 | 1.08 | 3.11 | 0.766 | 2.61 | 6.41
28/08/2017 | 6MeV_15x15app_11x11field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 6 | 15 | 11x11 | 1.38 | 2.4 | 1.67 | 1.8 | 2 | 0.96 | 2.99 | 0.894 | 2.5 | 6.16
28/08/2017 | 6MeV_15x15app_15x15field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 6 | 15 | 15x15 | 1.4 | 2.38 | 1.66 | 1.79 | 1.98 | 0.95 | 2.98 | 0.908 | 2.49 | 6.14
28/08/2017 | 6MeV_20x20app_20x20field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 6 | 20 | 20x20 | 1.29 | 2.36 | 1.63 | 1.76 | 1.95 |  | 2.95 | 0.969 | 2.44 | 6.09
25/08/2017 | 6MeV_25x25app_21x21field_PDD&Profiles_25082017 | 4297200% | TR2 | 100 | 6 | 25 | 21x21 | 1.41 | 2.48 | 1.75 | 1.88 | 2.07 | 0.99 | 3.06 | 0.889 | 2.59 | 6.3
25/08/2017 | 6MeV_25x25app_25x25field_PDD&Profiles_25082017 | 4298500% | TR2 | 100 | 6 | 25 | 25x25 | 1.4 | 2.45 | 1.71 | 1.85 | 2.04 | 0.96 | 3.03 | 0.792 | 2.55 | 6.25
28/08/2017 | 9MeV_6x6app_3x3field_PDD&Profiles_28082017 | 4298500% | TR2 | 100 | 9 | 6 | 3x3 | 1.65 | 3.53 | 2.29 | 2.51 | 2.84 | 0.87 | 4.5 | 0.868 | 2.29 | 9.17
28/08/2017 | 9MeV_6x6app_4x4field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 9 | 6 | 4x4 | 1.92 | 3.61 | 2.52 | 2.72 | 3.02 |  | 4.44 | 0.935 | 2.63 | 9.07
28/08/2017 | 9MeV_6x6app_6x6field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 9 | 6 | 6x6 | 2.04 | 3.62 | 2.64 | 2.82 | 3.09 | 1.37 | 4.42 | 0.926 | 2.76 | 9.02
24/08/2017 | 9MeV_10x10app_10x10field_PDD&Profiles_24082017 | 4297100% | TR2 | 100 | 9 | 10 | 10x10 | 2.17 | 3.64 | 2.68 | 2.84 | 3.11 | 1.4 | 4.42 | 0.889 | 2.78 | 9.03
28/08/2017 | 9MeV_10x10app_3x3field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 9 | 10 | 3x3 | 1.56 | 3.45 | 2.19 | 2.42 | 2.74 | 0.65 | 4.43 | 0.906 | 2.24 | 9.04
25/08/2017 | 9MeV_10x10app_7x7field_PDD&Profiles_25082017 | 4297200% | TR2 | 100 | 9 | 10 | 7x7 | 2.26 | 3.74 | 2.74 | 2.94 | 3.2 | 1.48 | 4.51 | 0.767 | 2.87 | 9.2
28/08/2017 | 9MeV_15x15app_11x11field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 9 | 15 | 11x11 | 2.16 | 3.61 | 2.62 | 2.81 | 3.07 | 1.39 | 4.4 | 0.886 | 2.77 | 8.98
28/08/2017 | 9MeV_15x15app_15x15field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 9 | 15 | 15x15 | 2.14 | 3.6 | 2.59 | 2.8 | 3.06 |  | 4.38 | 0.884 | 2.78 | 8.94
28/08/2017 | 9MeV_20x20app_20x20field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 9 | 20 | 20x20 | 2.02 | 3.58 | 2.57 | 2.76 | 3.03 | 1.29 | 4.35 | 0.922 | 2.77 | 8.88
25/08/2017 | 9MeV_25x25app_21x21field_PDD&Profiles_25082017 | 4297200% | TR2 | 100 | 9 | 25 | 21x21 | 2.19 | 3.7 | 2.69 | 2.89 | 3.15 | 1.45 | 4.48 | 0.865 | 2.82 | 9.14
25/08/2017 | 9MeV_25x25app_25x25field_PDD&Profiles_25082017 | 4297200% | TR2 | 100 | 9 | 25 | 25x25 | 2.04 | 3.67 | 2.66 | 2.85 | 3.12 | 1.43 | 4.44 | 0.902 | 2.85 | 9.06
28/08/2017 | 12MeV_6x6app_3x3field_PDD&Profiles_28082017 | 4298500% | TR2 | 100 | 12 | 6 | 3x3 | 2 | 4.62 | 2.81 | 3.15 | 3.6 | 0.51 | 6.12 | 0.907 | 2 | 12.43
28/08/2017 | 12MeV_6x6app_4x4field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 12 | 6 | 4x4 | 2.25 | 4.86 | 3.21 | 3.54 | 3.96 | 0.75 | 6.12 | 0.96 | 2.38 | 12.44
28/08/2017 | 12MeV_6x6app_6x6field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 12 | 6 | 6x6 | 2.72 | 4.99 | 3.6 | 3.87 | 4.25 |  | 6.04 | 0.947 | 2.84 | 12.27
24/08/2017 | 12MeV_10x10app_10x10field_PDD&Profiles_24082017 | 4297100% | TR2 | 100 | 12 | 10 | 10x10 | 2.85 | 5.03 | 3.67 | 3.94 | 4.31 |  | 6.06 | 0.863 | 2.89 | 12.31
28/08/2017 | 12MeV_10x10app_3x3field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 12 | 10 | 3x3 | 1.89 | 4.49 | 2.68 | 2.99 | 3.45 | 0.37 | 6.06 | 0.933 | 1.9 | 12.31
25/08/2017 | 12MeV_10x10app_7x7field_PDD&Profiles_25082017 | 4297200% | TR2 | 100 | 12 | 10 | 7x7 | 2.98 | 5.11 | 3.74 | 4.02 | 4.39 | 1.46 | 6.16 | 0.781 | 2.89 | 12.51
28/08/2017 | 12MeV_15x15app_11x11field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 12 | 15 | 11x11 | 2.88 | 5 | 3.64 | 3.91 | 4.28 | 1.31 | 6.04 | 0.892 | 2.85 | 12.27
28/08/2017 | 12MeV_15x15app_15x15field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 12 | 15 | 15x15 | 2.75 | 4.99 | 3.62 | 3.89 | 4.26 | 1.34 | 6.02 | 0.881 | 2.85 | 12.23
28/08/2017 | 12MeV_20x20app_20x20field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 12 | 20 | 20x20 | 2.84 | 4.96 | 3.6 | 3.87 | 4.23 | 1.05 | 6 | 0.916 | 2.83 | 12.18
25/08/2017 | 12MeV_25x25app_21x21field_PDD&Profiles_25082017 | 4297200% | TR2 | 100 | 12 | 25 | 21x21 | 2.85 | 5.09 | 3.71 | 4 | 4.35 | 1.13 | 6.11 | 0.871 | 2.91 | 12.42
25/08/2017 | 12MeV_25x25app_25x25field_PDD&Profiles_25082017 | 4297200% | TR2 | 100 | 12 | 25 | 25x25 | 2.73 | 5.05 | 3.67 | 3.96 | 4.33 | 1.19 | 6.09 | 0.905 | 2.88 | 12.36
28/08/2017 | 16MeV_6x6app_3x3field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 16 | 6 | 3x3 | 1.58 | 5.64 | 3.19 | 3.65 | 4.25 |  | 7.83 | 0.982 | 1.75 | 15.87
28/08/2017 | 16MeV_6x6app_4x4field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 16 | 6 | 4x4 | 1.96 | 6.09 | 3.64 | 4.14 | 4.75 | 0.36 | 8.02 | 0.978 | 2.02 | 16.25
28/08/2017 | 16MeV_6x6app_6x6field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 16 | 6 | 6x6 | 2.9 | 6.53 | 4.34 | 4.82 | 5.42 | 0.44 | 8.02 | 0.959 | 2.61 | 16.26
24/08/2017 | 16MeV_10x10app_10x10field_PDD&Profiles_24082017 | 4297100% | TR2 | 100 | 16 | 10 | 10x10 | 2.92 | 6.67 | 4.64 | 5.11 | 5.64 | 0.48 | 8.06 | 0.874 | 2.81 | 16.35
28/08/2017 | 16MeV_10x10app_3x3field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 16 | 10 | 3x3 | 1.58 | 5.47 | 3.01 | 3.48 | 4.09 | 0.21 | 7.71 | 0.95 | 1.69 | 15.63
25/08/2017 | 16MeV_10x10app_7x7field_PDD&Profiles_25082017 | 4297200% | TR2 | 100 | 16 | 10 | 7x7 | 3.18 | 6.68 | 4.57 | 5.05 | 5.63 | 0.53 | 8.13 | 0.793 | 2.72 | 16.49
28/08/2017 | 16MeV_15x15app_11x11field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 16 | 15 | 11x11 | 3.03 | 6.61 | 4.63 | 5.07 | 5.6 | 0.49 | 7.98 | 0.916 | 2.82 | 16.17
28/08/2017 | 16MeV_15x15app_15x15field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 16 | 15 | 15x15 | 2.66 | 6.6 | 4.61 | 5.06 | 5.59 | 0.5 | 7.96 | 0.897 | 2.81 | 16.15
28/08/2017 | 16MeV_20x20app_20x20field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 16 | 20 | 20x20 | 2.75 | 6.56 | 4.54 | 5.01 | 5.56 | 0.41 | 7.95 | 0.926 | 2.77 | 16.12
25/08/2017 | 16MeV_25x25app_21x21field_PDD&Profiles_25082017 | 4297200% | TR2 | 100 | 16 | 25 | 21x21 | 3.14 | 6.69 | 4.67 | 5.12 | 5.67 | 0.51 | 8.06 | 0.923 | 2.83 | 16.35
25/08/2017 | 16MeV_25x25app_25x25field_PDD&Profiles_25082017 | 4297200% | TR2 | 100 | 16 | 25 | 25x25 | 3.39 | 6.66 | 4.66 | 5.1 | 5.66 | 0.55 | 8.03 | 0.921 | 2.84 | 16.27
28/08/2017 | 20MeV_6x6app_3x3field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 20 | 6 | 3x3 | 1.65 | 6.51 | 3.42 | 4.04 | 4.79 | 0.22 | 9.32 | 0.993 | 1.61 | 18.9
28/08/2017 | 20MeV_6x6app_4x4field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 20 | 6 | 4x4 | 1.78 | 7.13 | 3.85 | 4.56 | 5.36 | 0.24 | 9.8 | 0.986 | 1.77 | 19.87
28/08/2017 | 20MeV_6x6app_6x6field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 20 | 6 | 6x6 | 1.91 | 7.91 | 4.55 | 5.34 | 6.25 |  | 10.04 | 0.964 | 2.25 | 20.36
25/08/2017 | 20MeV_10x10app_10x10field_PDD&Profiles_25082017 | 4297200% | TR2 | 100 | 20 | 10 | 10x10 | 2.25 | 8.26 | 5.02 | 5.85 | 6.77 | 0.38 | 10.05 | 0.8 | 2.65 | 20.37
28/08/2017 | 20MeV_10x10app_3x3field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 20 | 10 | 3x3 | 1.54 | 6.35 | 3.32 | 3.88 | 4.64 | 0.1 | 9.16 | 0.98 | 1.59 | 18.57
25/08/2017 | 20MeV_10x10app_7x7field_PDD&Profiles_25082017 | 4297200% | TR2 | 100 | 20 | 10 | 7x7 | 2.02 | 8.14 | 4.74 | 5.61 | 6.52 | 0.39 | 10.13 | 0.81 | 2.41 | 20.54
28/08/2017 | 20MeV_15x15app_11x11field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 20 | 15 | 11x11 | 2.26 | 8.17 | 5.08 | 5.83 | 6.72 | 0.32 | 9.98 | 0.934 | 2.62 | 20.23
28/08/2017 | 20MeV_15x15app_15x15field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 20 | 15 | 15x15 | 2.61 | 8.16 | 5.07 | 5.86 | 6.72 | 0.35 | 9.96 | 0.907 | 2.62 | 20.19
28/08/2017 | 20MeV_20x20app_20x20field_PDD&Profiles_28082017 | 4297500% | TR2 | 100 | 20 | 20 | 20x20 | 2.87 | 8.14 | 4.99 | 5.8 | 6.68 | 0.32 | 9.93 | 0.929 | 2.62 | 20.13
25/08/2017 | 20MeV_25x25app_21x21field_PDD&Profiles_25082017 | 4297200% | TR2 | 100 | 20 | 25 | 21x21 | 2.49 | 8.26 | 5.1 | 5.91 | 6.8 | 0.42 | 10.05 | 0.932 | 2.65 | 20.36
25/08/2017 | 20MeV_25x25app_25x25field_PDD&Profiles_25082017 | 4297200% | TR2 | 100 | 20 | 25 | 25x25 | 2.39 | 8.23 | 5.14 | 5.91 | 6.79 | 0.41 | 10.01 | 0.926 | 2.65 | 20.3
29/08/2017 | 6MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017 | 4297600% | TR2 | 110 | 6 | 6 | 6x6 | 1.4 | 2.42 | 1.75 | 1.86 | 2.04 | 1 | 3 | 0.996 | 2.59 | 6.18
29/08/2017 | 6MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017 | 4298500% | TR2 | 110 | 6 | 10 | 10x10 | 1.38 | 2.4 | 1.71 | 1.84 | 2.02 | 0.98 | 2.97 | 0.762 | 2.61 | 6.12
29/08/2017 | 6MeV_15x15app_15x15field_PDD&Profiles_110SSD_29082017 | 4297600% | TR2 | 110 | 6 | 15 | 15x15 | 1.39 | 2.39 | 1.69 | 1.82 | 2 | 0.96 | 2.95 | 1.163 | 2.61 | 6.08
29/08/2017 | 6MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017 | 4297600% | TR2 | 110 | 6 | 20 | 20x20 | 1.29 | 2.36 | 1.66 | 1.78 | 1.97 | 0.93 | 2.93 | 1.178 | 2.56 | 6.04
29/08/2017 | 6MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017 | 4297600% | TR2 | 110 | 6 | 25 | 25x25 | 1.26 | 2.34 | 1.64 | 1.77 | 1.95 | 0.86 | 2.91 | 0.91 | 2.53 | 6
29/08/2017 | 9MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017 | 4297600% | TR2 | 110 | 9 | 6 | 6x6 | 2.16 | 3.62 | 2.64 | 2.82 | 3.09 | 1.46 | 4.4 | 1.039 | 2.78 | 8.98
29/08/2017 | 9MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017 | 4298500% | TR2 | 110 | 9 | 10 | 10x10 | 2.04 | 3.6 | 2.65 | 2.83 | 3.07 |  | 4.38 | 0.8 | 2.8 | 8.93
29/08/2017 | 9MeV_15x15app_15x15field_PDD&Profiles_SSD110_29082017 | 4297600% | TR2 | 110 | 9 | 15 | 15x15 | 2.15 | 3.58 | 2.64 | 2.81 | 3.06 | 1.41 | 4.36 | 1.113 | 2.79 | 8.9
29/08/2017 | 9MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017 | 4297600% | TR2 | 110 | 9 | 20 | 20x20 | 2.14 | 3.57 | 2.59 | 2.78 | 3.03 | 1.37 | 4.34 | 1.116 | 2.78 | 8.86
29/08/2017 | 9MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017 | 4298500% | TR2 | 110 | 9 | 25 | 25x25 | 2.03 | 3.54 | 2.55 | 2.75 | 3.01 | 1.32 | 4.31 | 0.821 | 2.75 | 8.81
29/08/2017 | 12MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017 | 4297600% | TR2 | 110 | 12 | 6 | 6x6 | 2.85 | 4.99 | 3.63 | 3.91 | 4.26 |  | 6.03 | 1.065 | 2.85 | 12.26
29/08/2017 | 12MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017 | 4298500% | TR2 | 110 | 12 | 10 | 10x10 | 2.85 | 4.98 | 3.64 | 3.92 | 4.27 |  | 6.02 | 0.836 | 2.87 | 12.22
29/08/2017 | 12MeV_15x15app_15x15field_PDD&Profiles_SSD110_29082017 | 4297600% | TR2 | 110 | 12 | 15 | 15x15 | 2.95 | 4.97 | 3.64 | 3.9 | 4.26 | 1.49 | 6.01 | 1.094 | 2.84 | 12.21
29/08/2017 | 12MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017 | 4297600% | TR2 | 110 | 12 | 20 | 20x20 | 2.97 | 4.95 | 3.58 | 3.85 | 4.22 | 1.4 | 5.98 | 1.097 | 2.84 | 12.15
29/08/2017 | 12MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017 | 4297600% | TR2 | 110 | 12 | 25 | 25x25 | 2.85 | 4.94 | 3.57 | 3.83 | 4.21 |  | 5.96 | 0.884 | 2.84 | 12.11
29/08/2017 | 16MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017 | 4297600% | TR2 | 110 | 16 | 6 | 6x6 | 3.04 | 6.61 | 4.52 | 4.97 | 5.54 | 0.85 | 8.06 | 1.069 | 2.68 | 16.35
29/08/2017 | 16MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017 | 4297600% | TR2 | 110 | 16 | 10 | 10x10 | 3.52 | 6.64 | 4.72 | 5.11 | 5.63 | 0.9 | 8.01 | 1.127 | 2.83 | 16.24
29/08/2017 | 16MeV_15x15app_15x15field_PDD&Profiles_110SSD_29082017 | 4297600% | TR2 | 110 | 16 | 15 | 15x15 | 2.92 | 6.63 | 4.68 | 5.09 | 5.61 | 0.84 | 8 | 1.117 | 2.81 | 16.23
29/08/2017 | 16MeV_20x20app_20x20field_PDD&Profiles_SSD110_29082017 | 4297600% | TR2 | 110 | 16 | 20 | 20x20 | 3.29 | 6.61 | 4.61 | 5.04 | 5.57 | 0.7 | 7.97 | 1.119 | 2.82 | 16.16
29/08/2017 | 16MeV_25x25app_25x25field_PDD&Profiles_SSD110_29082017 | 4297600% | TR2 | 110 | 16 | 25 | 25x25 | 3.28 | 6.59 | 4.57 | 4.99 | 5.56 | 0.65 | 7.95 | 0.9 | 2.82 | 16.11
29/08/2017 | 20MeV_6x6app_6x6field_PDD&Profiles_110SSD_29082017 | 4297600% | TR2 | 110 | 20 | 6 | 6x6 | 3.25 | 8.05 | 4.97 | 5.7 | 6.56 | 0.6 | 9.95 | 1.073 | 2.48 | 20.18
29/08/2017 | 20MeV_10x10app_10x10field_PDD&Profiles_110SSD_29082017 | 4297600% | TR2 | 110 | 20 | 10 | 10x10 | 3.12 | 8.18 | 5.29 | 5.98 | 6.81 | 0.62 | 9.88 | 1.145 | 2.77 | 20.02
29/08/2017 | 20MeV_15x15app_15x15field_PDD&Profiles_SSD110_29082017 | 4297600% | TR2 | 110 | 20 | 15 | 15x15 | 2.62 | 8.18 | 5.34 | 6.03 | 6.83 | 0.53 | 9.9 | 1.134 | 2.72 | 20.07
29/08/2017 | 20MeV_20x20app_20x20field_PDD&Profiles_110SSD_29082017 | 4297600% | TR2 | 110 | 20 | 20 | 20x20 | 2.89 | 8.15 | 5.23 | 5.95 | 6.78 |  | 9.9 | 1.132 | 2.7 | 20.06
29/08/2017 | 20MeV_25x25app_25x25field_PDD&Profiles_110SSD_29082017 | 4297600% | TR2 | 110 | 20 | 25 | 25x25 | 2.98 | 8.14 | 5.12 | 5.91 | 6.75 | 0.43 | 9.86 | 0.921 | 2.71 | 19.99

## R50
SSD | 100 | [15 empty] | SSD | 100
[18 empty] | R50 |  |  |  |  | Average Regional R50
Average of R50 |  | Energy | [14 empty] | Field Size | Energy |  |  |  |  | Energy
Field Size | Applicator | 6 | 9 | 12 | 16 | 20 | [11 empty] | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV | 6 MeV Fit | 9 MeV Fit | 12 MeV Fit | 16 MeV Fit | 20 MeV Fit
3x3 | 6 | 2.42 | 3.53 | 4.62 | 5.64 | 6.51 | [10 empty] | 3 | 2.385 | 3.49 | 4.555 | 5.555 | 6.43 | 2.4145 | 3.49 | 4.555 | 5.555 | 6.43
3x3 | 10 | 2.35 | 3.45 | 4.49 | 5.47 | 6.35 | [10 empty] | 4 | 2.43 | 3.61 | 4.86 | 6.09 | 7.13 | 2.4145 | 3.61 | 4.86 | 6.09 | 7.13
4x4 | 6 | 2.43 | 3.61 | 4.86 | 6.09 | 7.13 | [10 empty] | 6 | 2.42 | 3.62 | 4.99 | 6.53 | 7.91 | 2.4145 | 3.62 | 4.99 | 6.53 | 7.91
6x6 | 6 | 2.42 | 3.62 | 4.99 | 6.53 | 7.91 | [10 empty] | 7 | 2.52 | 3.74 | 5.11 | 6.68 | 8.14 | 2.4145 | 3.64 | 5.01 | 6.58 | 8.04
7x7 | 10 | 2.52 | 3.74 | 5.11 | 6.68 | 8.14 | [10 empty] | 10 | 2.42 | 3.64 | 5.03 | 6.67 | 8.26 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
10x10 | 10 | 2.42 | 3.64 | 5.03 | 6.67 | 8.26 | [10 empty] | 11 | 2.4 | 3.61 | 5 | 6.61 | 8.17 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
11x11 | 15 | 2.4 | 3.61 | 5 | 6.61 | 8.17 | [10 empty] | 15 | 2.38 | 3.6 | 4.99 | 6.6 | 8.16 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
15x15 | 15 | 2.38 | 3.6 | 4.99 | 6.6 | 8.16 | [10 empty] | 20 | 2.36 | 3.58 | 4.96 | 6.56 | 8.14 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
20x20 | 20 | 2.36 | 3.58 | 4.96 | 6.56 | 8.14 | [10 empty] | 21 | 2.48 | 3.7 | 5.09 | 6.69 | 8.26 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
21x21 | 25 | 2.48 | 3.7 | 5.09 | 6.69 | 8.26 | [10 empty] | 25 | 2.45 | 3.67 | 5.05 | 6.66 | 8.23 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
25x25 | 25 | 2.45 | 3.67 | 5.05 | 6.66 | 8.23
[17 empty] | Smoothing Parameters | R50 Fit Parameters
[17 empty] | Fit Parameter | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV
[17 empty] | a |  |  |  | -0.0006 | -0.0005
[17 empty] | b |  | 0.0108 | 0.0163 | 0.0235 | 0.0196
[17 empty] | c |  | -0.1792 | -0.2913 | -0.345 | -0.3128
[17 empty] | d |  | 0.9733 | 1.7425 | 2.1799 | 2.2587
[17 empty] | e |  | 1.89 | 1.51 | 1.5317 | 1.9748
[17 empty] | Max Field Size | 2 | 6 | 7 | 10 | 10
[17 empty] | Average | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
[17 empty] | Square Field Size | Calculated R50
[18 empty] | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV
[17 empty] | 3 | 2.4145 | 3.4887 | 4.5559 | 5.5523 | 6.4244
[17 empty] | 3 | 2.4145 | 3.4887 | 4.5559 | 5.5523 | 6.4244
[17 empty] | 4 | 2.4145 | 3.6072 | 4.8624 | 6.0817 | 7.1312
[17 empty] | 6 | 2.4145 | 3.6333 | 4.999 | 6.5 | 7.8518
[17 empty] | 7 | 2.4145 | 3.6333 | 5.02 | 6.5 | 7.9808
[17 empty] | 10 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
[17 empty] | 11 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
[17 empty] | 15 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
[17 empty] | 20 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
[17 empty] | 21 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
[17 empty] | 25 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
SSD | 100
Shifted Depth
Average of R50 |  | Energy |  |  |  |  | Energy |  |  |  |  | Calculated R50
Field Size | Applicator | 6 | 9 | 12 | 16 | 20 | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV
3x3 | 6 | 2.42 | 3.53 | 4.62 | 5.64 | 6.51 | -0.0055 | -0.0413 | -0.0641 | -0.0877 | -0.0856 | 2.4145 | 3.4887 | 4.5559 | 5.5523 | 6.4244
3x3 | 10 | 2.35 | 3.45 | 4.49 | 5.47 | 6.35 | 0.0645 | 0.0387 | 0.0659 | 0.0823 | 0.0744 | 2.4145 | 3.4887 | 4.5559 | 5.5523 | 6.4244
4x4 | 6 | 2.43 | 3.61 | 4.86 | 6.09 | 7.13 | -0.0155 | -0.0028 | 0.0024 | -0.0083 | 0.0012 | 2.4145 | 3.6072 | 4.8624 | 6.0817 | 7.1312
6x6 | 6 | 2.42 | 3.62 | 4.99 | 6.53 | 7.91 | -0.0055 | 0.013333 | 0.009 | -0.03 | -0.0582 | 2.4145 | 3.6333 | 4.999 | 6.5 | 7.8518
7x7 | 10 | 2.52 | 3.74 | 5.11 | 6.68 | 8.14 | -0.1055 | -0.10667 | -0.09 | -0.18 | -0.1592 | 2.4145 | 3.6333 | 5.02 | 6.5 | 7.9808
10x10 | 10 | 2.42 | 3.64 | 5.03 | 6.67 | 8.26 | -0.0055 | -0.0066667 | -0.01 | -0.038333 | -0.056667 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
11x11 | 15 | 2.4 | 3.61 | 5 | 6.61 | 8.17 | 0.0145 | 0.023333 | 0.02 | 0.021667 | 0.033333 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
15x15 | 15 | 2.38 | 3.6 | 4.99 | 6.6 | 8.16 | 0.0345 | 0.033333 | 0.03 | 0.031667 | 0.043333 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
20x20 | 20 | 2.36 | 3.58 | 4.96 | 6.56 | 8.14 | 0.0545 | 0.053333 | 0.06 | 0.071667 | 0.063333 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
21x21 | 25 | 2.48 | 3.7 | 5.09 | 6.69 | 8.26 | -0.0655 | -0.066667 | -0.07 | -0.058333 | -0.056667 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
25x25 | 25 | 2.45 | 3.67 | 5.05 | 6.66 | 8.23 | -0.0355 | -0.036667 | -0.03 | -0.028333 | -0.026667 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033

## d_max_fit
SSD | 100 | [15 empty] | SSD | 100
[18 empty] | R50 |  |  |  |  | Average Regional R50
Sum of R100 |  | Energy | [14 empty] | Field Size | Energy |  |  |  |  | Energy
Field Size | Applicator | 6 | 9 | 12 | 16 | 20 | [11 empty] | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV | 6 MeV Fit | 9 MeV Fit | 12 MeV Fit | 16 MeV Fit | 20 MeV Fit
3x3 | 6 | 1.27 | 1.65 | 2 | 1.58 | 1.65 | [10 empty] | 3 | 1.22 | 1.605 | 1.945 | 1.58 | 1.595 | 1.22 | 1.605 | 1.945 | 1.58 | 1.595
3x3 | 10 | 1.17 | 1.56 | 1.89 | 1.58 | 1.54 | [10 empty] | 4 | 1.4 | 1.92 | 2.25 | 1.96 | 1.78 | 1.4 | 1.92 | 2.25 | 1.96 | 1.78
4x4 | 6 | 1.4 | 1.92 | 2.25 | 1.96 | 1.78 | [10 empty] | 6 | 1.39 | 2.04 | 2.72 | 2.9 | 1.91 | 1.39 | 2.04 | 2.72 | 2.9 | 1.91
6x6 | 6 | 1.39 | 2.04 | 2.72 | 2.9 | 1.91 | [10 empty] | 7 | 1.5 | 2.26 | 2.98 | 3.18 | 2.02 | 1.5 | 2.26 | 2.98 | 3.18 | 2.02
7x7 | 10 | 1.5 | 2.26 | 2.98 | 3.18 | 2.02 | [10 empty] | 10 | 1.38 | 2.17 | 2.85 | 2.92 | 2.25 | 1.38 | 2.17 | 2.85 | 2.92 | 2.25
10x10 | 10 | 1.38 | 2.17 | 2.85 | 2.92 | 2.25 | [10 empty] | 11 | 1.38 | 2.16 | 2.88 | 3.03 | 2.26 | 1.38 | 2.16 | 2.88 | 3.03 | 2.26
11x11 | 15 | 1.38 | 2.16 | 2.88 | 3.03 | 2.26 | [10 empty] | 15 | 1.4 | 2.14 | 2.75 | 2.66 | 2.61 | 1.4 | 2.14 | 2.75 | 2.66 | 2.61
15x15 | 15 | 1.4 | 2.14 | 2.75 | 2.66 | 2.61 | [10 empty] | 20 | 1.29 | 2.02 | 2.84 | 2.75 | 2.87 | 1.29 | 2.02 | 2.84 | 2.75 | 2.87
20x20 | 20 | 1.29 | 2.02 | 2.84 | 2.75 | 2.87 | [10 empty] | 21 | 1.41 | 2.19 | 2.85 | 3.14 | 2.49 | 1.41 | 2.19 | 2.85 | 3.14 | 2.49
21x21 | 25 | 1.41 | 2.19 | 2.85 | 3.14 | 2.49 | [10 empty] | 25 | 1.4 | 2.04 | 2.73 | 3.39 | 2.39 | 1.4 | 2.04 | 2.73 | 3.39 | 2.39
25x25 | 25 | 1.4 | 2.04 | 2.73 | 3.39 | 2.39
[17 empty] | Smoothing Parameters | R50 Fit Parameters
[17 empty] | Fit Parameter | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV
[17 empty] | a |  |  |  | -0.0006 | -0.0005
[17 empty] | b |  | 0.0108 | 0.0163 | 0.0235 | 0.0196
[17 empty] | c |  | -0.1792 | -0.2913 | -0.345 | -0.3128
[17 empty] | d |  | 0.9733 | 1.7425 | 2.1799 | 2.2587
[17 empty] | e |  | 1.89 | 1.51 | 1.5317 | 1.9748
[17 empty] | Max Field Size | 2 | 6 | 7 | 10 | 10
[17 empty] | Average | 1.4 | 2.04 | 2.73 | 3.39 | 2.39
[17 empty] | Square Field Size | Calculated R50
[18 empty] | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV
[17 empty] | 3 | 1.4 | 3.4887 | 4.5559 | 5.5523 | 6.4244
[17 empty] | 3 | 1.4 | 3.4887 | 4.5559 | 5.5523 | 6.4244
[17 empty] | 4 | 1.4 | 3.6072 | 4.8624 | 6.0817 | 7.1312
[17 empty] | 6 | 1.4 | 2.04 | 4.999 | 6.5 | 7.8518
[17 empty] | 7 | 1.4 | 2.04 | 2.73 | 6.5 | 7.9808
[17 empty] | 10 | 1.4 | 2.04 | 2.73 | 3.39 | 2.39
[17 empty] | 11 | 1.4 | 2.04 | 2.73 | 3.39 | 2.39
[17 empty] | 15 | 1.4 | 2.04 | 2.73 | 3.39 | 2.39
[17 empty] | 20 | 1.4 | 2.04 | 2.73 | 3.39 | 2.39
[17 empty] | 21 | 1.4 | 2.04 | 2.73 | 3.39 | 2.39
[17 empty] | 25 | 1.4 | 2.04 | 2.73 | 3.39 | 2.39
SSD | 100
Shifted Depth
Average of R50 |  | Energy |  |  |  |  | Energy |  |  |  |  | Calculated R50
Field Size | Applicator | 6 | 9 | 12 | 16 | 20 | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV
3x3 | 6 | 2.42 | 3.53 | 4.62 | 5.64 | 6.51 | -1.02 | -0.0413 | -0.0641 | -0.0877 | -0.0856 | 2.4145 | 3.4887 | 4.5559 | 5.5523 | 6.4244
3x3 | 10 | 2.35 | 3.45 | 4.49 | 5.47 | 6.35 | -0.95 | 0.0387 | 0.0659 | 0.0823 | 0.0744 | 2.4145 | 3.4887 | 4.5559 | 5.5523 | 6.4244
4x4 | 6 | 2.43 | 3.61 | 4.86 | 6.09 | 7.13 | -1.03 | -0.0028 | 0.0024 | -0.0083 | 0.0012 | 2.4145 | 3.6072 | 4.8624 | 6.0817 | 7.1312
6x6 | 6 | 2.42 | 3.62 | 4.99 | 6.53 | 7.91 | -1.02 | -1.58 | 0.009 | -0.03 | -0.0582 | 2.4145 | 3.6333 | 4.999 | 6.5 | 7.8518
7x7 | 10 | 2.52 | 3.74 | 5.11 | 6.68 | 8.14 | -1.12 | -1.7 | -2.38 | -0.18 | -0.1592 | 2.4145 | 3.6333 | 5.02 | 6.5 | 7.9808
10x10 | 10 | 2.42 | 3.64 | 5.03 | 6.67 | 8.26 | -1.02 | -1.6 | -2.3 | -3.28 | -5.87 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
11x11 | 15 | 2.4 | 3.61 | 5 | 6.61 | 8.17 | -1 | -1.57 | -2.27 | -3.22 | -5.78 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
15x15 | 15 | 2.38 | 3.6 | 4.99 | 6.6 | 8.16 | -0.98 | -1.56 | -2.26 | -3.21 | -5.77 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
20x20 | 20 | 2.36 | 3.58 | 4.96 | 6.56 | 8.14 | -0.96 | -1.54 | -2.23 | -3.17 | -5.75 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
21x21 | 25 | 2.48 | 3.7 | 5.09 | 6.69 | 8.26 | -1.08 | -1.66 | -2.36 | -3.3 | -5.87 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033
25x25 | 25 | 2.45 | 3.67 | 5.05 | 6.66 | 8.23 | -1.05 | -1.63 | -2.32 | -3.27 | -5.84 | 2.4145 | 3.6333 | 5.02 | 6.6317 | 8.2033

## R50 110
SSD | 110 | [15 empty] | SSD | 110 |  | R50
Energy
Average of R50 |  | Energy | [14 empty] | Field Size | Applicator | Field Size | 6 Mev | 9 Mev | 12 Mev | 16 Mev | 20 Mev
Field Size | Applicator | 6 | 9 | 12 | 16 | 20 | [13 empty] | 6 | 9 | 12 | 16 | 20
6x6 | 6 | 2.42 | 3.62 | 4.99 | 6.61 | 8.05 | [10 empty] | 6x6 | 6 | 6 | 2.42 | 3.62 | 4.99 | 6.61 | 8.05
10x10 | 10 | 2.4 | 3.6 | 4.98 | 6.64 | 8.18 | [10 empty] | 10x10 | 10 | 10 | 2.4 | 3.6 | 4.98 | 6.64 | 8.18
15x15 | 15 | 2.39 | 3.58 | 4.97 | 6.63 | 8.18 | [10 empty] | 15x15 | 15 | 15 | 2.39 | 3.58 | 4.97 | 6.63 | 8.18
20x20 | 20 | 2.36 | 3.57 | 4.95 | 6.61 | 8.15 | [10 empty] | 20x20 | 20 | 20 | 2.36 | 3.57 | 4.95 | 6.61 | 8.15
25x25 | 25 | 2.34 | 3.54 | 4.94 | 6.59 | 8.14 | [10 empty] | 25x25 | 25 | 25 | 2.34 | 3.54 | 4.94 | 6.59 | 8.14
[19 empty] | Smoothing Parameters | R50 Fit Parameters
[19 empty] | Fit Parameter | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV
[19 empty] | a | -0.0041681 | -0.0039515 | -0.0027123 | -0.0034 | -0.003
[19 empty] | b | 2.4454 | 3.6421 | 5.0072 | 6.677 | 8.215
[19 empty] | Square Field Size | Calculated R50
[20 empty] | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV
[19 empty] | 6 | 2.4203 | 3.6184 | 4.991 | 6.6566 | 8.197
[19 empty] | 10 | 2.4037 | 3.6025 | 4.9801 | 6.643 | 8.185
[19 empty] | 15 | 2.3828 | 3.5828 | 4.9665 | 6.626 | 8.17
[19 empty] | 20 | 2.362 | 3.563 | 4.953 | 6.609 | 8.155
[19 empty] | 25 | 2.3412 | 3.5433 | 4.9394 | 6.592 | 8.14
SSD | 110
Shifted Depth
Average of R50 |  | Energy |  |  |  |  | Energy |  |  |  |  | Averaged R50
Field Size | Applicator | 6 | 9 | 12 | 16 | 20 | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV
6x6 | 6 | 2.42 | 3.62 | 4.99 | 6.61 | 8.05 | -0.038 | -0.038 | -0.024 | -0.0406 | -0.0345 | 2.382 | 3.582 | 4.966 | 6.5694 | 8.0155
10x10 | 10 | 2.4 | 3.6 | 4.98 | 6.64 | 8.18 | -0.018 | -0.018 | -0.014 | -0.024 | -0.0175 | 2.382 | 3.582 | 4.966 | 6.616 | 8.1625
15x15 | 15 | 2.39 | 3.58 | 4.97 | 6.63 | 8.18 | -0.008 | 0.002 | -0.004 | -0.014 | -0.0175 | 2.382 | 3.582 | 4.966 | 6.616 | 8.1625
20x20 | 20 | 2.36 | 3.57 | 4.95 | 6.61 | 8.15 | 0.022 | 0.012 | 0.016 | 0.006 | 0.0125 | 2.382 | 3.582 | 4.966 | 6.616 | 8.1625
25x25 | 25 | 2.34 | 3.54 | 4.94 | 6.59 | 8.14 | 0.042 | 0.042 | 0.026 | 0.026 | 0.0225 | 2.382 | 3.582 | 4.966 | 6.616 | 8.1625

## SSD 110 Dmax
SSD | 110
Average of R100 |  | Energy
Field Size | Applicator | 6 | 9 | 12 | 16 | 20
6x6 | 6 | 1.4 | 2.16 | 2.85 | 3.04 | 3.25
10x10 | 10 | 1.38 | 2.04 | 2.85 | 3.52 | 3.12
15x15 | 15 | 1.39 | 2.15 | 2.95 | 2.92 | 2.62
20x20 | 20 | 1.29 | 2.14 | 2.97 | 3.29 | 2.89
25x25 | 25 | 1.26 | 2.03 | 2.85 | 3.28 | 2.98

## SSD 110 PDDs
PDD
SSD: | 110 cm
Energy: | 6 MeV | 9 MeV | 12 MeV | 16 MeV |  | 20 MeV
Depth | 6x6+ | 6x6+ | 6x6+ | 6x6 | 10x10+ | 6x6 | 10x10+ | Energy | SSD | Field size inline | R100 | R50 | R95 | R90 | R80 | Rp | Ds
0 | 76.68% | 80.24% | 84.57% | 88.5% | 89.1% | 90.69% | 90.22% | 6 MeV | 110.00 cm | 6.00 cm | 1.35 cm | 2.38 cm | 1.69 cm | 1.82 cm | 2.00 cm | 2.95 cm | 77.1%
0.01 | 76.78% | 80.24% | 84.57% | 88.59% | 88.89% | 90.69% | 90.32% | 9 MeV | 110.00 cm | 6.00 cm | 2.12 cm | 3.58 cm | 2.62 cm | 2.80 cm | 3.05 cm | 4.35 cm | 80.4%
0.02 | 76.88% | 80.34% | 84.67% | 88.59% | 88.89% | 90.69% | 90.42% | 12 MeV | 110.00 cm | 6.00 cm | 2.90 cm | 4.97 cm | 3.63 cm | 3.89 cm | 4.25 cm | 6.00 cm | 84.7%
0.03 | 76.98% | 80.44% | 84.77% | 88.59% | 88.99% | 90.69% | 90.52% | 16 MeV | 110.00 cm | 6.00 cm | 2.97 cm | 6.57 cm | 4.49 cm | 4.93 cm | 5.51 cm | 8.01 cm | 88.5%
0.04 | 77.08% | 80.54% | 84.77% | 88.59% | 89.09% | 90.69% | 90.62% | 16 MeV | 110.00 cm | 10.00 cm | 3.39 cm | 6.62 cm | 4.66 cm | 5.07 cm | 5.61 cm | 7.98 cm | 89.1%
0.05 | 77.18% | 80.64% | 84.87% | 88.59% | 89.19% | 90.69% | 90.72% | 20 MeV | 110.00 cm | 6.00 cm | 3.16 cm | 8.02 cm | 4.97 cm | 5.68 cm | 6.53 cm | 9.93 cm | 90.6%
0.06 | 77.28% | 80.74% | 84.97% | 88.59% | 89.19% | 90.69% | 90.72% | 20 MeV | 110.00 cm | 10.00 cm | 3.11 cm | 8.16 cm | 5.27 cm | 5.98 cm | 6.79 cm | 9.87 cm | 90.9%
0.07 | 77.38% | 80.84% | 84.97% | 88.59% | 89.29% | 90.69% | 90.82%
0.08 | 77.48% | 80.94% | 85.07% | 88.59% | 89.39% | 90.69% | 90.92%
0.09 | 77.58% | 80.94% | 85.17% | 88.59% | 89.49% | 90.69% | 91.02%
0.1 | 77.68% | 81.04% | 85.27% | 88.59% | 89.59% | 90.69% | 91.12%
0.11 | 77.78% | 81.14% | 85.27% | 88.69% | 89.59% | 90.69% | 91.22%
0.12 | 77.88% | 81.24% | 85.37% | 88.89% | 89.69% | 90.79% | 91.32%
0.13 | 77.98% | 81.34% | 85.47% | 88.99% | 89.79% | 90.89% | 91.32%
0.14 | 78.08% | 81.44% | 85.47% | 89.09% | 89.89% | 90.99% | 91.42%
0.15 | 78.18% | 81.54% | 85.57% | 89.19% | 89.89% | 90.99% | 91.52%
0.16 | 78.28% | 81.64% | 85.67% | 89.39% | 89.99% | 91.09% | 91.62%
0.17 | 78.38% | 81.75% | 85.77% | 89.49% | 90.09% | 91.19% | 91.72%
0.18 | 78.58% | 81.75% | 85.77% | 89.59% | 90.19% | 91.29% | 91.82%
0.19 | 78.78% | 81.85% | 85.87% | 89.79% | 90.29% | 91.39% | 91.82%
0.2 | 78.98% | 81.95% | 85.97% | 89.89% | 90.29% | 91.49% | 91.92%
0.21 | 79.08% | 82.05% | 86.07% | 89.99% | 90.39% | 91.59% | 92.02%
0.22 | 79.28% | 82.15% | 86.27% | 90.09% | 90.49% | 91.69% | 92.12%
0.23 | 79.48% | 82.25% | 86.37% | 90.29% | 90.69% | 91.79% | 92.22%
0.24 | 79.68% | 82.45% | 86.47% | 90.39% | 90.79% | 91.99% | 92.32%
0.25 | 79.98% | 82.55% | 86.57% | 90.49% | 90.89% | 92.09% | 92.42%
0.26 | 80.18% | 82.65% | 86.77% | 90.59% | 90.99% | 92.19% | 92.51%
0.27 | 80.38% | 82.75% | 86.87% | 90.69% | 91.09% | 92.29% | 92.71%
0.28 | 80.58% | 82.95% | 86.97% | 90.89% | 91.19% | 92.39% | 92.81%
0.29 | 80.78% | 83.05% | 87.07% | 90.99% | 91.29% | 92.49% | 92.91%
0.3 | 80.98% | 83.15% | 87.27% | 91.09% | 91.49% | 92.59% | 93.11%
0.31 | 81.18% | 83.35% | 87.37% | 91.19% | 91.59% | 92.69% | 93.21%
0.32 | 81.48% | 83.45% | 87.47% | 91.29% | 91.69% | 92.79% | 93.31%
0.33 | 81.68% | 83.55% | 87.58% | 91.39% | 91.79% | 92.89% | 93.41%
0.34 | 81.88% | 83.65% | 87.68% | 91.49% | 91.89% | 92.99% | 93.61%
0.35 | 82.08% | 83.85% | 87.78% | 91.59% | 91.99% | 93.09% | 93.71%
0.36 | 82.28% | 83.95% | 87.88% | 91.69% | 92.09% | 93.19% | 93.81%
0.37 | 82.48% | 84.05% | 87.98% | 91.79% | 92.19% | 93.29% | 93.91%
0.38 | 82.68% | 84.15% | 88.08% | 91.89% | 92.29% | 93.39% | 93.91%
0.39 | 82.88% | 84.25% | 88.18% | 91.99% | 92.39% | 93.49% | 94.01%
0.4 | 83.08% | 84.45% | 88.28% | 92.09% | 92.39% | 93.59% | 94.11%
0.41 | 83.28% | 84.55% | 88.38% | 92.19% | 92.49% | 93.69% | 94.21%
0.42 | 83.48% | 84.65% | 88.48% | 92.29% | 92.59% | 93.79% | 94.21%
0.43 | 83.68% | 84.75% | 88.58% | 92.39% | 92.69% | 93.89% | 94.31%
0.44 | 83.88% | 84.85% | 88.68% | 92.49% | 92.79% | 93.99% | 94.41%
0.45 | 84.08% | 84.95% | 88.78% | 92.69% | 92.79% | 94.09% | 94.41%
0.46 | 84.28% | 85.16% | 88.88% | 92.69% | 92.89% | 94.19% | 94.51%
0.47 | 84.48% | 85.26% | 88.98% | 92.79% | 92.99% | 94.29% | 94.61%
0.48 | 84.68% | 85.36% | 88.98% | 92.89% | 93.09% | 94.29% | 94.71%
0.49 | 84.88% | 85.46% | 89.08% | 92.99% | 93.09% | 94.39% | 94.71%
0.5 | 85.09% | 85.56% | 89.18% | 93.09% | 93.19% | 94.49% | 94.81%
0.51 | 85.29% | 85.66% | 89.28% | 93.19% | 93.29% | 94.59% | 94.91%
0.52 | 85.59% | 85.76% | 89.38% | 93.29% | 93.39% | 94.69% | 94.91%
0.53 | 85.79% | 85.86% | 89.48% | 93.39% | 93.39% | 94.79% | 95.01%
0.54 | 85.99% | 85.96% | 89.58% | 93.39% | 93.49% | 94.89% | 95.11%
0.55 | 86.19% | 86.06% | 89.68% | 93.49% | 93.59% | 94.89% | 95.11%
0.56 | 86.49% | 86.16% | 89.78% | 93.59% | 93.69% | 94.99% | 95.21%
0.57 | 86.69% | 86.26% | 89.78% | 93.69% | 93.79% | 95.1% | 95.31%
0.58 | 86.89% | 86.46% | 89.88% | 93.69% | 93.79% | 95.2% | 95.31%
0.59 | 87.19% | 86.56% | 89.98% | 93.79% | 93.89% | 95.3% | 95.41%
0.6 | 87.39% | 86.66% | 90.08% | 93.89% | 93.99% | 95.3% | 95.51%
0.61 | 87.59% | 86.76% | 90.08% | 93.89% | 94.09% | 95.4% | 95.51%
0.62 | 87.89% | 86.86% | 90.18% | 93.99% | 94.09% | 95.5% | 95.61%
0.63 | 88.09% | 86.96% | 90.28% | 94.09% | 94.19% | 95.5% | 95.71%
0.64 | 88.29% | 87.06% | 90.28% | 94.19% | 94.29% | 95.6% | 95.71%
0.65 | 88.49% | 87.16% | 90.38% | 94.19% | 94.29% | 95.7% | 95.81%
0.66 | 88.79% | 87.26% | 90.48% | 94.29% | 94.39% | 95.7% | 95.91%
0.67 | 88.99% | 87.36% | 90.48% | 94.29% | 94.39% | 95.8% | 95.91%
0.68 | 89.19% | 87.46% | 90.58% | 94.39% | 94.49% | 95.8% | 96.01%
0.69 | 89.39% | 87.56% | 90.68% | 94.49% | 94.49% | 95.9% | 96.01%
0.7 | 89.59% | 87.66% | 90.68% | 94.49% | 94.59% | 96% | 96.11%
0.71 | 89.79% | 87.76% | 90.78% | 94.59% | 94.69% | 96% | 96.21%
0.72 | 90.09% | 87.86% | 90.88% | 94.59% | 94.69% | 96.1% | 96.21%
0.73 | 90.29% | 87.96% | 90.98% | 94.69% | 94.79% | 96.1% | 96.31%
0.74 | 90.49% | 88.06% | 90.98% | 94.69% | 94.79% | 96.2% | 96.31%
0.75 | 90.69% | 88.16% | 91.08% | 94.79% | 94.89% | 96.2% | 96.41%
0.76 | 90.89% | 88.26% | 91.18% | 94.79% | 94.99% | 96.3% | 96.41%
0.77 | 91.09% | 88.37% | 91.18% | 94.89% | 94.99% | 96.3% | 96.51%
0.78 | 91.29% | 88.47% | 91.28% | 94.89% | 95.1% | 96.4% | 96.61%
0.79 | 91.49% | 88.57% | 91.38% | 94.99% | 95.1% | 96.4% | 96.61%
0.8 | 91.69% | 88.67% | 91.38% | 94.99% | 95.2% | 96.5% | 96.71%
0.81 | 91.99% | 88.77% | 91.48% | 95.1% | 95.2% | 96.5% | 96.71%
0.82 | 92.19% | 88.87% | 91.48% | 95.1% | 95.3% | 96.6% | 96.81%
0.83 | 92.39% | 88.97% | 91.58% | 95.2% | 95.3% | 96.6% | 96.81%
0.84 | 92.59% | 89.07% | 91.58% | 95.2% | 95.4% | 96.7% | 96.91%
0.85 | 92.79% | 89.17% | 91.68% | 95.3% | 95.4% | 96.7% | 96.91%
0.86 | 92.99% | 89.27% | 91.68% | 95.3% | 95.5% | 96.8% | 96.91%
0.87 | 93.19% | 89.37% | 91.78% | 95.4% | 95.5% | 96.8% | 97.01%
0.88 | 93.49% | 89.47% | 91.78% | 95.4% | 95.5% | 96.9% | 97.01%
0.89 | 93.69% | 89.57% | 91.78% | 95.4% | 95.6% | 96.9% | 97.11%
0.9 | 93.89% | 89.67% | 91.88% | 95.5% | 95.6% | 96.9% | 97.11%
0.91 | 94.09% | 89.77% | 91.88% | 95.5% | 95.6% | 97% | 97.11%
0.92 | 94.39% | 89.87% | 91.98% | 95.6% | 95.7% | 97% | 97.21%
0.93 | 94.59% | 89.97% | 92.08% | 95.6% | 95.7% | 97.1% | 97.21%
0.94 | 94.79% | 90.07% | 92.08% | 95.7% | 95.8% | 97.1% | 97.31%
0.95 | 94.99% | 90.17% | 92.18% | 95.7% | 95.8% | 97.1% | 97.31%
0.96 | 95.2% | 90.27% | 92.18% | 95.7% | 95.8% | 97.2% | 97.31%
0.97 | 95.5% | 90.37% | 92.28% | 95.8% | 95.9% | 97.2% | 97.41%
0.98 | 95.7% | 90.47% | 92.38% | 95.8% | 95.9% | 97.2% | 97.41%
0.99 | 95.9% | 90.57% | 92.38% | 95.8% | 96% | 97.3% | 97.41%
1 | 96.1% | 90.67% | 92.48% | 95.9% | 96% | 97.3% | 97.41%
1.01 | 96.3% | 90.77% | 92.48% | 95.9% | 96% | 97.4% | 97.5%
1.02 | 96.5% | 90.97% | 92.59% | 95.9% | 96.1% | 97.4% | 97.5%
1.03 | 96.7% | 91.07% | 92.59% | 96% | 96.1% | 97.4% | 97.5%
1.04 | 96.9% | 91.17% | 92.69% | 96% | 96.2% | 97.5% | 97.5%
1.05 | 97.1% | 91.27% | 92.69% | 96% | 96.2% | 97.5% | 97.5%
1.06 | 97.2% | 91.37% | 92.79% | 96.1% | 96.2% | 97.5% | 97.6%
1.07 | 97.4% | 91.47% | 92.79% | 96.1% | 96.3% | 97.6% | 97.6%
1.08 | 97.6% | 91.57% | 92.89% | 96.1% | 96.3% | 97.6% | 97.6%
1.09 | 97.7% | 91.68% | 92.89% | 96.2% | 96.3% | 97.6% | 97.6%
1.1 | 97.9% | 91.78% | 92.89% | 96.2% | 96.4% | 97.6% | 97.7%
1.11 | 98% | 91.88% | 92.99% | 96.2% | 96.4% | 97.7% | 97.7%
1.12 | 98.2% | 91.98% | 92.99% | 96.3% | 96.4% | 97.7% | 97.8%
1.13 | 98.3% | 92.08% | 93.09% | 96.3% | 96.4% | 97.7% | 97.8%
1.14 | 98.5% | 92.18% | 93.09% | 96.3% | 96.5% | 97.8% | 97.9%
1.15 | 98.6% | 92.28% | 93.19% | 96.4% | 96.5% | 97.8% | 97.9%
1.16 | 98.7% | 92.38% | 93.29% | 96.4% | 96.5% | 97.8% | 98%
1.17 | 98.9% | 92.48% | 93.29% | 96.4% | 96.6% | 97.9% | 98%
1.18 | 99% | 92.58% | 93.39% | 96.5% | 96.6% | 97.9% | 98.1%
1.19 | 99.1% | 92.68% | 93.39% | 96.5% | 96.6% | 97.9% | 98.1%
1.2 | 99.2% | 92.88% | 93.49% | 96.5% | 96.7% | 97.9% | 98.1%
1.21 | 99.3% | 92.98% | 93.49% | 96.6% | 96.7% | 98% | 98.2%
1.22 | 99.4% | 93.08% | 93.59% | 96.6% | 96.7% | 98% | 98.2%
1.23 | 99.5% | 93.18% | 93.69% | 96.6% | 96.8% | 98% | 98.2%
1.24 | 99.6% | 93.28% | 93.69% | 96.7% | 96.8% | 98.1% | 98.2%
1.25 | 99.7% | 93.38% | 93.79% | 96.7% | 96.8% | 98.1% | 98.3%
1.26 | 99.8% | 93.48% | 93.79% | 96.8% | 96.9% | 98.1% | 98.3%
1.27 | 99.8% | 93.68% | 93.89% | 96.8% | 96.9% | 98.2% | 98.3%
1.28 | 99.9% | 93.78% | 93.99% | 96.8% | 96.9% | 98.2% | 98.3%
1.29 | 99.9% | 93.88% | 93.99% | 96.9% | 97% | 98.2% | 98.3%
1.3 | 99.9% | 93.98% | 94.09% | 96.9% | 97% | 98.3% | 98.3%
1.31 | 99.9% | 94.08% | 94.09% | 96.9% | 97% | 98.3% | 98.3%
1.32 | 100% | 94.18% | 94.19% | 97% | 97% | 98.3% | 98.3%
1.33 | 100% | 94.28% | 94.29% | 97% | 97.1% | 98.3% | 98.3%
1.34 | 100% | 94.38% | 94.29% | 97% | 97.1% | 98.4% | 98.4%
1.35 | 100% | 94.48% | 94.39% | 97.1% | 97.1% | 98.4% | 98.4%
1.36 | 100% | 94.58% | 94.39% | 97.1% | 97.2% | 98.4% | 98.4%
1.37 | 100% | 94.68% | 94.49% | 97.1% | 97.2% | 98.5% | 98.4%
1.38 | 99.9% | 94.78% | 94.49% | 97.2% | 97.2% | 98.5% | 98.5%
1.39 | 99.9% | 94.88% | 94.49% | 97.2% | 97.3% | 98.5% | 98.5%
1.4 | 99.9% | 94.98% | 94.59% | 97.2% | 97.3% | 98.5% | 98.5%
1.41 | 99.8% | 95.09% | 94.59% | 97.3% | 97.3% | 98.6% | 98.5%
1.42 | 99.8% | 95.19% | 94.69% | 97.3% | 97.3% | 98.6% | 98.6%
1.43 | 99.7% | 95.29% | 94.69% | 97.3% | 97.4% | 98.6% | 98.6%
1.44 | 99.7% | 95.39% | 94.79% | 97.4% | 97.4% | 98.7% | 98.6%
1.45 | 99.6% | 95.49% | 94.79% | 97.4% | 97.4% | 98.7% | 98.6%
1.46 | 99.5% | 95.59% | 94.89% | 97.4% | 97.4% | 98.7% | 98.7%
1.47 | 99.4% | 95.69% | 94.89% | 97.5% | 97.5% | 98.7% | 98.7%
1.48 | 99.3% | 95.79% | 94.99% | 97.5% | 97.5% | 98.8% | 98.7%
1.49 | 99.2% | 95.89% | 94.99% | 97.5% | 97.5% | 98.8% | 98.8%
1.5 | 99.1% | 96.09% | 95.09% | 97.5% | 97.5% | 98.8% | 98.8%
1.51 | 98.9% | 96.19% | 95.09% | 97.6% | 97.5% | 98.8% | 98.8%
1.52 | 98.8% | 96.29% | 95.19% | 97.6% | 97.5% | 98.8% | 98.9%
1.53 | 98.6% | 96.39% | 95.19% | 97.6% | 97.6% | 98.9% | 98.9%
1.54 | 98.5% | 96.49% | 95.29% | 97.7% | 97.6% | 98.9% | 98.9%
1.55 | 98.3% | 96.59% | 95.29% | 97.7% | 97.6% | 98.9% | 99%
1.56 | 98.1% | 96.69% | 95.29% | 97.7% | 97.7% | 98.9% | 99%
1.57 | 97.9% | 96.79% | 95.39% | 97.8% | 97.7% | 98.9% | 99%
1.58 | 97.7% | 96.79% | 95.39% | 97.8% | 97.7% | 99% | 99%
1.59 | 97.5% | 96.89% | 95.49% | 97.8% | 97.8% | 99% | 99%
1.6 | 97.3% | 96.99% | 95.49% | 97.8% | 97.8% | 99% | 99%
1.61 | 97.1% | 97.09% | 95.59% | 97.9% | 97.8% | 99% | 99%
1.62 | 96.9% | 97.19% | 95.59% | 97.9% | 97.9% | 99% | 99.1%
1.63 | 96.7% | 97.29% | 95.69% | 97.9% | 97.9% | 99% | 99.1%
1.64 | 96.4% | 97.29% | 95.69% | 98% | 97.9% | 99.1% | 99.1%
1.65 | 96.2% | 97.39% | 95.79% | 98% | 97.9% | 99.1% | 99.1%
1.66 | 95.9% | 97.49% | 95.79% | 98% | 98% | 99.1% | 99.1%
1.67 | 95.6% | 97.59% | 95.89% | 98.1% | 98% | 99.1% | 99.1%
1.68 | 95.3% | 97.69% | 95.89% | 98.1% | 98% | 99.1% | 99.1%
1.69 | 94.99% | 97.79% | 95.99% | 98.1% | 98% | 99.1% | 99.2%
1.7 | 94.59% | 97.89% | 95.99% | 98.2% | 98% | 99.2% | 99.2%
1.71 | 94.29% | 97.89% | 96.09% | 98.2% | 98% | 99.2% | 99.2%
1.72 | 93.99% | 97.99% | 96.09% | 98.2% | 98.1% | 99.2% | 99.2%
1.73 | 93.59% | 98.09% | 96.19% | 98.2% | 98.1% | 99.2% | 99.3%
1.74 | 93.19% | 98.19% | 96.19% | 98.3% | 98.1% | 99.2% | 99.3%
1.75 | 92.79% | 98.29% | 96.19% | 98.3% | 98.1% | 99.2% | 99.3%
1.76 | 92.39% | 98.4% | 96.29% | 98.3% | 98.1% | 99.3% | 99.3%
1.77 | 91.99% | 98.5% | 96.29% | 98.3% | 98.1% | 99.3% | 99.4%
1.78 | 91.59% | 98.6% | 96.39% | 98.4% | 98.1% | 99.3% | 99.4%
1.79 | 91.19% | 98.7% | 96.39% | 98.4% | 98.2% | 99.3% | 99.4%
1.8 | 90.69% | 98.7% | 96.39% | 98.4% | 98.2% | 99.3% | 99.4%
1.81 | 90.29% | 98.8% | 96.49% | 98.4% | 98.2% | 99.3% | 99.4%
1.82 | 89.79% | 98.9% | 96.49% | 98.5% | 98.2% | 99.3% | 99.4%
1.83 | 89.39% | 98.9% | 96.59% | 98.5% | 98.3% | 99.3% | 99.4%
1.84 | 88.89% | 99% | 96.59% | 98.5% | 98.3% | 99.4% | 99.4%
1.85 | 88.49% | 99% | 96.69% | 98.5% | 98.3% | 99.4% | 99.4%
1.86 | 87.99% | 99.1% | 96.69% | 98.5% | 98.4% | 99.4% | 99.4%
1.87 | 87.49% | 99.2% | 96.69% | 98.6% | 98.4% | 99.4% | 99.4%
1.88 | 86.99% | 99.2% | 96.79% | 98.6% | 98.4% | 99.4% | 99.5%
1.89 | 86.39% | 99.3% | 96.79% | 98.6% | 98.4% | 99.4% | 99.5%
1.9 | 85.89% | 99.3% | 96.89% | 98.6% | 98.5% | 99.4% | 99.5%
1.91 | 85.29% | 99.4% | 96.89% | 98.6% | 98.5% | 99.5% | 99.5%
1.92 | 84.68% | 99.5% | 96.99% | 98.6% | 98.5% | 99.5% | 99.5%
1.93 | 84.08% | 99.5% | 96.99% | 98.7% | 98.6% | 99.5% | 99.5%
1.94 | 83.48% | 99.6% | 96.99% | 98.7% | 98.6% | 99.5% | 99.5%
1.95 | 82.88% | 99.6% | 97.09% | 98.7% | 98.6% | 99.5% | 99.5%
1.96 | 82.28% | 99.6% | 97.09% | 98.7% | 98.6% | 99.5% | 99.5%
1.97 | 81.68% | 99.7% | 97.09% | 98.7% | 98.6% | 99.5% | 99.6%
1.98 | 81.08% | 99.7% | 97.19% | 98.8% | 98.7% | 99.5% | 99.6%
1.99 | 80.38% | 99.8% | 97.19% | 98.8% | 98.7% | 99.5% | 99.6%
2 | 79.78% | 99.8% | 97.19% | 98.8% | 98.7% | 99.6% | 99.6%
2.01 | 79.08% | 99.8% | 97.29% | 98.8% | 98.7% | 99.6% | 99.6%
2.02 | 78.38% | 99.9% | 97.29% | 98.9% | 98.7% | 99.6% | 99.7%
2.03 | 77.78% | 99.9% | 97.39% | 98.9% | 98.7% | 99.6% | 99.7%
2.04 | 77.08% | 99.9% | 97.39% | 98.9% | 98.7% | 99.6% | 99.7%
2.05 | 76.38% | 99.9% | 97.49% | 98.9% | 98.8% | 99.6% | 99.7%
2.06 | 75.68% | 100% | 97.49% | 98.9% | 98.8% | 99.6% | 99.7%
2.07 | 74.97% | 100% | 97.6% | 99% | 98.8% | 99.6% | 99.7%
2.08 | 74.27% | 100% | 97.6% | 99% | 98.8% | 99.6% | 99.7%
2.09 | 73.57% | 100% | 97.7% | 99% | 98.8% | 99.6% | 99.7%
2.1 | 72.87% | 100% | 97.7% | 99% | 98.8% | 99.6% | 99.7%
2.11 | 72.07% | 100% | 97.8% | 99% | 98.8% | 99.6% | 99.7%
2.12 | 71.37% | 100% | 97.8% | 99.1% | 98.8% | 99.6% | 99.7%
2.13 | 70.57% | 100% | 97.9% | 99.1% | 98.8% | 99.6% | 99.7%
2.14 | 69.87% | 100% | 97.9% | 99.1% | 98.8% | 99.6% | 99.7%
2.15 | 69.07% | 100% | 97.9% | 99.1% | 98.9% | 99.6% | 99.7%
2.16 | 68.27% | 100% | 98% | 99.1% | 98.9% | 99.7% | 99.7%
2.17 | 67.57% | 100% | 98% | 99.1% | 98.9% | 99.7% | 99.7%
2.18 | 66.77% | 99.9% | 98.1% | 99.1% | 98.9% | 99.7% | 99.7%
2.19 | 65.97% | 99.9% | 98.1% | 99.2% | 98.9% | 99.7% | 99.7%
2.2 | 65.27% | 99.9% | 98.2% | 99.2% | 98.9% | 99.7% | 99.7%
2.21 | 64.46% | 99.8% | 98.2% | 99.2% | 99% | 99.7% | 99.7%
2.22 | 63.66% | 99.8% | 98.2% | 99.2% | 99% | 99.7% | 99.7%
2.23 | 62.86% | 99.8% | 98.3% | 99.2% | 99% | 99.7% | 99.7%
2.24 | 62.06% | 99.7% | 98.3% | 99.2% | 99% | 99.7% | 99.7%
2.25 | 61.26% | 99.7% | 98.4% | 99.2% | 99.1% | 99.7% | 99.7%
2.26 | 60.46% | 99.6% | 98.4% | 99.2% | 99.1% | 99.7% | 99.7%
2.27 | 59.66% | 99.6% | 98.5% | 99.2% | 99.1% | 99.7% | 99.7%
2.28 | 58.76% | 99.5% | 98.5% | 99.3% | 99.1% | 99.7% | 99.7%
2.29 | 57.96% | 99.5% | 98.6% | 99.3% | 99.1% | 99.7% | 99.8%
2.3 | 57.06% | 99.4% | 98.6% | 99.3% | 99.1% | 99.7% | 99.8%
2.31 | 56.26% | 99.3% | 98.6% | 99.3% | 99.1% | 99.7% | 99.8%
2.32 | 55.36% | 99.2% | 98.7% | 99.3% | 99.2% | 99.7% | 99.8%
2.33 | 54.55% | 99.2% | 98.7% | 99.3% | 99.2% | 99.7% | 99.8%
2.34 | 53.65% | 99.1% | 98.7% | 99.3% | 99.2% | 99.7% | 99.8%
2.35 | 52.75% | 99% | 98.8% | 99.4% | 99.2% | 99.8% | 99.8%
2.36 | 51.85% | 98.9% | 98.8% | 99.4% | 99.2% | 99.8% | 99.8%
2.37 | 50.95% | 98.8% | 98.8% | 99.4% | 99.2% | 99.8% | 99.8%
2.38 | 50.05% | 98.7% | 98.9% | 99.4% | 99.2% | 99.8% | 99.8%
2.39 | 49.15% | 98.6% | 98.9% | 99.4% | 99.2% | 99.8% | 99.8%
2.4 | 48.25% | 98.5% | 98.9% | 99.4% | 99.2% | 99.8% | 99.8%
2.41 | 47.35% | 98.4% | 98.9% | 99.4% | 99.2% | 99.8% | 99.8%
2.42 | 46.45% | 98.19% | 99% | 99.5% | 99.2% | 99.8% | 99.8%
2.43 | 45.55% | 98.09% | 99% | 99.5% | 99.3% | 99.8% | 99.9%
2.44 | 44.64% | 97.99% | 99% | 99.5% | 99.3% | 99.8% | 99.9%
2.45 | 43.74% | 97.79% | 99% | 99.5% | 99.3% | 99.8% | 99.9%
2.46 | 42.94% | 97.69% | 99.1% | 99.5% | 99.3% | 99.8% | 99.9%
2.47 | 42.04% | 97.59% | 99.1% | 99.5% | 99.3% | 99.8% | 99.9%
2.48 | 41.24% | 97.39% | 99.1% | 99.5% | 99.3% | 99.8% | 99.9%
2.49 | 40.34% | 97.29% | 99.1% | 99.6% | 99.4% | 99.8% | 99.9%
2.5 | 39.54% | 97.09% | 99.2% | 99.6% | 99.4% | 99.9% | 99.9%
2.51 | 38.74% | 96.99% | 99.2% | 99.6% | 99.4% | 99.9% | 99.9%
2.52 | 37.94% | 96.79% | 99.2% | 99.6% | 99.4% | 99.9% | 99.9%
2.53 | 37.14% | 96.59% | 99.2% | 99.6% | 99.5% | 99.9% | 99.9%
2.54 | 36.34% | 96.39% | 99.3% | 99.6% | 99.5% | 99.9% | 99.9%
2.55 | 35.54% | 96.29% | 99.3% | 99.6% | 99.5% | 99.9% | 99.9%
2.56 | 34.73% | 96.09% | 99.3% | 99.6% | 99.6% | 99.9% | 99.9%
2.57 | 33.93% | 95.89% | 99.3% | 99.7% | 99.6% | 99.9% | 99.9%
2.58 | 33.13% | 95.69% | 99.4% | 99.7% | 99.6% | 99.9% | 99.9%
2.59 | 32.33% | 95.49% | 99.4% | 99.7% | 99.7% | 99.9% | 99.9%
2.6 | 31.53% | 95.39% | 99.4% | 99.7% | 99.7% | 99.9% | 100%
2.61 | 30.73% | 95.19% | 99.4% | 99.7% | 99.7% | 99.9% | 100%
2.62 | 30.03% | 94.98% | 99.4% | 99.7% | 99.7% | 99.9% | 100%
2.63 | 29.23% | 94.78% | 99.4% | 99.7% | 99.7% | 99.9% | 100%
2.64 | 28.43% | 94.48% | 99.5% | 99.7% | 99.8% | 99.9% | 99.9%
2.65 | 27.73% | 94.28% | 99.5% | 99.7% | 99.8% | 99.9% | 99.9%
2.66 | 26.93% | 93.98% | 99.5% | 99.7% | 99.8% | 99.9% | 99.9%
2.67 | 26.23% | 93.78% | 99.5% | 99.8% | 99.8% | 99.9% | 99.9%
2.68 | 25.43% | 93.48% | 99.6% | 99.8% | 99.8% | 100% | 99.9%
2.69 | 24.72% | 93.28% | 99.6% | 99.8% | 99.8% | 100% | 99.9%
2.7 | 23.92% | 92.98% | 99.6% | 99.8% | 99.8% | 100% | 99.9%
2.71 | 23.22% | 92.78% | 99.6% | 99.8% | 99.8% | 100% | 99.9%
2.72 | 22.52% | 92.48% | 99.7% | 99.8% | 99.7% | 100% | 99.9%
2.73 | 21.92% | 92.18% | 99.7% | 99.8% | 99.7% | 100% | 99.9%
2.74 | 21.22% | 91.88% | 99.7% | 99.8% | 99.7% | 100% | 99.9%
2.75 | 20.52% | 91.57% | 99.7% | 99.8% | 99.7% | 100% | 99.9%
2.76 | 19.92% | 91.27% | 99.8% | 99.8% | 99.7% | 100% | 99.9%
2.77 | 19.22% | 90.97% | 99.8% | 99.9% | 99.7% | 100% | 99.9%
2.78 | 18.62% | 90.67% | 99.8% | 99.9% | 99.7% | 100% | 99.8%
2.79 | 17.92% | 90.27% | 99.8% | 99.9% | 99.7% | 100% | 99.8%
2.8 | 17.32% | 89.97% | 99.9% | 99.9% | 99.7% | 100% | 99.8%
2.81 | 16.72% | 89.67% | 99.9% | 99.9% | 99.7% | 99.9% | 99.8%
2.82 | 16.12% | 89.27% | 99.9% | 99.9% | 99.7% | 99.9% | 99.8%
2.83 | 15.52% | 88.97% | 99.9% | 99.9% | 99.7% | 99.9% | 99.8%
2.84 | 14.91% | 88.57% | 99.9% | 99.9% | 99.7% | 99.9% | 99.8%
2.85 | 14.41% | 88.26% | 99.9% | 99.9% | 99.8% | 99.9% | 99.8%
2.86 | 13.81% | 87.86% | 99.9% | 99.9% | 99.8% | 99.9% | 99.8%
2.87 | 13.31% | 87.56% | 99.9% | 99.9% | 99.8% | 99.9% | 99.8%
2.88 | 12.81% | 87.16% | 99.9% | 99.9% | 99.8% | 99.9% | 99.8%
2.89 | 12.41% | 86.76% | 100% | 99.9% | 99.8% | 99.9% | 99.8%
2.9 | 11.91% | 86.36% | 100% | 100% | 99.8% | 99.9% | 99.8%
2.91 | 11.41% | 85.96% | 100% | 100% | 99.8% | 99.9% | 99.9%
2.92 | 10.91% | 85.56% | 100% | 100% | 99.9% | 99.9% | 99.9%
2.93 | 10.51% | 85.16% | 100% | 100% | 99.9% | 99.9% | 99.9%
2.94 | 10.01% | 84.75% | 100% | 100% | 99.9% | 99.9% | 99.9%
2.95 | 9.61% | 84.35% | 99.9% | 100% | 99.9% | 99.9% | 99.9%
2.96 | 9.209% | 83.95% | 99.9% | 100% | 99.9% | 99.9% | 99.9%
2.97 | 8.809% | 83.55% | 99.9% | 100% | 99.9% | 99.9% | 99.9%
2.98 | 8.509% | 83.15% | 99.9% | 100% | 99.9% | 99.9% | 99.9%
2.99 | 8.108% | 82.65% | 99.9% | 100% | 99.9% | 99.9% | 99.9%
3 | 7.708% | 82.25% | 99.9% | 100% | 99.9% | 99.9% | 100%
3.01 | 7.407% | 81.85% | 99.9% | 100% | 99.8% | 99.9% | 100%
3.02 | 7.107% | 81.34% | 99.8% | 100% | 99.8% | 99.9% | 100%
3.03 | 6.707% | 80.94% | 99.8% | 100% | 99.8% | 99.9% | 100%
3.04 | 6.406% | 80.54% | 99.8% | 100% | 99.8% | 99.9% | 100%
3.05 | 6.106% | 80.04% | 99.7% | 100% | 99.8% | 99.9% | 100%
3.06 | 5.806% | 79.64% | 99.7% | 100% | 99.8% | 99.9% | 100%
3.07 | 5.506% | 79.14% | 99.6% | 100% | 99.8% | 100% | 100%
3.08 | 5.305% | 78.74% | 99.6% | 100% | 99.9% | 100% | 100%
3.09 | 5.005% | 78.23% | 99.6% | 100% | 99.9% | 100% | 100%
3.1 | 4.805% | 77.83% | 99.5% | 99.9% | 99.9% | 100% | 100%
3.11 | 4.605% | 77.33% | 99.5% | 99.9% | 99.9% | 100% | 100%
3.12 | 4.304% | 76.83% | 99.4% | 99.9% | 99.9% | 100% | 100%
3.13 | 4.204% | 76.43% | 99.4% | 99.9% | 99.9% | 100% | 100%
3.14 | 4.004% | 75.93% | 99.4% | 99.9% | 99.9% | 100% | 100%
3.15 | 3.804% | 75.43% | 99.3% | 99.9% | 99.9% | 100% | 99.9%
3.16 | 3.604% | 74.92% | 99.3% | 99.9% | 99.9% | 100% | 99.9%
3.17 | 3.403% | 74.32% | 99.3% | 99.9% | 99.9% | 100% | 99.9%
3.18 | 3.203% | 73.82% | 99.2% | 99.9% | 99.9% | 100% | 99.9%
3.19 | 3.103% | 73.32% | 99.2% | 99.9% | 99.9% | 100% | 99.9%
3.2 | 2.903% | 72.82% | 99.2% | 99.9% | 99.9% | 100% | 99.9%
3.21 | 2.803% | 72.32% | 99.1% | 99.9% | 99.9% | 100% | 99.9%
3.22 | 2.703% | 71.72% | 99.1% | 99.9% | 99.9% | 100% | 99.9%
3.23 | 2.603% | 71.21% | 99% | 99.9% | 99.9% | 100% | 99.9%
3.24 | 2.402% | 70.61% | 98.9% | 99.9% | 99.9% | 100% | 99.8%
3.25 | 2.302% | 70.01% | 98.9% | 99.9% | 99.9% | 100% | 99.8%
3.26 | 2.202% | 69.41% | 98.8% | 99.9% | 99.9% | 100% | 99.8%
3.27 | 2.102% | 68.81% | 98.7% | 99.8% | 99.9% | 100% | 99.8%
3.28 | 2.002% | 68.3% | 98.6% | 99.8% | 99.9% | 99.9% | 99.8%
3.29 | 2.002% | 67.7% | 98.6% | 99.8% | 99.9% | 99.9% | 99.8%
3.3 | 1.902% | 67.1% | 98.5% | 99.8% | 100% | 99.9% | 99.8%
3.31 | 1.802% | 66.5% | 98.4% | 99.8% | 100% | 99.9% | 99.8%
3.32 | 1.702% | 65.9% | 98.3% | 99.8% | 100% | 99.9% | 99.7%
3.33 | 1.702% | 65.3% | 98.2% | 99.8% | 100% | 99.9% | 99.7%
3.34 | 1.602% | 64.69% | 98.2% | 99.8% | 100% | 99.9% | 99.7%
3.35 | 1.602% | 64.09% | 98.1% | 99.7% | 100% | 99.9% | 99.7%
3.36 | 1.502% | 63.49% | 98% | 99.7% | 100% | 99.9% | 99.7%
3.37 | 1.502% | 62.89% | 97.9% | 99.7% | 100% | 99.9% | 99.7%
3.38 | 1.401% | 62.39% | 97.9% | 99.7% | 100% | 99.9% | 99.7%
3.39 | 1.401% | 61.79% | 97.8% | 99.7% | 100% | 99.9% | 99.7%
3.4 | 1.401% | 61.18% | 97.7% | 99.7% | 100% | 99.9% | 99.7%
3.41 | 1.301% | 60.58% | 97.6% | 99.7% | 100% | 99.8% | 99.7%
3.42 | 1.301% | 59.98% | 97.49% | 99.6% | 100% | 99.8% | 99.7%
3.43 | 1.301% | 59.38% | 97.39% | 99.6% | 100% | 99.8% | 99.7%
3.44 | 1.201% | 58.88% | 97.29% | 99.6% | 100% | 99.8% | 99.6%
3.45 | 1.201% | 58.17% | 97.19% | 99.6% | 100% | 99.8% | 99.6%
3.46 | 1.201% | 57.57% | 97.09% | 99.6% | 99.9% | 99.8% | 99.6%
3.47 | 1.201% | 56.97% | 96.99% | 99.6% | 99.9% | 99.8% | 99.6%
3.48 | 1.201% | 56.37% | 96.89% | 99.6% | 99.9% | 99.7% | 99.6%
3.49 | 1.201% | 55.67% | 96.79% | 99.6% | 99.9% | 99.7% | 99.6%
3.5 | 1.201% | 55.07% | 96.69% | 99.5% | 99.9% | 99.7% | 99.6%
3.51 | 1.201% | 54.36% | 96.59% | 99.5% | 99.9% | 99.7% | 99.6%
3.52 | 1.101% | 53.76% | 96.39% | 99.5% | 99.9% | 99.7% | 99.6%
3.53 | 1.101% | 53.06% | 96.29% | 99.5% | 99.8% | 99.7% | 99.6%
3.54 | 1.101% | 52.46% | 96.19% | 99.5% | 99.8% | 99.6% | 99.6%
3.55 | 1.101% | 51.76% | 96.09% | 99.5% | 99.8% | 99.6% | 99.6%
3.56 | 1.101% | 51.15% | 95.99% | 99.4% | 99.8% | 99.6% | 99.6%
3.57 | 1.101% | 50.45% | 95.89% | 99.4% | 99.8% | 99.6% | 99.6%
3.58 | 1.101% | 49.85% | 95.69% | 99.4% | 99.7% | 99.6% | 99.6%
3.59 | 1.101% | 49.15% | 95.59% | 99.4% | 99.7% | 99.5% | 99.6%
3.6 | 1.101% | 48.55% | 95.39% | 99.4% | 99.7% | 99.5% | 99.6%
3.61 | 1.101% | 47.94% | 95.29% | 99.4% | 99.7% | 99.5% | 99.6%
3.62 | 1.101% | 47.34% | 95.09% | 99.3% | 99.7% | 99.5% | 99.6%
3.63 | 1.101% | 46.64% | 94.89% | 99.3% | 99.7% | 99.5% | 99.6%
3.64 | 1.101% | 46.04% | 94.79% | 99.3% | 99.6% | 99.5% | 99.6%
3.65 | 1.101% | 45.44% | 94.59% | 99.3% | 99.6% | 99.4% | 99.5%
3.66 | 1.101% | 44.83% | 94.49% | 99.2% | 99.6% | 99.4% | 99.5%
3.67 | 1.101% | 44.13% | 94.29% | 99.2% | 99.6% | 99.4% | 99.5%
3.68 | 1.101% | 43.53% | 94.09% | 99.2% | 99.6% | 99.4% | 99.5%
3.69 | 1.001% | 42.93% | 93.99% | 99.2% | 99.6% | 99.4% | 99.5%
3.7 | 1.001% | 42.23% | 93.79% | 99.2% | 99.5% | 99.4% | 99.4%
3.71 | 1.001% | 41.62% | 93.69% | 99.1% | 99.5% | 99.3% | 99.4%
3.72 | 1.001% | 40.92% | 93.49% | 99.1% | 99.5% | 99.3% | 99.4%
3.73 | 1.001% | 40.32% | 93.29% | 99.1% | 99.5% | 99.3% | 99.4%
3.74 | 1.001% | 39.72% | 93.19% | 99.1% | 99.5% | 99.3% | 99.3%
3.75 | 1.001% | 39.02% | 92.99% | 99% | 99.4% | 99.3% | 99.3%
3.76 | 1.001% | 38.42% | 92.79% | 99% | 99.4% | 99.2% | 99.3%
3.77 | 1.001% | 37.81% | 92.59% | 99% | 99.4% | 99.2% | 99.3%
3.78 | 1.001% | 37.21% | 92.38% | 99% | 99.4% | 99.2% | 99.3%
3.79 | 1.001% | 36.61% | 92.28% | 98.9% | 99.4% | 99.2% | 99.3%
3.8 | 1.001% | 36.01% | 92.08% | 98.9% | 99.4% | 99.2% | 99.3%
3.81 | 1.001% | 35.41% | 91.88% | 98.9% | 99.4% | 99.2% | 99.3%
3.82 | 1.001% | 34.8% | 91.68% | 98.8% | 99.4% | 99.1% | 99.3%
3.83 | 1.001% | 34.2% | 91.38% | 98.8% | 99.4% | 99.1% | 99.3%
3.84 | 1.001% | 33.6% | 91.18% | 98.8% | 99.4% | 99.1% | 99.3%
3.85 | 1.001% | 33% | 90.98% | 98.7% | 99.3% | 99.1% | 99.3%
3.86 | 1.001% | 32.4% | 90.78% | 98.7% | 99.3% | 99.1% | 99.2%
3.87 | 1.001% | 31.8% | 90.48% | 98.7% | 99.3% | 99.1% | 99.2%
3.88 | 1.001% | 31.19% | 90.28% | 98.6% | 99.3% | 99% | 99.2%
3.89 | 1.001% | 30.59% | 90.08% | 98.6% | 99.3% | 99% | 99.2%
3.9 | 1.001% | 29.99% | 89.88% | 98.6% | 99.3% | 99% | 99.2%
3.91 | 1.001% | 29.39% | 89.58% | 98.5% | 99.2% | 99% | 99.2%
3.92 | 1.001% | 28.89% | 89.38% | 98.5% | 99.2% | 99% | 99.2%
3.93 | 1.001% | 28.28% | 89.18% | 98.4% | 99.2% | 98.9% | 99.2%
3.94 | 1.001% | 27.68% | 88.88% | 98.4% | 99.2% | 98.9% | 99.2%
3.95 | 1.001% | 27.08% | 88.68% | 98.4% | 99.2% | 98.9% | 99.2%
3.96 | 1.001% | 26.58% | 88.48% | 98.3% | 99.1% | 98.9% | 99.2%
3.97 | 1.001% | 25.98% | 88.18% | 98.3% | 99.1% | 98.9% | 99.1%
3.98 | 1.001% | 25.48% | 87.98% | 98.3% | 99.1% | 98.8% | 99.1%
3.99 | 1.001% | 24.87% | 87.68% | 98.2% | 99.1% | 98.8% | 99.1%
4 | 1.001% | 24.37% | 87.47% | 98.2% | 99.1% | 98.8% | 99.1%
4.01 | 1.001% | 23.87% | 87.17% | 98.1% | 99% | 98.8% | 99.1%
4.02 | 1.001% | 23.37% | 86.97% | 98.1% | 99% | 98.7% | 99.1%
4.03 | 1.001% | 22.87% | 86.67% | 98% | 99% | 98.7% | 99.1%
4.04 | 1.001% | 22.27% | 86.47% | 98% | 99% | 98.7% | 99%
4.05 | 1.001% | 21.77% | 86.17% | 97.9% | 98.9% | 98.6% | 99%
4.06 | 1.001% | 21.26% | 85.97% | 97.9% | 98.9% | 98.6% | 99%
4.07 | 1.001% | 20.76% | 85.67% | 97.8% | 98.8% | 98.6% | 99%
4.08 | 1.001% | 20.26% | 85.37% | 97.8% | 98.8% | 98.6% | 99%
4.09 | 1.001% | 19.76% | 85.07% | 97.7% | 98.7% | 98.5% | 98.9%
4.1 | 1.001% | 19.26% | 84.77% | 97.7% | 98.6% | 98.5% | 98.9%
4.11 | 1.001% | 18.86% | 84.47% | 97.6% | 98.6% | 98.5% | 98.9%
4.12 | 1.001% | 18.36% | 84.17% | 97.6% | 98.5% | 98.4% | 98.8%
4.13 | 1.001% | 17.85% | 83.87% | 97.5% | 98.5% | 98.4% | 98.8%
4.14 | 1.001% | 17.45% | 83.67% | 97.5% | 98.4% | 98.4% | 98.8%
4.15 | 1.001% | 16.95% | 83.37% | 97.4% | 98.4% | 98.3% | 98.8%
4.16 | 1.001% | 16.45% | 83.07% | 97.4% | 98.3% | 98.3% | 98.7%
4.17 | 1.001% | 16.05% | 82.77% | 97.3% | 98.3% | 98.3% | 98.7%
4.18 | 1.001% | 15.55% | 82.36% | 97.2% | 98.2% | 98.3% | 98.7%
4.19 | 1.001% | 15.15% | 82.06% | 97.2% | 98.2% | 98.2% | 98.7%
4.2 | 1.001% | 14.74% | 81.76% | 97.1% | 98.1% | 98.2% | 98.7%
4.21 | 1.001% | 14.34% | 81.46% | 97.1% | 98.1% | 98.2% | 98.6%
4.22 | 1.001% | 13.94% | 81.06% | 97% | 98% | 98.1% | 98.6%
4.23 | 1.001% | 13.54% | 80.76% | 96.9% | 98% | 98.1% | 98.6%
4.24 | 1.001% | 13.24% | 80.36% | 96.9% | 97.9% | 98% | 98.6%
4.25 | 1.001% | 12.84% | 80.06% | 96.8% | 97.9% | 98% | 98.5%
4.26 | 1.001% | 12.44% | 79.76% | 96.7% | 97.8% | 98% | 98.5%
4.27 | 1.001% | 12.14% | 79.36% | 96.7% | 97.8% | 97.9% | 98.5%
4.28 | 1.001% | 11.74% | 79.06% | 96.6% | 97.7% | 97.9% | 98.5%
4.29 | 1.001% | 11.33% | 78.66% | 96.5% | 97.7% | 97.9% | 98.4%
4.3 | 1.001% | 11.03% | 78.36% | 96.5% | 97.6% | 97.8% | 98.4%
4.31 | 1.001% | 10.73% | 77.96% | 96.4% | 97.6% | 97.8% | 98.4%
4.32 | 1.001% | 10.33% | 77.66% | 96.3% | 97.5% | 97.8% | 98.3%
4.33 | 1.001% | 10.03% | 77.25% | 96.3% | 97.4% | 97.7% | 98.3%
4.34 | 1.001% | 9.729% | 76.95% | 96.2% | 97.4% | 97.7% | 98.3%
4.35 | 1.001% | 9.428% | 76.55% | 96.1% | 97.3% | 97.6% | 98.3%
4.36 | 1.001% | 9.127% | 76.25% | 96.1% | 97.2% | 97.6% | 98.2%
4.37 | 1.001% | 8.927% | 75.85% | 96% | 97.2% | 97.6% | 98.2%
4.38 | 1.001% | 8.626% | 75.55% | 95.9% | 97.1% | 97.5% | 98.2%
4.39 | 1.001% | 8.325% | 75.15% | 95.8% | 97% | 97.5% | 98.2%
4.4 | 1.001% | 8.024% | 74.85% | 95.7% | 97% | 97.4% | 98.1%
4.41 | 1.001% | 7.823% | 74.45% | 95.7% | 96.9% | 97.4% | 98.1%
4.42 | 1.001% | 7.523% | 74.15% | 95.6% | 96.9% | 97.4% | 98.1%
4.43 | 1.001% | 7.322% | 73.75% | 95.5% | 96.8% | 97.3% | 98.1%
4.44 | 1.001% | 7.121% | 73.45% | 95.4% | 96.8% | 97.3% | 98.1%
4.45 | 1.001% | 6.82% | 73.05% | 95.3% | 96.7% | 97.3% | 98%
4.46 | 1.001% | 6.62% | 72.65% | 95.2% | 96.6% | 97.2% | 98%
4.47 | 1.001% | 6.419% | 72.24% | 95.2% | 96.6% | 97.2% | 98%
4.48 | 1.001% | 6.219% | 71.94% | 95.1% | 96.5% | 97.1% | 98%
4.49 | 1.001% | 6.018% | 71.54% | 94.99% | 96.5% | 97.1% | 97.9%
4.5 | 1.001% | 5.817% | 71.14% | 94.89% | 96.4% | 97.1% | 97.9%
4.51 | 1.001% | 5.617% | 70.74% | 94.79% | 96.3% | 97% | 97.9%
4.52 | 1.001% | 5.416% | 70.34% | 94.69% | 96.3% | 97% | 97.9%
4.53 | 1.001% | 5.216% | 69.94% | 94.59% | 96.2% | 97% | 97.8%
4.54 | 1.001% | 5.015% | 69.54% | 94.49% | 96.1% | 96.9% | 97.8%
4.55 | 1.001% | 4.814% | 69.14% | 94.39% | 96% | 96.9% | 97.8%
4.56 | 1.001% | 4.714% | 68.64% | 94.39% | 95.9% | 96.8% | 97.8%
4.57 | 1.001% | 4.514% | 68.24% | 94.29% | 95.9% | 96.8% | 97.7%
4.58 | 1.001% | 4.413% | 67.84% | 94.19% | 95.8% | 96.8% | 97.7%
4.59 | 1.001% | 4.313% | 67.33% | 94.09% | 95.7% | 96.7% | 97.7%
4.6 | 1.001% | 4.112% | 66.93% | 93.99% | 95.6% | 96.7% | 97.6%
4.61 | 1.001% | 4.012% | 66.53% | 93.89% | 95.5% | 96.6% | 97.6%
4.62 | 1.001% | 3.912% | 66.03% | 93.79% | 95.4% | 96.6% | 97.6%
4.63 | 1.001% | 3.811% | 65.63% | 93.69% | 95.3% | 96.6% | 97.5%
4.64 | 1.001% | 3.611% | 65.23% | 93.59% | 95.2% | 96.5% | 97.5%
4.65 | 1.001% | 3.511% | 64.73% | 93.49% | 95.1% | 96.5% | 97.5%
4.66 | 0.9009% | 3.41% | 64.33% | 93.29% | 94.99% | 96.4% | 97.41%
4.67 | 0.9009% | 3.31% | 63.93% | 93.19% | 94.89% | 96.4% | 97.41%
4.68 | 0.9009% | 3.21% | 63.43% | 93.09% | 94.79% | 96.4% | 97.31%
4.69 | 0.9009% | 3.109% | 63.03% | 92.99% | 94.69% | 96.3% | 97.31%
4.7 | 0.9009% | 3.009% | 62.53% | 92.89% | 94.59% | 96.3% | 97.31%
4.71 | 0.9009% | 3.009% | 62.12% | 92.79% | 94.49% | 96.2% | 97.21%
4.72 | 0.9009% | 2.909% | 61.62% | 92.69% | 94.39% | 96.2% | 97.21%
4.73 | 0.9009% | 2.808% | 61.22% | 92.59% | 94.29% | 96.1% | 97.21%
4.74 | 0.9009% | 2.708% | 60.72% | 92.49% | 94.19% | 96.1% | 97.21%
4.75 | 0.9009% | 2.708% | 60.32% | 92.29% | 94.09% | 96% | 97.11%
4.76 | 0.9009% | 2.608% | 59.82% | 92.19% | 93.99% | 96% | 97.11%
4.77 | 0.9009% | 2.508% | 59.42% | 92.09% | 93.89% | 96% | 97.11%
4.78 | 0.9009% | 2.508% | 58.92% | 91.99% | 93.79% | 95.9% | 97.11%
4.79 | 0.9009% | 2.407% | 58.52% | 91.89% | 93.69% | 95.9% | 97.11%
4.8 | 0.9009% | 2.307% | 58.02% | 91.69% | 93.49% | 95.8% | 97.01%
4.81 | 0.9009% | 2.307% | 57.52% | 91.59% | 93.39% | 95.8% | 97.01%
4.82 | 0.9009% | 2.307% | 57.11% | 91.49% | 93.29% | 95.7% | 97.01%
4.83 | 0.9009% | 2.207% | 56.61% | 91.39% | 93.19% | 95.7% | 96.91%
4.84 | 0.9009% | 2.207% | 56.21% | 91.29% | 93.09% | 95.6% | 96.91%
4.85 | 0.9009% | 2.106% | 55.71% | 91.09% | 92.99% | 95.6% | 96.91%
4.86 | 0.9009% | 2.106% | 55.21% | 90.99% | 92.89% | 95.5% | 96.81%
4.87 | 0.9009% | 2.106% | 54.81% | 90.89% | 92.69% | 95.5% | 96.81%
4.88 | 0.9009% | 2.006% | 54.31% | 90.69% | 92.59% | 95.5% | 96.81%
4.89 | 0.9009% | 2.006% | 53.91% | 90.59% | 92.49% | 95.4% | 96.71%
4.9 | 0.9009% | 2.006% | 53.41% | 90.49% | 92.29% | 95.4% | 96.71%
4.91 | 0.9009% | 1.906% | 52.91% | 90.29% | 92.19% | 95.3% | 96.71%
4.92 | 0.9009% | 1.906% | 52.4% | 90.19% | 92.09% | 95.3% | 96.61%
4.93 | 0.9009% | 1.906% | 52% | 90.09% | 91.99% | 95.2% | 96.61%
4.94 | 0.9009% | 1.906% | 51.5% | 89.89% | 91.79% | 95.1% | 96.61%
4.95 | 0.9009% | 1.805% | 51% | 89.79% | 91.69% | 95.1% | 96.51%
4.96 | 0.9009% | 1.805% | 50.5% | 89.59% | 91.59% | 94.99% | 96.51%
4.97 | 0.9009% | 1.805% | 50.1% | 89.49% | 91.39% | 94.99% | 96.41%
4.98 | 0.9009% | 1.805% | 49.6% | 89.29% | 91.29% | 94.89% | 96.41%
4.99 | 0.9009% | 1.805% | 49.1% | 89.19% | 91.19% | 94.89% | 96.31%
5 | 0.9009% | 1.705% | 48.6% | 88.99% | 90.99% | 94.79% | 96.31%
5.01 | 0.9009% | 1.705% | 48.1% | 88.89% | 90.89% | 94.79% | 96.21%
5.02 | 0.9009% | 1.705% | 47.7% | 88.69% | 90.79% | 94.69% | 96.11%
5.03 | 0.9009% | 1.705% | 47.19% | 88.59% | 90.59% | 94.59% | 96.11%
5.04 | 0.9009% | 1.705% | 46.69% | 88.39% | 90.49% | 94.59% | 96.01%
5.05 | 0.9009% | 1.705% | 46.29% | 88.29% | 90.39% | 94.49% | 96.01%
5.06 | 0.9009% | 1.705% | 45.79% | 88.09% | 90.19% | 94.49% | 95.91%
5.07 | 0.9009% | 1.705% | 45.29% | 87.89% | 90.09% | 94.39% | 95.91%
5.08 | 0.9009% | 1.605% | 44.89% | 87.79% | 89.89% | 94.39% | 95.81%
5.09 | 0.9009% | 1.605% | 44.39% | 87.59% | 89.79% | 94.29% | 95.81%
5.1 | 0.9009% | 1.605% | 43.89% | 87.49% | 89.69% | 94.19% | 95.71%
5.11 | 0.9009% | 1.605% | 43.39% | 87.29% | 89.49% | 94.19% | 95.71%
5.12 | 0.9009% | 1.605% | 42.99% | 87.19% | 89.39% | 94.09% | 95.71%
5.13 | 0.9009% | 1.605% | 42.48% | 86.99% | 89.19% | 93.99% | 95.61%
5.14 | 0.9009% | 1.605% | 41.98% | 86.79% | 88.99% | 93.99% | 95.61%
5.15 | 0.9009% | 1.605% | 41.48% | 86.69% | 88.89% | 93.89% | 95.61%
5.16 | 0.9009% | 1.605% | 41.08% | 86.49% | 88.69% | 93.89% | 95.51%
5.17 | 0.9009% | 1.605% | 40.58% | 86.39% | 88.59% | 93.79% | 95.51%
5.18 | 0.9009% | 1.605% | 40.18% | 86.19% | 88.39% | 93.69% | 95.41%
5.19 | 0.9009% | 1.605% | 39.68% | 85.99% | 88.19% | 93.69% | 95.41%
5.2 | 0.9009% | 1.605% | 39.28% | 85.79% | 88.09% | 93.59% | 95.31%
5.21 | 0.9009% | 1.605% | 38.78% | 85.69% | 87.89% | 93.59% | 95.31%
5.22 | 0.9009% | 1.605% | 38.38% | 85.49% | 87.69% | 93.49% | 95.21%
5.23 | 0.9009% | 1.605% | 37.88% | 85.29% | 87.59% | 93.39% | 95.21%
5.24 | 0.9009% | 1.605% | 37.47% | 85.09% | 87.39% | 93.39% | 95.11%
5.25 | 0.9009% | 1.605% | 36.97% | 84.98% | 87.19% | 93.29% | 95.11%
5.26 | 0.9009% | 1.605% | 36.57% | 84.78% | 87.09% | 93.29% | 95.01%
5.27 | 0.9009% | 1.605% | 36.07% | 84.58% | 86.89% | 93.19% | 94.91%
5.28 | 0.9009% | 1.605% | 35.67% | 84.38% | 86.69% | 93.09% | 94.91%
5.29 | 0.9009% | 1.505% | 35.27% | 84.18% | 86.49% | 93.09% | 94.81%
5.3 | 0.9009% | 1.505% | 34.77% | 84.08% | 86.29% | 92.99% | 94.71%
5.31 | 0.9009% | 1.505% | 34.37% | 83.88% | 86.19% | 92.89% | 94.71%
5.32 | 0.9009% | 1.505% | 33.87% | 83.68% | 85.99% | 92.89% | 94.61%
5.33 | 0.9009% | 1.505% | 33.47% | 83.48% | 85.79% | 92.79% | 94.61%
5.34 | 0.9009% | 1.505% | 33.07% | 83.28% | 85.59% | 92.69% | 94.51%
5.35 | 0.9009% | 1.505% | 32.67% | 83.08% | 85.39% | 92.69% | 94.41%
5.36 | 0.9009% | 1.505% | 32.16% | 82.88% | 85.29% | 92.59% | 94.41%
5.37 | 0.9009% | 1.505% | 31.76% | 82.68% | 85.09% | 92.49% | 94.31%
5.38 |  | 1.505% | 31.26% | 82.48% | 84.88% | 92.49% | 94.31%
5.39 |  | 1.505% | 30.86% | 82.28% | 84.68% | 92.39% | 94.21%
5.4 |  | 1.505% | 30.46% | 82.08% | 84.48% | 92.29% | 94.11%
5.41 |  | 1.505% | 30.06% | 81.98% | 84.28% | 92.19% | 94.11%
5.42 |  | 1.505% | 29.56% | 81.78% | 84.08% | 92.19% | 94.01%
5.43 |  | 1.505% | 29.16% | 81.58% | 83.88% | 92.09% | 93.91%
5.44 |  | 1.505% | 28.76% | 81.38% | 83.68% | 91.99% | 93.91%
5.45 |  | 1.505% | 28.26% | 81.18% | 83.38% | 91.89% | 93.81%
5.46 |  | 1.505% | 27.86% | 80.98% | 83.18% | 91.89% | 93.71%
5.47 |  | 1.505% | 27.45% | 80.68% | 82.98% | 91.79% | 93.71%
5.48 |  | 1.505% | 26.95% | 80.48% | 82.78% | 91.69% | 93.61%
5.49 |  | 1.505% | 26.55% | 80.28% | 82.58% | 91.59% | 93.51%
5.5 |  | 1.505% | 26.15% | 80.08% | 82.38% | 91.59% | 93.41%
5.51 |  | 1.505% | 25.75% | 79.88% | 82.18% | 91.49% | 93.41%
5.52 |  | 1.505% | 25.35% | 79.68% | 81.88% | 91.39% | 93.31%
5.53 |  | 1.505% | 24.95% | 79.48% | 81.68% | 91.29% | 93.31%
5.54 |  | 1.505% | 24.55% | 79.28% | 81.48% | 91.19% | 93.21%
5.55 |  | 1.505% | 24.15% | 78.98% | 81.28% | 91.19% | 93.21%
5.56 |  | 1.505% | 23.75% | 78.78% | 81.08% | 91.09% | 93.11%
5.57 |  | 1.505% | 23.35% | 78.58% | 80.78% | 90.99% | 93.01%
5.58 |  | 1.505% | 23.05% | 78.38% | 80.58% | 90.89% | 93.01%
5.59 |  | 1.505% | 22.65% | 78.08% | 80.38% | 90.79% | 92.91%
5.6 |  | 1.505% | 22.24% | 77.88% | 80.18% | 90.69% | 92.91%
5.61 |  | 1.505% | 21.84% | 77.68% | 79.98% | 90.69% | 92.81%
5.62 |  | 1.505% | 21.54% | 77.48% | 79.68% | 90.59% | 92.71%
5.63 |  | 1.505% | 21.14% | 77.18% | 79.48% | 90.49% | 92.71%
5.64 |  | 1.505% | 20.74% | 76.98% | 79.28% | 90.39% | 92.61%
5.65 |  | 1.505% | 20.44% | 76.68% | 79.08% | 90.29% | 92.51%
5.66 |  | 1.505% | 20.04% | 76.48% | 78.88% | 90.19% | 92.42%
5.67 |  | 1.505% | 19.74% | 76.28% | 78.58% | 90.09% | 92.42%
5.68 |  | 1.505% | 19.34% | 75.98% | 78.38% | 89.99% | 92.32%
5.69 |  | 1.505% | 19.04% | 75.78% | 78.18% | 89.89% | 92.22%
5.7 |  | 1.505% | 18.64% | 75.48% | 77.98% | 89.79% | 92.12%
5.71 |  | 1.505% | 18.34% | 75.28% | 77.78% | 89.69% | 92.12%
5.72 |  | 1.505% | 18.04% | 75.08% | 77.48% | 89.59% | 92.02%
5.73 |  | 1.505% | 17.64% | 74.77% | 77.28% | 89.49% | 91.92%
5.74 |  | 1.505% | 17.33% | 74.47% | 76.98% | 89.39% | 91.82%
5.75 |  | 1.505% | 17.03% | 74.27% | 76.78% | 89.29% | 91.82%
5.76 |  | 1.505% | 16.63% | 73.97% | 76.48% | 89.19% | 91.72%
5.77 |  | 1.505% | 16.33% | 73.77% | 76.28% | 89.09% | 91.62%
5.78 |  | 1.505% | 16.03% | 73.47% | 75.98% | 88.99% | 91.52%
5.79 |  | 1.505% | 15.73% | 73.27% | 75.78% | 88.89% | 91.42%
5.8 |  | 1.505% | 15.43% | 72.97% | 75.48% | 88.79% | 91.42%
5.81 |  | 1.505% | 15.13% | 72.77% | 75.28% | 88.69% | 91.32%
5.82 |  | 1.505% | 14.83% | 72.47% | 74.97% | 88.59% | 91.22%
5.83 |  | 1.505% | 14.53% | 72.27% | 74.77% | 88.49% | 91.22%
5.84 |  | 1.505% | 14.23% | 71.97% | 74.47% | 88.39% | 91.12%
5.85 |  | 1.505% | 13.93% | 71.67% | 74.17% | 88.19% | 91.02%
5.86 |  | 1.505% | 13.63% | 71.47% | 73.97% | 88.09% | 91.02%
5.87 |  | 1.505% | 13.33% | 71.17% | 73.67% | 87.99% | 90.92%
5.88 |  | 1.505% | 13.03% | 70.97% | 73.37% | 87.89% | 90.82%
5.89 |  | 1.505% | 12.83% | 70.67% | 73.07% | 87.79% | 90.82%
5.9 |  | 1.404% | 12.53% | 70.37% | 72.77% | 87.69% | 90.72%
5.91 |  | 1.404% | 12.32% | 70.17% | 72.57% | 87.59% | 90.62%
5.92 |  | 1.404% | 12.02% | 69.87% | 72.27% | 87.49% | 90.52%
5.93 |  | 1.404% | 11.82% | 69.57% | 71.97% | 87.39% | 90.42%
5.94 |  | 1.404% | 11.52% | 69.27% | 71.67% | 87.29% | 90.32%
5.95 |  | 1.404% | 11.32% | 68.97% | 71.37% | 87.19% | 90.22%
5.96 |  | 1.404% | 11.12% | 68.77% | 71.07% | 87.09% | 90.12%
5.97 |  | 1.404% | 10.82% | 68.47% | 70.77% | 86.89% | 90.02%
5.98 |  | 1.404% | 10.62% | 68.17% | 70.47% | 86.79% | 90.02%
5.99 |  | 1.404% | 10.42% | 67.87% | 70.27% | 86.69% | 89.82%
6 |  | 1.505% | 10.22% | 67.57% | 69.97% | 86.59% | 89.72%
6.01 |  | 1.505% | 10.02% | 67.27% | 69.67% | 86.49% | 89.62%
6.02 |  | 1.505% | 9.719% | 67.07% | 69.37% | 86.39% | 89.52%
6.03 |  | 1.505% | 9.519% | 66.77% | 69.07% | 86.19% | 89.42%
6.04 |  | 1.505% | 9.419% | 66.47% | 68.77% | 86.09% | 89.32%
6.05 |  | 1.505% | 9.218% | 66.17% | 68.47% | 85.99% | 89.22%
6.06 |  | 1.505% | 9.018% | 65.87% | 68.17% | 85.89% | 89.12%
6.07 |  | 1.404% | 8.818% | 65.57% | 67.97% | 85.79% | 89.02%
6.08 |  |  | 8.617% | 65.27% | 67.67% | 85.69% | 88.92%
6.09 |  |  | 8.417% | 64.96% | 67.37% | 85.49% | 88.82%
6.1 |  |  | 8.216% | 64.66% | 67.07% | 85.39% | 88.72%
6.11 |  |  | 8.016% | 64.36% | 66.77% | 85.29% | 88.62%
6.12 |  |  | 7.816% | 64.16% | 66.47% | 85.19% | 88.52%
6.13 |  |  | 7.615% | 63.86% | 66.17% | 85.09% | 88.42%
6.14 |  |  | 7.515% | 63.56% | 65.87% | 84.98% | 88.32%
6.15 |  |  | 7.315% | 63.26% | 65.57% | 84.78% | 88.22%
6.16 |  |  | 7.114% | 62.96% | 65.27% | 84.68% | 88.12%
6.17 |  |  | 7.014% | 62.66% | 64.96% | 84.58% | 88.02%
6.18 |  |  | 6.814% | 62.36% | 64.66% | 84.48% | 87.92%
6.19 |  |  | 6.613% | 62.06% | 64.36% | 84.38% | 87.82%
6.2 |  |  | 6.513% | 61.76% | 64.06% | 84.28% | 87.62%
6.21 |  |  | 6.313% | 61.46% | 63.76% | 84.08% | 87.52%
6.22 |  |  | 6.212% | 61.16% | 63.46% | 83.98% | 87.43%
6.23 |  |  | 6.112% | 60.86% | 63.06% | 83.88% | 87.33%
6.24 |  |  | 5.912% | 60.56% | 62.76% | 83.78% | 87.23%
6.25 |  |  | 5.812% | 60.26% | 62.46% | 83.68% | 87.13%
6.26 |  |  | 5.711% | 59.96% | 62.16% | 83.48% | 87.03%
6.27 |  |  | 5.611% | 59.66% | 61.86% | 83.38% | 87.03%
6.28 |  |  | 5.511% | 59.26% | 61.46% | 83.28% | 86.93%
6.29 |  |  | 5.411% | 58.96% | 61.16% | 83.18% | 86.83%
6.3 |  |  | 5.311% | 58.66% | 60.86% | 83.08% | 86.73%
6.31 |  |  | 5.11% | 58.36% | 60.56% | 82.88% | 86.63%
6.32 |  |  | 5.01% | 58.06% | 60.26% | 82.78% | 86.43%
6.33 |  |  | 4.91% | 57.76% | 59.86% | 82.68% | 86.33%
6.34 |  |  | 4.81% | 57.46% | 59.56% | 82.58% | 86.23%
6.35 |  |  | 4.709% | 57.16% | 59.16% | 82.38% | 86.13%
6.36 |  |  | 4.709% | 56.86% | 58.86% | 82.28% | 86.03%
6.37 |  |  | 4.609% | 56.46% | 58.56% | 82.18% | 85.93%
6.38 |  |  | 4.509% | 56.16% | 58.16% | 82.08% | 85.73%
6.39 |  |  | 4.409% | 55.86% | 57.86% | 81.88% | 85.63%
6.4 |  |  | 4.309% | 55.56% | 57.46% | 81.78% | 85.53%
6.41 |  |  | 4.208% | 55.26% | 57.16% | 81.68% | 85.33%
6.42 |  |  | 4.208% | 54.85% | 56.76% | 81.48% | 85.23%
6.43 |  |  | 4.108% | 54.55% | 56.46% | 81.38% | 85.13%
6.44 |  |  | 4.008% | 54.25% | 56.16% | 81.18% | 85.03%
6.45 |  |  | 3.908% | 53.95% | 55.76% | 81.08% | 84.93%
6.46 |  |  | 3.908% | 53.55% | 55.46% | 80.98% | 84.83%
6.47 |  |  | 3.808% | 53.25% | 55.16% | 80.78% | 84.63%
6.48 |  |  | 3.707% | 52.95% | 54.85% | 80.68% | 84.53%
6.49 |  |  | 3.707% | 52.55% | 54.55% | 80.48% | 84.43%
6.5 |  |  | 3.607% | 52.25% | 54.25% | 80.38% | 84.33%
6.51 |  |  | 3.607% | 51.95% | 53.95% | 80.18% | 84.13%
6.52 |  |  | 3.607% | 51.65% | 53.55% | 80.08% | 84.03%
6.53 |  |  | 3.507% | 51.25% | 53.25% | 79.88% | 83.93%
6.54 |  |  | 3.507% | 50.95% | 52.95% | 79.78% | 83.73%
6.55 |  |  | 3.407% | 50.65% | 52.65% | 79.58% | 83.63%
6.56 |  |  | 3.407% | 50.25% | 52.25% | 79.48% | 83.43%
6.57 |  |  | 3.307% | 49.95% | 51.95% | 79.28% | 83.33%
6.58 |  |  | 3.307% | 49.65% | 51.55% | 79.18% | 83.13%
6.59 |  |  | 3.307% | 49.35% | 51.15% | 78.98% | 83.03%
6.6 |  |  | 3.206% | 48.95% | 50.85% | 78.88% | 82.83%
6.61 |  |  | 3.206% | 48.65% | 50.45% | 78.68% | 82.73%
6.62 |  |  | 3.106% | 48.35% | 50.15% | 78.48% | 82.53%
6.63 |  |  | 3.106% | 47.95% | 49.75% | 78.38% | 82.44%
6.64 |  |  | 3.106% | 47.65% | 49.35% | 78.18% | 82.24%
6.65 |  |  | 3.106% | 47.35% | 49.05% | 77.98% | 82.14%
6.66 |  |  | 3.006% | 46.95% | 48.75% | 77.88% | 81.94%
6.67 |  |  | 3.006% | 46.65% | 48.35% | 77.68% | 81.84%
6.68 |  |  | 3.006% | 46.35% | 48.05% | 77.48% | 81.64%
6.69 |  |  | 2.906% | 45.95% | 47.65% | 77.38% | 81.54%
6.7 |  |  | 2.906% | 45.65% | 47.35% | 77.18% | 81.34%
6.71 |  |  | 2.906% | 45.35% | 46.95% | 76.98% | 81.24%
6.72 |  |  | 2.906% | 44.94% | 46.55% | 76.88% | 81.04%
6.73 |  |  | 2.906% | 44.64% | 46.25% | 76.68% | 80.94%
6.74 |  |  | 2.806% | 44.34% | 45.85% | 76.48% | 80.84%
6.75 |  |  | 2.806% | 43.94% | 45.55% | 76.38% | 80.64%
6.76 |  |  | 2.806% | 43.64% | 45.15% | 76.18% | 80.54%
6.77 |  |  | 2.806% | 43.34% | 44.84% | 75.98% | 80.34%
6.78 |  |  | 2.806% | 42.94% | 44.44% | 75.78% | 80.24%
6.79 |  |  | 2.705% | 42.64% | 44.04% | 75.68% | 80.04%
6.8 |  |  | 2.705% | 42.24% | 43.74% | 75.48% | 79.94%
6.81 |  |  | 2.705% | 41.94% | 43.34% | 75.28% | 79.74%
6.82 |  |  | 2.705% | 41.64% | 43.04% | 75.08% | 79.64%
6.83 |  |  | 2.705% | 41.24% | 42.64% | 74.97% | 79.44%
6.84 |  |  | 2.605% | 40.94% | 42.34% | 74.77% | 79.24%
6.85 |  |  | 2.605% | 40.64% | 42.04% | 74.57% | 79.14%
6.86 |  |  | 2.605% | 40.24% | 41.64% | 74.37% | 78.94%
6.87 |  |  | 2.605% | 39.94% | 41.34% | 74.27% | 78.74%
6.88 |  |  | 2.605% | 39.54% | 40.94% | 74.07% | 78.64%
6.89 |  |  | 2.605% | 39.24% | 40.64% | 73.87% | 78.44%
6.9 |  |  | 2.605% | 38.94% | 40.24% | 73.67% | 78.34%
6.91 |  |  | 2.605% | 38.54% | 39.94% | 73.57% | 78.14%
6.92 |  |  | 2.605% | 38.24% | 39.54% | 73.37% | 77.94%
6.93 |  |  | 2.605% | 37.84% | 39.24% | 73.17% | 77.74%
6.94 |  |  | 2.605% | 37.54% | 38.84% | 72.97% | 77.64%
6.95 |  |  | 2.505% | 37.24% | 38.54% | 72.87% | 77.45%
6.96 |  |  | 2.505% | 36.84% | 38.14% | 72.67% | 77.25%
6.97 |  |  | 2.505% | 36.54% | 37.84% | 72.47% | 77.05%
6.98 |  |  | 2.505% | 36.14% | 37.44% | 72.27% | 76.95%
6.99 |  |  | 2.505% | 35.84% | 37.14% | 72.07% | 76.75%
7 |  |  | 2.505% | 35.54% | 36.74% | 71.97% | 76.55%
7.01 |  |  | 2.505% | 35.14% | 36.44% | 71.77% | 76.35%
7.02 |  |  | 2.505% | 34.83% | 36.04% | 71.57% | 76.15%
7.03 |  |  | 2.505% | 34.53% | 35.74% | 71.37% | 75.95%
7.04 |  |  | 2.505% | 34.13% | 35.34% | 71.17% | 75.85%
7.05 |  |  | 2.505% | 33.83% | 35.04% | 70.97% | 75.65%
7.06 |  |  | 2.505% | 33.53% | 34.63% | 70.87% | 75.45%
7.07 |  |  | 2.505% | 33.13% | 34.33% | 70.67% | 75.25%
7.08 |  |  | 2.505% | 32.83% | 34.03% | 70.47% | 75.05%
7.09 |  |  | 2.505% | 32.53% | 33.63% | 70.27% | 74.85%
7.1 |  |  | 2.505% | 32.13% | 33.33% | 70.07% | 74.75%
7.11 |  |  | 2.505% | 31.83% | 32.93% | 69.87% | 74.55%
7.12 |  |  | 2.505% | 31.53% | 32.63% | 69.77% | 74.35%
7.13 |  |  | 2.505% | 31.23% | 32.23% | 69.57% | 74.15%
7.14 |  |  | 2.505% | 30.83% | 31.93% | 69.37% | 73.95%
7.15 |  |  | 2.405% | 30.53% | 31.53% | 69.17% | 73.75%
7.16 |  |  | 2.405% | 30.23% | 31.23% | 68.97% | 73.55%
7.17 |  |  | 2.405% | 29.83% | 30.93% | 68.77% | 73.35%
7.18 |  |  | 2.405% | 29.53% | 30.63% | 68.57% | 73.25%
7.19 |  |  | 2.405% | 29.23% | 30.23% | 68.47% | 73.05%
7.2 |  |  | 2.405% | 28.93% | 29.93% | 68.27% | 72.85%
7.21 |  |  | 2.405% | 28.53% | 29.63% | 68.07% | 72.65%
7.22 |  |  | 2.405% | 28.23% | 29.33% | 67.87% | 72.46%
7.23 |  |  | 2.405% | 27.93% | 29.03% | 67.67% | 72.26%
7.24 |  |  | 2.405% | 27.63% | 28.63% | 67.47% | 71.96%
7.25 |  |  | 2.405% | 27.33% | 28.33% | 67.27% | 71.76%
7.26 |  |  | 2.405% | 26.93% | 28.03% | 67.07% | 71.56%
7.27 |  |  | 2.405% | 26.63% | 27.73% | 66.87% | 71.36%
7.28 |  |  | 2.405% | 26.33% | 27.43% | 66.67% | 71.16%
7.29 |  |  | 2.405% | 26.03% | 27.13% | 66.47% | 70.86%
7.3 |  |  | 2.405% | 25.73% | 26.73% | 66.27% | 70.66%
7.31 |  |  | 2.405% | 25.43% | 26.43% | 66.07% | 70.46%
7.32 |  |  | 2.405% | 25.13% | 26.13% | 65.87% | 70.26%
7.33 |  |  | 2.405% | 24.82% | 25.83% | 65.57% | 70.06%
7.34 |  |  | 2.405% | 24.52% | 25.53% | 65.37% | 69.76%
7.35 |  |  | 2.405% | 24.22% | 25.23% | 65.17% | 69.56%
7.36 |  |  | 2.405% | 23.92% | 24.92% | 64.96% | 69.36%
7.37 |  |  | 2.405% | 23.62% | 24.62% | 64.76% | 69.16%
7.38 |  |  | 2.405% | 23.32% | 24.32% | 64.56% | 68.96%
7.39 |  |  | 2.405% | 23.02% | 24.02% | 64.36% | 68.76%
7.4 |  |  | 2.405% | 22.72% | 23.72% | 64.16% | 68.56%
7.41 |  |  | 2.405% | 22.52% | 23.42% | 63.96% | 68.36%
7.42 |  |  | 2.405% | 22.22% | 23.12% | 63.76% | 68.16%
7.43 |  |  | 2.405% | 21.92% | 22.82% | 63.56% | 67.96%
7.44 |  |  | 2.405% | 21.62% | 22.52% | 63.36% | 67.76%
7.45 |  |  | 2.405% | 21.42% | 22.22% | 63.06% | 67.47%
7.46 |  |  | 2.405% | 21.12% | 21.92% | 62.86% | 67.27%
7.47 |  |  | 2.405% | 20.82% | 21.62% | 62.66% | 67.07%
7.48 |  |  | 2.405% | 20.52% | 21.32% | 62.46% | 66.87%
7.49 |  |  | 2.405% | 20.32% | 21.02% | 62.26% | 66.57%
7.5 |  |  | 2.405% | 20.02% | 20.72% | 62.06% | 66.37%
7.51 |  |  | 2.405% | 19.72% | 20.52% | 61.76% | 66.07%
7.52 |  |  | 2.405% | 19.52% | 20.22% | 61.56% | 65.87%
7.53 |  |  | 2.405% | 19.22% | 19.92% | 61.36% | 65.67%
7.54 |  |  | 2.405% | 19.02% | 19.72% | 61.16% | 65.37%
7.55 |  |  | 2.405% | 18.72% | 19.42% | 60.96% | 65.17%
7.56 |  |  | 2.405% | 18.52% | 19.12% | 60.76% | 64.97%
7.57 |  |  | 2.405% | 18.22% | 18.92% | 60.46% | 64.77%
7.58 |  |  | 2.405% | 17.92% | 18.62% | 60.26% | 64.47%
7.59 |  |  | 2.405% | 17.72% | 18.42% | 60.06% | 64.27%
7.6 |  |  | 2.405% | 17.52% | 18.12% | 59.86% | 64.07%
7.61 |  |  | 2.405% | 17.22% | 17.92% | 59.56% | 63.87%
7.62 |  |  | 2.405% | 17.02% | 17.62% | 59.36% | 63.67%
7.63 |  |  | 2.405% | 16.72% | 17.42% | 59.16% | 63.37%
7.64 |  |  | 2.405% | 16.52% | 17.12% | 58.96% | 63.17%
7.65 |  |  | 2.405% | 16.32% | 16.92% | 58.66% | 62.97%
7.66 |  |  | 2.305% | 16.02% | 16.72% | 58.46% | 62.77%
7.67 |  |  | 2.305% | 15.82% | 16.42% | 58.26% | 62.48%
7.68 |  |  | 2.305% | 15.52% | 16.22% | 58.06% | 62.28%
7.69 |  |  | 2.305% | 15.32% | 16.02% | 57.76% | 62.08%
7.7 |  |  | 2.305% | 15.12% | 15.72% | 57.56% | 61.78%
7.71 |  |  | 2.305% | 14.91% | 15.52% | 57.36% | 61.48%
7.72 |  |  | 2.305% | 14.71% | 15.32% | 57.16% | 61.28%
7.73 |  |  | 2.305% | 14.41% | 15.12% | 56.86% | 60.98%
7.74 |  |  | 2.305% | 14.21% | 14.81% | 56.66% | 60.78%
7.75 |  |  | 2.305% | 14.01% | 14.61% | 56.46% | 60.48%
7.76 |  |  | 2.305% | 13.81% | 14.41% | 56.16% | 60.28%
7.77 |  |  | 2.305% | 13.61% | 14.21% | 55.96% | 59.98%
7.78 |  |  | 2.305% | 13.41% | 14.01% | 55.76% | 59.78%
7.79 |  |  | 2.305% | 13.21% | 13.81% | 55.46% | 59.48%
7.8 |  |  | 2.305% | 13.01% | 13.61% | 55.26% | 59.28%
7.81 |  |  | 2.305% | 12.81% | 13.41% | 54.95% | 59.08%
7.82 |  |  | 2.305% | 12.61% | 13.21% | 54.75% | 58.78%
7.83 |  |  | 2.305% | 12.41% | 13.01% | 54.55% | 58.58%
7.84 |  |  | 2.305% | 12.21% | 12.81% | 54.25% | 58.38%
7.85 |  |  | 2.305% | 12.01% | 12.61% | 54.05% | 58.08%
7.86 |  |  | 2.305% | 11.81% | 12.51% | 53.75% | 57.88%
7.87 |  |  | 2.305% | 11.61% | 12.31% | 53.55% | 57.58%
7.88 |  |  | 2.305% | 11.41% | 12.11% | 53.35% | 57.39%
7.89 |  |  | 2.305% | 11.31% | 11.91% | 53.05% | 57.09%
7.9 |  |  | 2.305% | 11.11% | 11.71% | 52.85% | 56.89%
7.91 |  |  | 2.305% | 10.91% | 11.51% | 52.55% | 56.59%
7.92 |  |  | 2.305% | 10.81% | 11.31% | 52.35% | 56.29%
7.93 |  |  | 2.305% | 10.61% | 11.21% | 52.15% | 56.09%
7.94 |  |  | 2.305% | 10.41% | 11.01% | 51.85% | 55.79%
7.95 |  |  | 2.305% | 10.31% | 10.81% | 51.65% | 55.49%
7.96 |  |  | 2.305% | 10.11% | 10.61% | 51.35% | 55.19%
7.97 |  |  | 2.305% | 9.91% | 10.51% | 51.15% | 54.99%
7.98 |  |  | 2.305% | 9.81% | 10.31% | 50.95% | 54.69%
7.99 |  |  | 2.305% | 9.61% | 10.21% | 50.65% | 54.39%
8 |  |  | 2.305% | 9.51% | 10.01% | 50.45% | 54.09%
8.01 |  |  | 2.305% | 9.309% | 9.91% | 50.15% | 53.89%
8.02 |  |  | 2.305% | 9.209% | 9.71% | 49.95% | 53.59%
8.03 |  |  | 2.305% | 9.009% | 9.61% | 49.75% | 53.29%
8.04 |  |  | 2.305% | 8.909% | 9.409% | 49.45% | 52.99%
8.05 |  |  | 2.305% | 8.809% | 9.309% | 49.25% | 52.79%
8.06 |  |  | 2.305% | 8.609% | 9.209% | 48.95% | 52.5%
8.07 |  |  | 2.305% | 8.509% | 9.009% | 48.75% | 52.2%
8.08 |  |  | 2.305% | 8.408% | 8.909% | 48.45% | 52%
8.09 |  |  | 2.305% | 8.208% | 8.809% | 48.25% | 51.7%
8.1 |  |  | 2.305% | 8.108% | 8.609% | 48.05% | 51.5%
8.11 |  |  | 2.305% | 8.008% | 8.509% | 47.75% | 51.2%
8.12 |  |  | 2.305% | 7.908% | 8.408% | 47.55% | 50.9%
8.13 |  |  | 2.305% | 7.808% | 8.308% | 47.25% | 50.7%
8.14 |  |  | 2.305% | 7.608% | 8.208% | 47.05% | 50.4%
8.15 |  |  | 2.305% | 7.508% | 8.108% | 46.75% | 50.2%
8.16 |  |  | 2.305% | 7.407% | 8.008% | 46.55% | 49.9%
8.17 |  |  | 2.305% | 7.307% | 7.908% | 46.35% | 49.6%
8.18 |  |  | 2.305% | 7.207% | 7.808% | 46.05% | 49.4%
8.19 |  |  | 2.305% | 7.107% | 7.608% | 45.85% | 49.1%
8.2 |  |  | 2.305% | 7.007% | 7.508% | 45.55% | 48.8%
8.21 |  |  | 2.305% | 6.907% | 7.407% | 45.35% | 48.6%
8.22 |  |  | 2.305% | 6.807% | 7.307% | 45.05% | 48.3%
8.23 |  |  | 2.305% | 6.707% | 7.207% | 44.84% | 48%
8.24 |  |  | 2.305% | 6.607% | 7.107% | 44.54% | 47.8%
8.25 |  |  | 2.305% | 6.507% | 7.007% | 44.34% | 47.5%
8.26 |  |  | 2.305% | 6.406% | 7.007% | 44.04% | 47.31%
8.27 |  |  | 2.204% | 6.306% | 6.907% | 43.84% | 47.01%
8.28 |  |  | 2.204% | 6.206% | 6.807% | 43.64% | 46.71%
8.29 |  |  | 2.204% | 6.206% | 6.707% | 43.34% | 46.51%
8.3 |  |  | 2.204% | 6.106% | 6.607% | 43.14% | 46.21%
8.31 |  |  | 2.204% | 6.006% | 6.507% | 42.84% | 46.01%
8.32 |  |  | 2.204% | 5.906% | 6.406% | 42.64% | 45.71%
8.33 |  |  | 2.204% | 5.806% | 6.406% | 42.34% | 45.51%
8.34 |  |  | 2.204% | 5.806% | 6.306% | 42.14% | 45.21%
8.35 |  |  | 2.204% | 5.706% | 6.206% | 41.84% | 44.91%
8.36 |  |  | 2.204% | 5.606% | 6.106% | 41.64% | 44.71%
8.37 |  |  | 2.204% | 5.606% | 6.106% | 41.34% | 44.41%
8.38 |  |  | 2.204% | 5.506% | 6.006% | 41.14% | 44.11%
8.39 |  |  | 2.204% | 5.405% | 6.006% | 40.84% | 43.91%
8.4 |  |  | 2.204% | 5.405% | 5.906% | 40.64% | 43.61%
8.41 |  |  | 2.204% | 5.305% | 5.806% | 40.34% | 43.31%
8.42 |  |  | 2.204% | 5.205% | 5.806% | 40.14% | 43.01%
8.43 |  |  | 2.204% | 5.205% | 5.706% | 39.84% | 42.81%
8.44 |  |  | 2.204% | 5.105% | 5.706% | 39.64% | 42.51%
8.45 |  |  | 2.204% | 5.105% | 5.606% | 39.44% | 42.22%
8.46 |  |  |  | 5.005% | 5.606% | 39.14% | 41.92%
8.47 |  |  |  | 5.005% | 5.506% | 38.94% | 41.62%
8.48 |  |  |  | 4.905% | 5.405% | 38.64% | 41.32%
8.49 |  |  |  | 4.905% | 5.405% | 38.44% | 41.12%
8.5 |  |  |  | 4.805% | 5.405% | 38.14% | 40.82%
8.51 |  |  |  | 4.805% | 5.305% | 37.94% | 40.52%
8.52 |  |  |  | 4.705% | 5.305% | 37.74% | 40.22%
8.53 |  |  |  | 4.705% | 5.205% | 37.44% | 39.92%
8.54 |  |  |  | 4.605% | 5.205% | 37.24% | 39.72%
8.55 |  |  |  | 4.605% | 5.105% | 36.94% | 39.42%
8.56 |  |  |  | 4.505% | 5.105% | 36.74% | 39.12%
8.57 |  |  |  | 4.505% | 5.005% | 36.44% | 38.82%
8.58 |  |  |  | 4.505% | 5.005% | 36.24% | 38.62%
8.59 |  |  |  | 4.404% | 5.005% | 36.04% | 38.32%
8.6 |  |  |  | 4.404% | 4.905% | 35.74% | 38.02%
8.61 |  |  |  | 4.404% | 4.905% | 35.54% | 37.72%
8.62 |  |  |  | 4.304% | 4.905% | 35.24% | 37.52%
8.63 |  |  |  | 4.304% | 4.805% | 35.04% | 37.23%
8.64 |  |  |  | 4.304% | 4.805% | 34.73% | 36.93%
8.65 |  |  |  | 4.204% | 4.805% | 34.53% | 36.63%
8.66 |  |  |  | 4.204% | 4.705% | 34.23% | 36.43%
8.67 |  |  |  | 4.204% | 4.705% | 34.03% | 36.13%
8.68 |  |  |  | 4.204% | 4.705% | 33.83% | 35.83%
8.69 |  |  |  | 4.104% | 4.705% | 33.53% | 35.63%
8.7 |  |  |  | 4.104% | 4.605% | 33.33% | 35.33%
8.71 |  |  |  | 4.104% | 4.605% | 33.03% | 35.13%
8.72 |  |  |  | 4.104% | 4.605% | 32.83% | 34.83%
8.73 |  |  |  | 4.004% | 4.605% | 32.63% | 34.63%
8.74 |  |  |  | 4.004% | 4.605% | 32.33% | 34.33%
8.75 |  |  |  | 4.004% | 4.505% | 32.13% | 34.13%
8.76 |  |  |  | 4.004% | 4.505% | 31.93% | 33.83%
8.77 |  |  |  | 4.004% | 4.505% | 31.63% | 33.63%
8.78 |  |  |  | 3.904% | 4.505% | 31.43% | 33.33%
8.79 |  |  |  | 3.904% | 4.505% | 31.23% | 33.13%
8.8 |  |  |  | 3.904% | 4.404% | 30.93% | 32.83%
8.81 |  |  |  | 3.904% | 4.404% | 30.73% | 32.63%
8.82 |  |  |  | 3.904% | 4.404% | 30.43% | 32.34%
8.83 |  |  |  | 3.904% | 4.404% | 30.23% | 32.14%
8.84 |  |  |  | 3.804% | 4.404% | 30.03% | 31.84%
8.85 |  |  |  | 3.804% | 4.404% | 29.73% | 31.64%
8.86 |  |  |  | 3.804% | 4.304% | 29.53% | 31.34%
8.87 |  |  |  | 3.804% | 4.304% | 29.33% | 31.04%
8.88 |  |  |  | 3.804% | 4.304% | 29.03% | 30.84%
8.89 |  |  |  | 3.804% | 4.304% | 28.83% | 30.54%
8.9 |  |  |  | 3.804% | 4.304% | 28.63% | 30.34%
8.91 |  |  |  | 3.804% | 4.304% | 28.43% | 30.04%
8.92 |  |  |  | 3.704% | 4.304% | 28.13% | 29.74%
8.93 |  |  |  | 3.704% | 4.204% | 27.93% | 29.54%
8.94 |  |  |  | 3.704% | 4.204% | 27.73% | 29.34%
8.95 |  |  |  | 3.704% | 4.204% | 27.53% | 29.04%
8.96 |  |  |  | 3.704% | 4.204% | 27.23% | 28.84%
8.97 |  |  |  | 3.704% | 4.204% | 27.03% | 28.64%
8.98 |  |  |  | 3.704% | 4.204% | 26.83% | 28.34%
8.99 |  |  |  | 3.704% | 4.204% | 26.63% | 28.14%
9 |  |  |  | 3.704% | 4.204% | 26.43% | 27.94%
9.01 |  |  |  | 3.704% | 4.204% | 26.13% | 27.64%
9.02 |  |  |  | 3.704% | 4.204% | 25.93% | 27.45%
9.03 |  |  |  | 3.604% | 4.104% | 25.73% | 27.25%
9.04 |  |  |  | 3.604% | 4.104% | 25.53% | 27.05%
9.05 |  |  |  | 3.604% | 4.104% | 25.33% | 26.75%
9.06 |  |  |  | 3.604% | 4.104% | 25.03% | 26.55%
9.07 |  |  |  | 3.604% | 4.104% | 24.82% | 26.25%
9.08 |  |  |  | 3.604% | 4.104% | 24.62% | 26.05%
9.09 |  |  |  | 3.604% | 4.104% | 24.42% | 25.85%
9.1 |  |  |  | 3.604% | 4.104% | 24.22% | 25.55%
9.11 |  |  |  | 3.604% | 4.104% | 24.02% | 25.35%
9.12 |  |  |  | 3.604% | 4.104% | 23.82% | 25.15%
9.13 |  |  |  | 3.604% | 4.104% | 23.62% | 24.85%
9.14 |  |  |  | 3.604% | 4.104% | 23.42% | 24.65%
9.15 |  |  |  | 3.604% | 4.104% | 23.22% | 24.45%
9.16 |  |  |  | 3.604% | 4.104% | 23.02% | 24.25%
9.17 |  |  |  | 3.604% | 4.104% | 22.82% | 24.05%
9.18 |  |  |  | 3.604% | 4.104% | 22.62% | 23.85%
9.19 |  |  |  | 3.604% | 4.104% | 22.42% | 23.65%
9.2 |  |  |  | 3.504% | 4.104% | 22.22% | 23.45%
9.21 |  |  |  | 3.504% | 4.004% | 22.02% | 23.25%
9.22 |  |  |  | 3.504% | 4.004% | 21.82% | 22.95%
9.23 |  |  |  | 3.504% | 4.004% | 21.62% | 22.75%
9.24 |  |  |  | 3.504% | 4.004% | 21.42% | 22.55%
9.25 |  |  |  | 3.504% | 4.004% | 21.22% | 22.36%
9.26 |  |  |  | 3.504% | 4.004% | 21.02% | 22.16%
9.27 |  |  |  | 3.504% | 4.004% | 20.82% | 21.96%
9.28 |  |  |  | 3.504% | 4.004% | 20.62% | 21.76%
9.29 |  |  |  | 3.504% | 4.004% | 20.52% | 21.56%
9.3 |  |  |  | 3.504% | 4.004% | 20.32% | 21.36%
9.31 |  |  |  | 3.504% | 4.004% | 20.12% | 21.16%
9.32 |  |  |  | 3.504% | 4.004% | 19.92% | 20.96%
9.33 |  |  |  | 3.504% | 4.004% | 19.72% | 20.76%
9.34 |  |  |  | 3.504% | 4.004% | 19.62% | 20.56%
9.35 |  |  |  | 3.504% | 4.004% | 19.42% | 20.36%
9.36 |  |  |  | 3.504% | 4.004% | 19.22% | 20.16%
9.37 |  |  |  | 3.504% | 4.004% | 19.02% | 19.96%
9.38 |  |  |  | 3.504% | 4.004% | 18.82% | 19.76%
9.39 |  |  |  | 3.504% | 4.004% | 18.72% | 19.56%
9.4 |  |  |  | 3.504% | 4.004% | 18.52% | 19.46%
9.41 |  |  |  | 3.504% | 4.004% | 18.32% | 19.26%
9.42 |  |  |  | 3.504% | 4.004% | 18.12% | 19.06%
9.43 |  |  |  | 3.504% | 4.004% | 18.02% | 18.86%
9.44 |  |  |  | 3.504% | 4.004% | 17.82% | 18.66%
9.45 |  |  |  | 3.504% | 4.004% | 17.62% | 18.46%
9.46 |  |  |  | 3.504% | 4.004% | 17.42% | 18.36%
9.47 |  |  |  | 3.504% | 4.004% | 17.32% | 18.16%
9.48 |  |  |  | 3.504% | 4.004% | 17.12% | 17.96%
9.49 |  |  |  | 3.504% | 4.004% | 16.92% | 17.76%
9.5 |  |  |  | 3.504% | 4.004% | 16.82% | 17.56%
9.51 |  |  |  | 3.504% | 4.004% | 16.62% | 17.37%
9.52 |  |  |  | 3.504% | 4.004% | 16.42% | 17.17%
9.53 |  |  |  | 3.504% | 4.004% | 16.22% | 16.97%
9.54 |  |  |  | 3.504% | 3.904% | 16.12% | 16.87%
9.55 |  |  |  | 3.504% | 3.904% | 15.92% | 16.67%
9.56 |  |  |  | 3.504% | 3.904% | 15.82% | 16.47%
9.57 |  |  |  | 3.504% | 3.904% | 15.62% | 16.27%
9.58 |  |  |  | 3.504% | 3.904% | 15.42% | 16.17%
9.59 |  |  |  | 3.504% | 3.904% | 15.32% | 15.97%
9.6 |  |  |  | 3.504% | 3.904% | 15.12% | 15.77%
9.61 |  |  |  | 3.504% | 3.904% | 15.02% | 15.67%
9.62 |  |  |  | 3.504% | 3.904% | 14.81% | 15.47%
9.63 |  |  |  | 3.504% | 3.904% | 14.71% | 15.27%
9.64 |  |  |  | 3.504% | 3.904% | 14.51% | 15.17%
9.65 |  |  |  | 3.403% | 3.904% | 14.41% | 15.07%
9.66 |  |  |  | 3.403% | 3.904% | 14.21% | 14.87%
9.67 |  |  |  | 3.403% | 3.904% | 14.11% | 14.77%
9.68 |  |  |  | 3.403% | 3.904% | 13.91% | 14.57%
9.69 |  |  |  | 3.403% | 3.904% | 13.81% | 14.47%
9.7 |  |  |  | 3.403% | 3.904% | 13.61% | 14.27%
9.71 |  |  |  | 3.403% | 3.904% | 13.51% | 14.17%
9.72 |  |  |  | 3.403% | 3.904% | 13.31% | 14.07%
9.73 |  |  |  | 3.403% | 3.904% | 13.21% | 13.87%
9.74 |  |  |  | 3.403% | 3.904% | 13.11% | 13.77%
9.75 |  |  |  | 3.403% | 3.904% | 12.91% | 13.57%
9.76 |  |  |  | 3.403% | 3.904% | 12.81% | 13.47%
9.77 |  |  |  | 3.403% | 3.904% | 12.61% | 13.37%
9.78 |  |  |  | 3.403% | 3.904% | 12.51% | 13.17%
9.79 |  |  |  | 3.403% | 3.904% | 12.41% | 13.07%
9.8 |  |  |  | 3.403% | 3.904% | 12.21% | 12.97%
9.81 |  |  |  | 3.403% | 3.904% | 12.11% | 12.77%
9.82 |  |  |  | 3.403% | 3.904% | 12.01% | 12.67%
9.83 |  |  |  | 3.403% | 3.904% | 11.91% | 12.57%
9.84 |  |  |  | 3.403% | 3.904% | 11.71% | 12.38%
9.85 |  |  |  | 3.403% | 3.904% | 11.61% | 12.28%
9.86 |  |  |  | 3.403% | 3.904% | 11.51% | 12.18%
9.87 |  |  |  | 3.403% | 3.904% | 11.41% | 11.98%
9.88 |  |  |  | 3.403% | 3.904% | 11.21% | 11.88%
9.89 |  |  |  | 3.403% | 3.904% | 11.11% | 11.78%
9.9 |  |  |  | 3.403% | 3.904% | 11.01% | 11.68%
9.91 |  |  |  | 3.403% | 3.904% | 10.91% | 11.58%
9.92 |  |  |  | 3.403% | 3.904% | 10.81% | 11.48%
9.93 |  |  |  | 3.403% | 3.904% | 10.71% | 11.38%
9.94 |  |  |  | 3.403% | 3.904% | 10.61% | 11.28%
9.95 |  |  |  | 3.403% | 3.904% | 10.51% | 11.18%
9.96 |  |  |  | 3.303% | 3.804% | 10.41% | 11.08%
9.97 |  |  |  | 3.303% | 3.804% | 10.31% | 10.98%
9.98 |  |  |  | 3.303% | 3.804% | 10.21% | 10.78%
9.99 |  |  |  | 3.303% | 3.804% | 10.11% | 10.68%
10 |  |  |  | 3.303% | 3.804% | 10.01% | 10.58%
10.01 |  |  |  | 3.303% | 3.804% | 9.91% | 10.48%
10.02 |  |  |  | 3.303% | 3.804% | 9.81% | 10.38%
10.03 |  |  |  | 3.303% | 3.804% | 9.71% | 10.28%
10.04 | [5 empty] | 9.61% | 10.18%
10.05 | [5 empty] | 9.51% | 10.08%
10.06 | [5 empty] | 9.409% | 9.98%
10.07 | [5 empty] | 9.309% | 9.88%
10.08 | [5 empty] | 9.209% | 9.78%
10.09 | [5 empty] | 9.109% | 9.78%
10.1 | [5 empty] | 9.109% | 9.681%
10.11 | [5 empty] | 9.009% | 9.581%
10.12 | [5 empty] | 8.909% | 9.481%
10.13 | [5 empty] | 8.809% | 9.381%
10.14 | [5 empty] | 8.709% | 9.281%
10.15 | [5 empty] | 8.709% | 9.281%
10.16 | [5 empty] | 8.609% | 9.182%
10.17 | [5 empty] | 8.509% | 9.082%
10.18 | [5 empty] | 8.408% | 9.082%
10.19 | [5 empty] | 8.408% | 8.982%
10.2 | [5 empty] | 8.308% | 8.882%
10.21 | [5 empty] | 8.208% | 8.882%
10.22 | [5 empty] | 8.208% | 8.782%
10.23 | [5 empty] | 8.108% | 8.683%
10.24 | [5 empty] | 8.008% | 8.683%
10.25 | [5 empty] | 8.008% | 8.583%
10.26 | [5 empty] | 7.908% | 8.483%
10.27 | [5 empty] | 7.808% | 8.483%
10.28 | [5 empty] | 7.808% | 8.383%
10.29 | [5 empty] | 7.708% | 8.383%
10.3 | [5 empty] | 7.708% | 8.283%
10.31 | [5 empty] | 7.608% | 8.184%
10.32 | [5 empty] | 7.608% | 8.184%
10.33 | [5 empty] | 7.508% | 8.084%
10.34 | [5 empty] | 7.407% | 8.084%
10.35 | [5 empty] | 7.407% | 7.984%
10.36 | [5 empty] | 7.307% | 7.984%
10.37 | [5 empty] | 7.307% | 7.884%
10.38 | [5 empty] | 7.207% | 7.784%
10.39 | [5 empty] | 7.207% | 7.784%
10.4 | [5 empty] | 7.107% | 7.685%
10.41 | [5 empty] | 7.107% | 7.685%
10.42 | [5 empty] | 7.007% | 7.585%
10.43 | [5 empty] | 7.007% | 7.585%
10.44 | [5 empty] | 7.007% | 7.485%
10.45 | [5 empty] | 6.907% | 7.485%
10.46 | [5 empty] | 6.907% | 7.385%
10.47 | [5 empty] | 6.807% | 7.385%
10.48 | [5 empty] | 6.807% | 7.385%
10.49 | [5 empty] | 6.707% | 7.285%
10.5 | [5 empty] | 6.707% | 7.285%
10.51 | [5 empty] | 6.707% | 7.186%
10.52 | [5 empty] | 6.607% | 7.186%
10.53 | [5 empty] | 6.607% | 7.186%
10.54 | [5 empty] | 6.507% | 7.086%
10.55 | [5 empty] | 6.507% | 7.086%
10.56 | [5 empty] | 6.507% | 7.086%
10.57 | [5 empty] | 6.406% | 6.986%
10.58 | [5 empty] | 6.406% | 6.986%
10.59 | [5 empty] | 6.406% | 6.986%
10.6 | [5 empty] | 6.306% | 6.886%
10.61 | [5 empty] | 6.306% | 6.886%
10.62 | [5 empty] | 6.306% | 6.886%
10.63 | [5 empty] | 6.206% | 6.786%
10.64 | [5 empty] | 6.206% | 6.786%
10.65 | [5 empty] | 6.206% | 6.786%
10.66 | [5 empty] | 6.106% | 6.687%
10.67 | [5 empty] | 6.106% | 6.687%
10.68 | [5 empty] | 6.106% | 6.687%
10.69 | [5 empty] | 6.106% | 6.587%
10.7 | [5 empty] | 6.006% | 6.587%
10.71 | [5 empty] | 6.006% | 6.587%
10.72 | [5 empty] | 6.006% | 6.587%
10.73 | [5 empty] | 5.906% | 6.487%
10.74 | [5 empty] | 5.906% | 6.487%
10.75 | [5 empty] | 5.906% | 6.487%
10.76 | [5 empty] | 5.906% | 6.487%
10.77 | [5 empty] | 5.806% | 6.487%
10.78 | [5 empty] | 5.806% | 6.387%
10.79 | [5 empty] | 5.806% | 6.387%
10.8 | [5 empty] | 5.806% | 6.387%
10.81 | [5 empty] | 5.706% | 6.387%
10.82 | [5 empty] | 5.706% | 6.387%
10.83 | [5 empty] | 5.706% | 6.287%
10.84 | [5 empty] | 5.706% | 6.287%
10.85 | [5 empty] | 5.706% | 6.287%
10.86 | [5 empty] | 5.606% | 6.287%
10.87 | [5 empty] | 5.606% | 6.287%
10.88 | [5 empty] | 5.606% | 6.287%
10.89 | [5 empty] | 5.606% | 6.188%
10.9 | [5 empty] | 5.606% | 6.188%
10.91 | [5 empty] | 5.606% | 6.188%
10.92 | [5 empty] | 5.506% | 6.188%
10.93 | [5 empty] | 5.506% | 6.188%
10.94 | [5 empty] | 5.506% | 6.188%
10.95 | [5 empty] | 5.506% | 6.088%
10.96 | [5 empty] | 5.506% | 6.088%
10.97 | [5 empty] | 5.506% | 6.088%
10.98 | [5 empty] | 5.506% | 6.088%
10.99 | [5 empty] | 5.506% | 6.088%
11 | [5 empty] | 5.405% | 6.088%
11.01 | [5 empty] | 5.405% | 6.088%
11.02 | [5 empty] | 5.405% | 6.088%
11.03 | [5 empty] | 5.405% | 6.088%
11.04 | [5 empty] | 5.405% | 6.088%
11.05 | [5 empty] | 5.405% | 5.988%
11.06 | [5 empty] | 5.405% | 5.988%
11.07 | [5 empty] | 5.405% | 5.988%
11.08 | [5 empty] | 5.405% | 5.988%
11.09 | [5 empty] | 5.405% | 5.988%
11.1 | [5 empty] | 5.305% | 5.988%
11.11 | [5 empty] | 5.305% | 5.988%
11.12 | [5 empty] | 5.305% | 5.988%
11.13 | [5 empty] | 5.305% | 5.988%
11.14 | [5 empty] | 5.305% | 5.988%
11.15 | [5 empty] | 5.305% | 5.988%
11.16 | [5 empty] | 5.305% | 5.988%
11.17 | [5 empty] | 5.305% | 5.988%
11.18 | [5 empty] | 5.305% | 5.988%
11.19 | [5 empty] | 5.305% | 5.988%
11.2 | [5 empty] | 5.305% | 5.888%
11.21 | [5 empty] | 5.305% | 5.888%
11.22 | [5 empty] | 5.305% | 5.888%
11.23 | [5 empty] | 5.305% | 5.888%
11.24 | [5 empty] | 5.305% | 5.888%
11.25 | [5 empty] | 5.305% | 5.888%
11.26 | [5 empty] | 5.205% | 5.888%
11.27 | [5 empty] | 5.205% | 5.888%
11.28 | [5 empty] | 5.205% | 5.888%
11.29 | [5 empty] | 5.205% | 5.888%
11.3 | [5 empty] | 5.205% | 5.888%
11.31 | [5 empty] | 5.205% | 5.788%
11.32 | [5 empty] | 5.205% | 5.788%
11.33 | [5 empty] | 5.205% | 5.788%
11.34 | [5 empty] | 5.205% | 5.788%
11.35 | [5 empty] | 5.205% | 5.788%
11.36 | [5 empty] | 5.205% | 5.788%
11.37 | [5 empty] | 5.205% | 5.788%
11.38 | [5 empty] | 5.205% | 5.788%
11.39 | [5 empty] | 5.205% | 5.788%
11.4 | [5 empty] | 5.205% | 5.788%
11.41 | [5 empty] | 5.205% | 5.788%
11.42 | [5 empty] | 5.205% | 5.788%
11.43 | [5 empty] | 5.205% | 5.788%
11.44 | [5 empty] | 5.205% | 5.788%
11.45 | [5 empty] | 5.205% | 5.788%
11.46 | [5 empty] | 5.205% | 5.788%
11.47 | [5 empty] | 5.205% | 5.788%
11.48 | [5 empty] | 5.205% | 5.788%
11.49 | [5 empty] | 5.205% | 5.788%
11.5 | [5 empty] | 5.205% | 5.788%
11.51 | [5 empty] | 5.205% | 5.788%
11.52 | [5 empty] | 5.105% | 5.788%
11.53 | [5 empty] | 5.105% | 5.788%
11.54 | [5 empty] | 5.105% | 5.788%
11.55 | [5 empty] | 5.105% | 5.788%
11.56 | [5 empty] | 5.105% | 5.689%
11.57 | [5 empty] | 5.105% | 5.689%
11.58 | [5 empty] | 5.105% | 5.689%
11.59 | [5 empty] | 5.105% | 5.689%
11.6 | [5 empty] | 5.105% | 5.689%
11.61 | [5 empty] | 5.105% | 5.689%
11.62 | [5 empty] | 5.105% | 5.689%
11.63 | [5 empty] | 5.105% | 5.689%
11.64 | [5 empty] | 5.105% | 5.689%
11.65 | [5 empty] | 5.105% | 5.689%
11.66 | [5 empty] | 5.105% | 5.689%
11.67 | [5 empty] | 5.105% | 5.689%
11.68 | [5 empty] | 5.105% | 5.689%
11.69 | [5 empty] | 5.105% | 5.689%
11.7 | [5 empty] | 5.105% | 5.689%
11.71 | [5 empty] | 5.105% | 5.689%
11.72 | [5 empty] | 5.105% | 5.689%
11.73 | [5 empty] | 5.105% | 5.689%
11.74 | [5 empty] | 5.105% | 5.689%
11.75 | [5 empty] | 5.105% | 5.689%
11.76 | [5 empty] | 5.105% | 5.689%
11.77 | [5 empty] | 5.005% | 5.689%
11.78 | [5 empty] | 5.005% | 5.689%
11.79 | [5 empty] | 5.005% | 5.689%
11.8 | [5 empty] | 5.005% | 5.689%
11.81 | [5 empty] | 5.005% | 5.689%
11.82 | [5 empty] | 5.005% | 5.689%
11.83 | [5 empty] | 5.005% | 5.689%
11.84 | [5 empty] | 5.005% | 5.689%
11.85 | [5 empty] | 5.005% | 5.589%
11.86 | [5 empty] | 5.005% | 5.589%
11.87 | [5 empty] | 5.005% | 5.589%
11.88 | [5 empty] | 5.005% | 5.589%
11.89 | [5 empty] | 5.005% | 5.589%
11.9 | [5 empty] | 5.005% | 5.589%
11.91 | [5 empty] | 5.005% | 5.589%
11.92 | [5 empty] | 5.005% | 5.589%
11.93 | [5 empty] | 5.005% | 5.589%
11.94 | [5 empty] | 5.005% | 5.589%
11.95 | [5 empty] | 5.005% | 5.589%
11.96 | [5 empty] | 5.005% | 5.589%
11.97 | [5 empty] | 5.005% | 5.589%
11.98 | [5 empty] | 5.005% | 5.589%
11.99 | [5 empty] | 5.005% | 5.589%
12 | [5 empty] | 5.005% | 5.589%
12.01 | [5 empty] | 5.005% | 5.589%
12.02 | [5 empty] | 5.005% | 5.589%
12.03 | [5 empty] | 5.005% | 5.589%
12.04 | [5 empty] | 5.005% | 5.589%
12.05 | [5 empty] | 5.005% | 5.589%
12.06 | [5 empty] | 5.005% | 5.589%
12.07 | [5 empty] | 5.005% | 5.589%
12.08 | [5 empty] | 5.005% | 5.589%
12.09 | [5 empty] | 5.005% | 5.589%
12.1 | [5 empty] | 5.005% | 5.589%
12.11 | [5 empty] | 5.005% | 5.589%
12.12 | [5 empty] | 4.905% | 5.589%
12.13 | [5 empty] | 4.905% | 5.589%
12.14 | [5 empty] | 4.905% | 5.589%
12.15 | [5 empty] | 4.905% | 5.589%
12.16 | [5 empty] | 4.905% | 5.489%
12.17 | [5 empty] | 4.905% | 5.489%
12.18 | [5 empty] | 4.905% | 5.489%
12.19 | [5 empty] | 4.905% | 5.489%
12.2 | [5 empty] | 4.905% | 5.489%
12.21 | [5 empty] | 4.905% | 5.489%
12.22 | [5 empty] | 4.905% | 5.489%
12.23 | [5 empty] | 4.905% | 5.489%
12.24 | [5 empty] | 4.905% | 5.489%
12.25 | [5 empty] | 4.905% | 5.489%
12.26 | [5 empty] | 4.905% | 5.489%
12.27 | [5 empty] | 4.905% | 5.489%
12.28 | [5 empty] | 4.905% | 5.489%
12.29 | [5 empty] | 4.905% | 5.489%
12.3 | [5 empty] | 4.905% | 5.489%
12.31 | [5 empty] | 4.905% | 5.489%
12.32 | [5 empty] | 4.905% | 5.489%
12.33 | [5 empty] | 4.905% | 5.489%
12.34 | [5 empty] | 4.905% | 5.489%
12.35 | [5 empty] | 4.905% | 5.489%
12.36 | [5 empty] | 4.905% | 5.489%
12.37 | [5 empty] | 4.905% | 5.489%
12.38 | [5 empty] | 4.905% | 5.489%
12.39 | [5 empty] | 4.905% | 5.489%
12.4 | [5 empty] | 4.905% | 5.489%
12.41 | [5 empty] | 4.805% | 5.489%
12.42 | [5 empty] | 4.805% | 5.489%
12.43 | [5 empty] | 4.805% | 5.489%
12.44 | [5 empty] | 4.805% | 5.489%
12.45 | [5 empty] | 4.805% | 5.489%
12.46 | [5 empty] | 4.805% | 5.489%
12.47 | [5 empty] | 4.805% | 5.489%
12.48 | [5 empty] | 4.805% | 5.489%
12.49 | [5 empty] | 4.805% | 5.489%
12.5 | [5 empty] | 4.805% | 5.489%
12.51 | [5 empty] | 4.805% | 5.489%
12.52 | [5 empty] | 4.805% | 5.489%
12.53 | [5 empty] | 4.805% | 5.489%
12.54 | [5 empty] | 4.805% | 5.389%
12.55 | [5 empty] | 4.805% | 5.389%
12.56 | [5 empty] | 4.805% | 5.389%
12.57 | [5 empty] | 4.805% | 5.389%
12.58 | [5 empty] | 4.805% | 5.389%
12.59 | [5 empty] | 4.805% | 5.389%
12.6 | [5 empty] | 4.805% | 5.389%
12.61 | [5 empty] | 4.805% | 5.389%
12.62 | [5 empty] | 4.805% | 5.389%
12.63 | [5 empty] | 4.805% | 5.389%
12.64 | [5 empty] | 4.805% | 5.389%
12.65 | [5 empty] | 4.805% | 5.389%
12.66 | [5 empty] | 4.805% | 5.389%
12.67 | [5 empty] | 4.805% | 5.389%
12.68 | [5 empty] | 4.805% | 5.389%
12.69 | [5 empty] | 4.805% | 5.389%
12.7 | [5 empty] | 4.805% | 5.389%
12.71 | [5 empty] | 4.805% | 5.389%
12.72 | [5 empty] | 4.805% | 5.389%
12.73 | [5 empty] | 4.805% | 5.389%
12.74 | [5 empty] | 4.805% | 5.389%
12.75 | [5 empty] | 4.805% | 5.389%
12.76 | [5 empty] | 4.805% | 5.389%
12.77 | [5 empty] | 4.805% | 5.389%
12.78 | [5 empty] | 4.805% | 5.389%
12.79 | [5 empty] | 4.805% | 5.389%
12.8 | [5 empty] | 4.805% | 5.389%
12.81 | [5 empty] | 4.805% | 5.389%
12.82 | [5 empty] | 4.705% | 5.389%
12.83 | [5 empty] | 4.705% | 5.389%
12.84 | [5 empty] | 4.705% | 5.389%
12.85 | [5 empty] | 4.705% | 5.389%
12.86 | [5 empty] | 4.705% | 5.389%
12.87 | [5 empty] | 4.705% | 5.289%
12.88 | [5 empty] | 4.705% | 5.289%
12.89 | [5 empty] | 4.705% | 5.289%
12.9 | [5 empty] | 4.705% | 5.289%
12.91 | [5 empty] | 4.705% | 5.289%
12.92 | [5 empty] | 4.705% | 5.289%
12.93 | [5 empty] | 4.705% | 5.289%
12.94 | [5 empty] | 4.705% | 5.289%
12.95 | [5 empty] | 4.705% | 5.289%
12.96 | [5 empty] | 4.705% | 5.289%
12.97 | [5 empty] | 4.705% | 5.289%
12.98 | [5 empty] | 4.705% | 5.289%
12.99 | [5 empty] | 4.705% | 5.289%
13 | [5 empty] | 4.705% | 5.289%
13.01 | [5 empty] | 4.705% | 5.289%
13.02 | [5 empty] | 4.705% | 5.289%
13.03 | [5 empty] | 4.705% | 5.289%
13.04 | [5 empty] | 4.705% | 5.289%
13.05 | [5 empty] | 4.705% | 5.289%
13.06 | [5 empty] | 4.705% | 5.289%
13.07 | [5 empty] | 4.705% | 5.289%
13.08 | [5 empty] | 4.705% | 5.289%
13.09 | [5 empty] | 4.705% | 5.289%
13.1 | [5 empty] | 4.705% | 5.289%
13.11 | [5 empty] | 4.705% | 5.289%
13.12 | [5 empty] | 4.705% | 5.289%
13.13 | [5 empty] | 4.705% | 5.289%
13.14 | [5 empty] | 4.705% | 5.289%
13.15 | [5 empty] | 4.705% | 5.289%
13.16 | [5 empty] | 4.705% | 5.289%
13.17 | [5 empty] | 4.705% | 5.289%
13.18 | [5 empty] | 4.705% | 5.289%
13.19 | [5 empty] | 4.605% | 5.289%
13.2 | [5 empty] | 4.605% | 5.289%
13.21 | [5 empty] | 4.605% | 5.289%
13.22 | [5 empty] | 4.605% | 5.289%
13.23 | [5 empty] | 4.605% | 5.289%
13.24 | [5 empty] | 4.605% | 5.289%
13.25 | [5 empty] | 4.605% | 5.289%
13.26 | [5 empty] | 4.605% | 5.289%
13.27 | [5 empty] | 4.605% | 5.289%
13.28 | [5 empty] | 4.605% | 5.289%
13.29 | [5 empty] | 4.605% | 5.289%
13.3 | [5 empty] | 4.605% | 5.289%
13.31 | [5 empty] | 4.605% | 5.19%
13.32 | [5 empty] | 4.605% | 5.19%
13.33 | [5 empty] | 4.605% | 5.19%
13.34 | [5 empty] | 4.605% | 5.19%
13.35 | [5 empty] | 4.605% | 5.19%
13.36 | [5 empty] | 4.605% | 5.19%
13.37 | [5 empty] | 4.605% | 5.19%
13.38 | [5 empty] | 4.605% | 5.19%
13.39 | [5 empty] | 4.605% | 5.19%
13.4 | [5 empty] | 4.605% | 5.19%
13.41 | [5 empty] | 4.605% | 5.19%
13.42 | [5 empty] | 4.605% | 5.19%
13.43 | [5 empty] | 4.605% | 5.19%
13.44 | [5 empty] | 4.605% | 5.19%
13.45 | [5 empty] | 4.505% | 5.19%
13.46 | [5 empty] | 4.505% | 5.19%
13.47 | [5 empty] | 4.505% | 5.19%
13.48 | [5 empty] | 4.505% | 5.19%
13.49 | [5 empty] | 4.505% | 5.19%
13.5 | [5 empty] | 4.505% | 5.19%
13.51 | [5 empty] | 4.505% | 5.19%
13.52 | [5 empty] | 4.505% | 5.19%
13.53 | [5 empty] | 4.505% | 5.19%
13.54 | [5 empty] | 4.505% | 5.19%
13.55 | [5 empty] | 4.505% | 5.19%
13.56 | [5 empty] | 4.505% | 5.19%
13.57 | [5 empty] | 4.505% | 5.19%
13.58 | [5 empty] | 4.505% | 5.19%
13.59 | [5 empty] | 4.505% | 5.09%
13.6 | [5 empty] | 4.505% | 5.09%
13.61 | [5 empty] | 4.505% | 5.09%
13.62 | [5 empty] | 4.505% | 5.09%
13.63 | [5 empty] | 4.505% | 5.09%
13.64 | [5 empty] | 4.505% | 5.09%
13.65 | [5 empty] | 4.505% | 5.09%
13.66 | [5 empty] | 4.505% | 5.09%
13.67 | [5 empty] | 4.505% | 5.09%
13.68 | [5 empty] | 4.505% | 5.09%
13.69 | [5 empty] | 4.505% | 5.09%
13.7 | [5 empty] | 4.505% | 5.09%
13.71 | [5 empty] | 4.505% | 5.09%
13.72 | [5 empty] | 4.505% | 5.09%
13.73 | [5 empty] | 4.505% | 5.09%
13.74 | [5 empty] | 4.505% | 5.09%
13.75 | [5 empty] | 4.505% | 5.09%
13.76 | [5 empty] | 4.505% | 5.09%
13.77 | [5 empty] | 4.505% | 5.09%
13.78 | [5 empty] | 4.505% | 5.09%
13.79 | [5 empty] | 4.505% | 5.09%
13.8 | [5 empty] | 4.404% | 5.09%
13.81 | [5 empty] | 4.404% | 5.09%
13.82 | [5 empty] | 4.404% | 5.09%
13.83 | [5 empty] | 4.404% | 5.09%
13.84 | [5 empty] | 4.404% | 5.09%
13.85 | [5 empty] | 4.404% | 5.09%
13.86 | [5 empty] | 4.404% | 5.09%
13.87 | [5 empty] | 4.404% | 5.09%
13.88 | [5 empty] | 4.404% | 5.09%
13.89 | [5 empty] | 4.404% | 5.09%
13.9 | [5 empty] | 4.404% | 5.09%
13.91 | [5 empty] | 4.404% | 5.09%
13.92 | [5 empty] | 4.404% | 5.09%
13.93 | [5 empty] | 4.404% | 5.09%
13.94 | [5 empty] | 4.404% | 5.09%
13.95 | [5 empty] | 4.404% | 5.09%
13.96 | [5 empty] | 4.404% | 5.09%
13.97 | [5 empty] | 4.404% | 5.09%
13.98 | [5 empty] | 4.404% | 5.09%
13.99 | [5 empty] | 4.404% | 5.09%
14 | [5 empty] | 4.404% | 4.99%
14.01 | [5 empty] | 4.404% | 4.99%
14.02 | [5 empty] | 4.404% | 4.99%
14.03 | [5 empty] | 4.404% | 4.99%
14.04 | [5 empty] | 4.404% | 4.99%
14.05 | [5 empty] | 4.404% | 4.99%
14.06 | [5 empty] | 4.404% | 4.99%
14.07 | [5 empty] | 4.404% | 4.99%
14.08 | [5 empty] | 4.404% | 4.99%
14.09 | [5 empty] | 4.404% | 4.99%
14.1 | [5 empty] | 4.404% | 4.99%
14.11 | [5 empty] | 4.404% | 4.99%
14.12 | [5 empty] | 4.404% | 4.99%
14.13 | [5 empty] | 4.404% | 4.99%
14.14 | [5 empty] | 4.404% | 4.99%
14.15 | [5 empty] | 4.304% | 4.99%
14.16 | [5 empty] | 4.304% | 4.99%
14.17 | [5 empty] | 4.304% | 4.99%
14.18 | [5 empty] | 4.304% | 4.99%
14.19 | [5 empty] | 4.304% | 4.99%
14.2 | [5 empty] | 4.304% | 4.99%
14.21 | [5 empty] | 4.304% | 4.99%
14.22 | [5 empty] | 4.304% | 4.99%
14.23 | [5 empty] | 4.304% | 4.99%
14.24 | [5 empty] | 4.304% | 4.99%
14.25 | [5 empty] | 4.304% | 4.99%
14.26 | [5 empty] | 4.304% | 4.99%
14.27 | [5 empty] | 4.304% | 4.99%
14.28 | [5 empty] | 4.304% | 4.99%
14.29 | [5 empty] | 4.304% | 4.99%
14.3 | [5 empty] | 4.304% | 4.99%
14.31 | [5 empty] | 4.304% | 4.99%
14.32 | [5 empty] | 4.304% | 4.99%
14.33 | [5 empty] | 4.304% | 4.99%
14.34 | [5 empty] | 4.304% | 4.99%
14.35 | [5 empty] | 4.304% | 4.99%
14.36 | [5 empty] | 4.304% | 4.99%
14.37 | [5 empty] | 4.304% | 4.99%
14.38 | [5 empty] | 4.304% | 4.99%
14.39 | [5 empty] | 4.304% | 4.99%
14.4 | [5 empty] | 4.304% | 4.89%
14.41 | [5 empty] | 4.304% | 4.89%
14.42 | [5 empty] | 4.304% | 4.89%
14.43 | [5 empty] | 4.304% | 4.89%
14.44 | [5 empty] | 4.304% | 4.89%
14.45 | [5 empty] | 4.304% | 4.89%
14.46 | [5 empty] | 4.304% | 4.89%
14.47 | [5 empty] | 4.304% | 4.89%
14.48 | [5 empty] | 4.304% | 4.89%
14.49 | [5 empty] | 4.304% | 4.89%
14.5 | [5 empty] | 4.304% | 4.89%
14.51 | [5 empty] | 4.304% | 4.89%
14.52 | [5 empty] | 4.304% | 4.89%
14.53 | [5 empty] | 4.304% | 4.89%
14.54 | [5 empty] | 4.304% | 4.89%
14.55 | [5 empty] | 4.304% | 4.89%
14.56 | [5 empty] | 4.304% | 4.89%
14.57 | [5 empty] | 4.304% | 4.89%
14.58 | [5 empty] | 4.304% | 4.89%
14.59 | [5 empty] | 4.304% | 4.89%
14.6 | [5 empty] | 4.304% | 4.89%
14.61 | [5 empty] | 4.304% | 4.89%
14.62 | [5 empty] | 4.304% | 4.89%
14.63 | [5 empty] | 4.304% | 4.89%
14.64 | [5 empty] | 4.304% | 4.89%
14.65 | [5 empty] | 4.204% | 4.89%
14.66 | [5 empty] | 4.204% | 4.89%
14.67 | [5 empty] | 4.204% | 4.89%
14.68 | [5 empty] | 4.204% | 4.89%
14.69 | [5 empty] | 4.204% | 4.89%
14.7 | [5 empty] | 4.204% | 4.89%
14.71 | [5 empty] | 4.204% | 4.89%
14.72 | [5 empty] | 4.204% | 4.89%
14.73 | [5 empty] | 4.204% | 4.89%
14.74 | [5 empty] | 4.204% | 4.89%
14.75 | [5 empty] | 4.204% | 4.89%
14.76 | [5 empty] | 4.204% | 4.89%
14.77 | [5 empty] | 4.204% | 4.89%
14.78 | [5 empty] | 4.204% | 4.89%
14.79 | [5 empty] | 4.204% | 4.89%
14.8 | [5 empty] | 4.204% | 4.89%
14.81 | [5 empty] | 4.204% | 4.89%
14.82 | [5 empty] | 4.204% | 4.89%
14.83 | [5 empty] | 4.204% | 4.79%
14.84 | [5 empty] | 4.204% | 4.79%
14.85 | [5 empty] | 4.204% | 4.79%
14.86 | [5 empty] | 4.204% | 4.79%
14.87 | [5 empty] | 4.204% | 4.79%
14.88 | [5 empty] | 4.204% | 4.79%
14.89 | [5 empty] | 4.204% | 4.79%
14.9 | [5 empty] | 4.104% | 4.79%
14.91 | [5 empty] | 4.104% | 4.79%
14.92 | [5 empty] | 4.104% | 4.79%
14.93 | [5 empty] | 4.104% | 4.79%
14.94 | [5 empty] | 4.104% | 4.79%
14.95 | [5 empty] | 4.104% | 4.79%
14.96 | [5 empty] | 4.104% | 4.79%
14.97 | [5 empty] | 4.104% | 4.79%
14.98 | [5 empty] | 4.104% | 4.79%
14.99 | [5 empty] | 4.104% | 4.79%
15 | [5 empty] | 4.104% | 4.79%
15.01 | [5 empty] | 4.104% | 4.79%
15.02 | [5 empty] | 4.104% | 4.79%
15.03 | [5 empty] | 4.104% | 4.79%
15.04 | [5 empty] | 4.104% | 4.79%
15.05 | [5 empty] | 4.104% | 4.79%
15.06 | [5 empty] | 4.104% | 4.79%
15.07 | [5 empty] | 4.104% | 4.79%
15.08 | [5 empty] | 4.104% | 4.79%
15.09 | [5 empty] | 4.104% | 4.79%
15.1 | [5 empty] | 4.104% | 4.79%
15.11 | [5 empty] | 4.104% | 4.79%
15.12 | [5 empty] | 4.104% | 4.79%
15.13 | [5 empty] | 4.104% | 4.79%
15.14 | [5 empty] | 4.104% | 4.79%
15.15 | [5 empty] | 4.104% | 4.691%
15.16 | [5 empty] | 4.104% | 4.691%
15.17 | [5 empty] | 4.104% | 4.691%
15.18 | [5 empty] | 4.104% | 4.691%
15.19 | [5 empty] | 4.104% | 4.691%
15.2 | [5 empty] | 4.104% | 4.691%
15.21 | [5 empty] | 4.104% | 4.691%
15.22 | [5 empty] | 4.104% | 4.691%
15.23 | [5 empty] | 4.104% | 4.691%
15.24 | [5 empty] | 4.104% | 4.691%
15.25 | [5 empty] | 4.104% | 4.691%
15.26 | [5 empty] | 4.104% | 4.691%
15.27 | [5 empty] | 4.104% | 4.691%
15.28 | [5 empty] | 4.104% | 4.691%
15.29 | [5 empty] | 4.104% | 4.691%
15.3 | [5 empty] | 4.104% | 4.691%
15.31 | [5 empty] | 4.104% | 4.691%
15.32 | [5 empty] | 4.104% | 4.691%
15.33 | [5 empty] | 4.104% | 4.691%
15.34 | [5 empty] | 4.104% | 4.691%
15.35 | [5 empty] | 4.104% | 4.691%
15.36 | [5 empty] | 4.104% | 4.691%
15.37 | [5 empty] | 4.104% | 4.691%
15.38 | [5 empty] | 4.104% | 4.691%
15.39 | [5 empty] | 4.104% | 4.691%
15.4 | [5 empty] | 4.104% | 4.691%
15.41 | [5 empty] | 4.104% | 4.691%
15.42 | [5 empty] | 4.104% | 4.691%
15.43 | [5 empty] | 4.104% | 4.691%
15.44 | [5 empty] | 4.104% | 4.691%
15.45 | [5 empty] | 4.104% | 4.691%
15.46 | [5 empty] | 4.004% | 4.691%
15.47 | [5 empty] | 4.004% | 4.691%
15.48 | [5 empty] | 4.004% | 4.691%
15.49 | [5 empty] | 4.004% | 4.691%
15.5 | [5 empty] | 4.004% | 4.691%
15.51 | [5 empty] | 4.004% | 4.691%
15.52 | [5 empty] | 4.004% | 4.691%
15.53 | [5 empty] | 4.004% | 4.691%
15.54 | [5 empty] | 4.004% | 4.691%
15.55 | [5 empty] | 4.004% | 4.691%
15.56 | [5 empty] | 4.004% | 4.691%
15.57 | [5 empty] | 4.004% | 4.691%
15.58 | [5 empty] | 4.004% | 4.691%
15.59 | [5 empty] | 4.004% | 4.691%
15.6 | [5 empty] | 4.004% | 4.691%
15.61 | [5 empty] | 4.004% | 4.691%
15.62 | [5 empty] | 4.004% | 4.691%
15.63 | [5 empty] | 4.004% | 4.691%
15.64 | [5 empty] | 4.004% | 4.691%
15.65 | [5 empty] | 4.004% | 4.691%
15.66 | [5 empty] | 4.004% | 4.691%
15.67 | [5 empty] | 4.004% | 4.691%
15.68 | [5 empty] | 4.004% | 4.591%
15.69 | [5 empty] | 4.004% | 4.591%
15.7 | [5 empty] | 4.004% | 4.591%
15.71 | [5 empty] | 4.004% | 4.591%
15.72 | [5 empty] | 4.004% | 4.591%
15.73 | [5 empty] | 4.004% | 4.591%
15.74 | [5 empty] | 4.004% | 4.591%
15.75 | [5 empty] | 4.004% | 4.591%
15.76 | [5 empty] | 4.004% | 4.591%
15.77 | [5 empty] | 4.004% | 4.591%
15.78 | [5 empty] | 4.004% | 4.591%
15.79 | [5 empty] | 4.004% | 4.591%
15.8 | [5 empty] | 4.004% | 4.591%
15.81 | [5 empty] | 4.004% | 4.591%
15.82 | [5 empty] | 4.004% | 4.591%
15.83 | [5 empty] | 4.004% | 4.591%
15.84 | [5 empty] | 4.004% | 4.591%
15.85 | [5 empty] | 4.004% | 4.591%
15.86 | [5 empty] | 4.004% | 4.591%
15.87 | [5 empty] | 4.004% | 4.591%
15.88 | [5 empty] | 4.004% | 4.591%
15.89 | [5 empty] | 4.004% | 4.591%
15.9 | [5 empty] | 4.004% | 4.591%

## 6 MeV PDDs
Energy: | 6 MeV
SSD: | 100 cm
| PDD | PDD
Depth | 4x4+ | 3x3
0.02 | 77.6% | 80.7% | Energy | SSD | Field size inline | R100 | R50 | R95 | R90 | R80 | Rp | Ds | PDD(1.38) | PDD(1.3)
0.03 | 77.7% | 80.8% | 6 MeV | 100.00 cm | 3x3 | 1.26 cm | 2.41 cm | 1.64 cm | 1.78 cm | 1.98 cm | 3.03 cm | 81% | 99.5% | 99.8%
0.04 | 77.8% | 80.9% | 6 MeV | 100.00 cm | 4x4+ | 1.38 cm | 2.41 cm | 1.69 cm | 1.82 cm | 2.01 cm | 2.99 cm | 77.8% | 100% | 99.8%
0.05 | 77.8% | 81%
0.06 | 77.9% | 81%
0.07 | 77.9% | 81.1%
0.08 | 78% | 81.2%
0.09 | 78.1% | 81.3%
0.1 | 78.2% | 81.4%
0.11 | 78.3% | 81.5%
0.12 | 78.3% | 81.5%
0.13 | 78.4% | 81.6%
0.14 | 78.5% | 81.7%
0.15 | 78.6% | 81.8%
0.16 | 78.7% | 81.9%
0.17 | 78.8% | 82%
0.18 | 79% | 82%
0.19 | 79.1% | 82.2%
0.2 | 79.3% | 82.4%
0.21 | 79.5% | 82.6%
0.22 | 79.7% | 82.7%
0.23 | 79.9% | 82.9%
0.24 | 80.1% | 83.1%
0.25 | 80.2% | 83.3%
0.26 | 80.5% | 83.4%
0.27 | 80.7% | 83.6%
0.28 | 80.9% | 83.8%
0.29 | 81.1% | 84%
0.3 | 81.3% | 84.2%
0.31 | 81.5% | 84.5%
0.32 | 81.7% | 84.7%
0.33 | 81.9% | 84.9%
0.34 | 82.1% | 85.1%
0.35 | 82.3% | 85.3%
0.36 | 82.5% | 85.5%
0.37 | 82.7% | 85.7%
0.38 | 82.9% | 85.9%
0.39 | 83.1% | 86.1%
0.4 | 83.4% | 86.3%
0.41 | 83.6% | 86.5%
0.42 | 83.8% | 86.7%
0.43 | 84% | 86.9%
0.44 | 84.2% | 87.2%
0.45 | 84.4% | 87.4%
0.46 | 84.6% | 87.6%
0.47 | 84.8% | 87.8%
0.48 | 85% | 88%
0.49 | 85.2% | 88.3%
0.5 | 85.4% | 88.5%
0.51 | 85.7% | 88.7%
0.52 | 85.9% | 88.9%
0.53 | 86.1% | 89.1%
0.54 | 86.3% | 89.3%
0.55 | 86.5% | 89.5%
0.56 | 86.7% | 89.7%
0.57 | 86.9% | 89.9%
0.58 | 87.2% | 90.1%
0.59 | 87.4% | 90.3%
0.6 | 87.6% | 90.5%
0.61 | 87.8% | 90.6%
0.62 | 88% | 90.8%
0.63 | 88.2% | 91%
0.64 | 88.4% | 91.2%
0.65 | 88.6% | 91.4%
0.66 | 88.8% | 91.6%
0.67 | 89% | 91.8%
0.68 | 89.2% | 92%
0.69 | 89.4% | 92.2%
0.7 | 89.7% | 92.4%
0.71 | 89.9% | 92.6%
0.72 | 90.1% | 92.8%
0.73 | 90.3% | 93%
0.74 | 90.5% | 93.2%
0.75 | 90.7% | 93.4%
0.76 | 91% | 93.6%
0.77 | 91.2% | 93.8%
0.78 | 91.4% | 93.9%
0.79 | 91.6% | 94.1%
0.8 | 91.8% | 94.3%
0.81 | 92% | 94.4%
0.82 | 92.3% | 94.6%
0.83 | 92.5% | 94.8%
0.84 | 92.7% | 95%
0.85 | 92.9% | 95.1%
0.86 | 93.1% | 95.3%
0.87 | 93.3% | 95.5%
0.88 | 93.5% | 95.7%
0.89 | 93.7% | 95.9%
0.9 | 93.9% | 96%
0.91 | 94.1% | 96.2%
0.92 | 94.3% | 96.4%
0.93 | 94.5% | 96.5%
0.94 | 94.7% | 96.7%
0.95 | 94.9% | 96.8%
0.96 | 95.2% | 97%
0.97 | 95.4% | 97.1%
0.98 | 95.6% | 97.3%
0.99 | 95.8% | 97.4%
1 | 96% | 97.6%
1.01 | 96.2% | 97.7%
1.02 | 96.3% | 97.8%
1.03 | 96.5% | 98%
1.04 | 96.7% | 98.1%
1.05 | 96.9% | 98.3%
1.06 | 97.1% | 98.4%
1.07 | 97.2% | 98.6%
1.08 | 97.4% | 98.7%
1.09 | 97.6% | 98.9%
1.1 | 97.7% | 99%
1.11 | 97.9% | 99.2%
1.12 | 98% | 99.3%
1.13 | 98.2% | 99.4%
1.14 | 98.3% | 99.5%
1.15 | 98.5% | 99.5%
1.16 | 98.6% | 99.6%
1.17 | 98.7% | 99.6%
1.18 | 98.8% | 99.7%
1.19 | 98.9% | 99.7%
1.2 | 99% | 99.8%
1.21 | 99.1% | 99.8%
1.22 | 99.2% | 99.9%
1.23 | 99.3% | 99.9%
1.24 | 99.4% | 100%
1.25 | 99.5% | 100%
1.26 | 99.6% | 100%
1.27 | 99.7% | 100%
1.28 | 99.7% | 99.9%
1.29 | 99.7% | 99.9%
1.3 | 99.8% | 99.8%
1.31 | 99.8% | 99.8%
1.32 | 99.9% | 99.7%
1.33 | 99.9% | 99.7%
1.34 | 99.9% | 99.6%
1.35 | 100% | 99.6%
1.36 | 100% | 99.5%
1.37 | 100% | 99.5%
1.38 | 100% | 99.4%
1.39 | 100% | 99.4%
1.4 | 100% | 99.3%
1.41 | 99.9% | 99.2%
1.42 | 99.8% | 99.1%
1.43 | 99.8% | 99%
1.44 | 99.7% | 98.9%
1.45 | 99.7% | 98.8%
1.46 | 99.6% | 98.7%
1.47 | 99.5% | 98.6%
1.48 | 99.4% | 98.4%
1.49 | 99.3% | 98.3%
1.5 | 99.2% | 98.1%
1.51 | 99% | 97.9%
1.52 | 98.9% | 97.7%
1.53 | 98.7% | 97.5%
1.54 | 98.5% | 97.3%
1.55 | 98.4% | 97%
1.56 | 98.2% | 96.8%
1.57 | 98% | 96.6%
1.58 | 97.8% | 96.4%
1.59 | 97.6% | 96.2%
1.6 | 97.4% | 95.9%
1.61 | 97.2% | 95.7%
1.62 | 96.9% | 95.4%
1.63 | 96.7% | 95.2%
1.64 | 96.4% | 94.9%
1.65 | 96.1% | 94.6%
1.66 | 95.9% | 94.3%
1.67 | 95.6% | 93.9%
1.68 | 95.2% | 93.6%
1.69 | 94.9% | 93.3%
1.7 | 94.6% | 93%
1.71 | 94.3% | 92.7%
1.72 | 94% | 92.3%
1.73 | 93.6% | 92%
1.74 | 93.2% | 91.6%
1.75 | 92.9% | 91.2%
1.76 | 92.5% | 90.7%
1.77 | 92.1% | 90.3%
1.78 | 91.7% | 89.9%
1.79 | 91.3% | 89.4%
1.8 | 90.8% | 89%
1.81 | 90.4% | 88.6%
1.82 | 90% | 88.2%
1.83 | 89.6% | 87.7%
1.84 | 89.1% | 87.3%
1.85 | 88.7% | 86.8%
1.86 | 88.2% | 86.3%
1.87 | 87.7% | 85.8%
1.88 | 87.2% | 85.3%
1.89 | 86.7% | 84.8%
1.9 | 86.2% | 84.3%
1.91 | 85.7% | 83.8%
1.92 | 85.1% | 83.3%
1.93 | 84.6% | 82.8%
1.94 | 84% | 82.3%
1.95 | 83.5% | 81.8%
1.96 | 82.9% | 81.3%
1.97 | 82.3% | 80.7%
1.98 | 81.8% | 80.2%
1.99 | 81.2% | 79.6%
2 | 80.5% | 79%
2.01 | 79.9% | 78.3%
2.02 | 79.3% | 77.7%
2.03 | 78.6% | 77.1%
2.04 | 78% | 76.5%
2.05 | 77.3% | 75.8%
2.06 | 76.7% | 75.2%
2.07 | 76% | 74.6%
2.08 | 75.3% | 73.9%
2.09 | 74.7% | 73.2%
2.1 | 74% | 72.5%
2.11 | 73.3% | 71.8%
2.12 | 72.6% | 71.1%
2.13 | 71.9% | 70.4%
2.14 | 71.1% | 69.8%
2.15 | 70.4% | 69.1%
2.16 | 69.7% | 68.4%
2.17 | 68.9% | 67.7%
2.18 | 68.2% | 67%
2.19 | 67.4% | 66.3%
2.2 | 66.7% | 65.6%
2.21 | 65.9% | 64.9%
2.22 | 65.2% | 64.2%
2.23 | 64.4% | 63.5%
2.24 | 63.6% | 62.8%
2.25 | 62.9% | 62.1%
2.26 | 62.1% | 61.4%
2.27 | 61.3% | 60.7%
2.28 | 60.5% | 60%
2.29 | 59.7% | 59.3%
2.3 | 58.9% | 58.6%
2.31 | 58.2% | 57.9%
2.32 | 57.4% | 57.2%
2.33 | 56.6% | 56.4%
2.34 | 55.8% | 55.7%
2.35 | 55% | 54.9%
2.36 | 54.1% | 54.1%
2.37 | 53.3% | 53.3%
2.38 | 52.5% | 52.5%
2.39 | 51.6% | 51.7%
2.4 | 50.8% | 50.9%
2.41 | 49.9% | 50.1%
2.42 | 49.1% | 49.3%
2.43 | 48.3% | 48.5%
2.44 | 47.4% | 47.7%
2.45 | 46.5% | 46.9%
2.46 | 45.7% | 46%
2.47 | 44.8% | 45.2%
2.48 | 43.9% | 44.4%
2.49 | 43% | 43.6%
2.5 | 42.2% | 42.8%
2.51 | 41.3% | 42%
2.52 | 40.4% | 41.2%
2.53 | 39.5% | 40.3%
2.54 | 38.7% | 39.5%
2.55 | 37.8% | 38.7%
2.56 | 36.9% | 37.9%
2.57 | 36.1% | 37.1%
2.58 | 35.2% | 36.3%
2.59 | 34.4% | 35.5%
2.6 | 33.6% | 34.7%
2.61 | 32.8% | 34%
2.62 | 32.1% | 33.2%
2.63 | 31.3% | 32.4%
2.64 | 30.5% | 31.7%
2.65 | 29.8% | 30.9%
2.66 | 29% | 30.1%
2.67 | 28.3% | 29.3%
2.68 | 27.5% | 28.6%
2.69 | 26.8% | 27.9%
2.7 | 26.1% | 27.2%
2.71 | 25.3% | 26.5%
2.72 | 24.6% | 25.8%
2.73 | 24% | 25.1%
2.74 | 23.3% | 24.4%
2.75 | 22.6% | 23.7%
2.76 | 21.9% | 23.1%
2.77 | 21.3% | 22.4%
2.78 | 20.6% | 21.7%
2.79 | 19.9% | 21%
2.8 | 19.3% | 20.3%
2.81 | 18.6% | 19.7%
2.82 | 18% | 19%
2.83 | 17.4% | 18.4%
2.84 | 16.8% | 17.8%
2.85 | 16.3% | 17.3%
2.86 | 15.7% | 16.7%
2.87 | 15.2% | 16.2%
2.88 | 14.6% | 15.6%
2.89 | 14.1% | 15.1%
2.9 | 13.6% | 14.5%
2.91 | 13.1% | 14%
2.92 | 12.6% | 13.5%
2.93 | 12.1% | 13%
2.94 | 11.6% | 12.6%
2.95 | 11.1% | 12.2%
2.96 | 10.7% | 11.7%
2.97 | 10.3% | 11.3%
2.98 | 9.9% | 10.9%
2.99 | 9.5% | 10.5%
3 | 9.1% | 10.1%
3.01 | 8.7% | 9.7%
3.02 | 8.4% | 9.2%
3.03 | 8% | 8.8%
3.04 | 7.6% | 8.4%
3.05 | 7.3% | 8%
3.06 | 6.9% | 7.6%
3.07 | 6.6% | 7.3%
3.08 | 6.3% | 6.9%
3.09 | 6.1% | 6.6%
3.1 | 5.8% | 6.3%
3.11 | 5.5% | 6%
3.12 | 5.3% | 5.7%
3.13 | 5% | 5.4%
3.14 | 4.8% | 5.1%
3.15 | 4.6% | 4.9%
3.16 | 4.3% | 4.7%
3.17 | 4.1% | 4.5%
3.18 | 3.9% | 4.3%
3.19 | 3.7% | 4.1%
3.2 | 3.5% | 3.9%
3.21 | 3.4% | 3.8%
3.22 | 3.2% | 3.6%
3.23 | 3.1% | 3.4%
3.24 | 3% | 3.3%
3.25 | 2.9% | 3.1%
3.26 | 2.7% | 3%
3.27 | 2.6% | 2.8%
3.28 | 2.5% | 2.7%
3.29 | 2.4% | 2.6%
3.3 | 2.3% | 2.5%
3.31 | 2.2% | 2.4%
3.32 | 2.1% | 2.3%
3.33 | 2% | 2.2%
3.34 | 1.9% | 2.1%
3.35 | 1.9% | 2%
3.36 | 1.8% | 1.9%
3.37 | 1.8% | 1.8%
3.38 | 1.7% | 1.7%
3.39 | 1.6% | 1.7%
3.4 | 1.6% | 1.6%
3.41 | 1.5% | 1.6%
3.42 | 1.5% | 1.5%
3.43 | 1.4% | 1.4%
3.44 | 1.4% | 1.4%
3.45 | 1.4% | 1.4%
3.46 | 1.3% | 1.3%
3.47 | 1.3% | 1.3%
3.48 | 1.3% | 1.3%
3.49 | 1.3% | 1.2%
3.5 | 1.2% | 1.2%
3.51 | 1.2% | 1.2%
3.52 | 1.2% | 1.1%
3.53 | 1.2% | 1.1%
3.54 | 1.2% | 1.1%
3.55 | 1.1% | 1.1%
3.56 | 1.1% | 1.1%
3.57 | 1.1% | 1.1%
3.58 | 1.1% | 1.1%
3.59 | 1.1% | 1%
3.6 | 1.1% | 1%
3.61 | 1.1% | 1%
3.62 | 1.1% | 1%
3.63 | 1.1% | 1%
3.64 | 1.1% | 1%
3.65 | 1.1% | 1%
3.66 | 1.1% | 1%
3.67 | 1.1% | 1%
3.68 | 1.1% | 1%
3.69 | 1.1% | 1%
3.7 | 1% | 1%
3.71 | 1% | 1%
3.72 | 1% | 1%
3.73 | 1% | 1%
3.74 | 1% | 1%
3.75 | 1% | 0.9%
3.76 | 1% | 0.9%
3.77 | 1% | 0.9%
3.78 | 1% | 0.9%
3.79 | 1% | 0.9%
3.8 | 1% | 0.9%
3.81 | 1% | 0.9%
3.82 | 1% | 0.9%
3.83 | 1% | 0.9%
3.84 | 1% | 0.9%
3.85 | 1% | 0.9%
3.86 | 1% | 0.9%
3.87 | 1% | 0.9%
3.88 | 1% | 0.9%
3.89 | 1% | 0.9%
3.9 | 1% | 0.9%
3.91 | 1% | 0.9%
3.92 | 1% | 0.9%
3.93 | 1% | 0.9%
3.94 | 1% | 0.9%
3.95 | 1% | 0.9%
3.96 | 1% | 0.9%
3.97 | 1% | 0.9%
3.98 | 1% | 0.9%
3.99 | 1% | 0.9%
4 | 1% | 0.9%
4.01 | 1% | 0.9%
4.02 | 1% | 0.9%
4.03 | 1% | 0.9%
4.04 | 1% | 0.9%
4.05 | 1% | 0.9%
4.06 | 1% | 0.9%
4.07 | 1% | 0.9%
4.08 | 1% | 0.9%
4.09 | 1% | 0.9%
4.1 | 1% | 0.9%
4.11 | 1% | 0.9%
4.12 | 1% | 0.9%
4.13 | 1% | 0.9%
4.14 | 1% | 0.9%
4.15 | 1% | 0.9%
4.16 | 1% | 0.9%
4.17 | 1% | 0.9%
4.18 | 1% | 0.9%
4.19 | 1% | 0.9%
4.2 | 1% | 0.9%
4.21 | 1% | 0.9%
4.22 | 1% | 0.9%
4.23 | 1% | 0.9%
4.24 | 1% | 0.9%
4.25 | 1% | 0.9%
4.26 | 1% | 0.9%
4.27 | 1% | 0.9%
4.28 | 1% | 0.9%
4.29 | 1% | 0.9%
4.3 | 1% | 0.9%
4.31 | 1% | 0.9%
4.32 | 1% | 0.9%
4.33 | 1% | 0.9%
4.34 | 1% | 0.9%
4.35 | 1% | 0.9%
4.36 | 1% | 0.9%
4.37 | 1% | 0.9%
4.38 | 1% | 0.9%
4.39 | 1% | 0.9%
4.4 | 1% | 0.9%
4.41 | 1% | 0.9%
4.42 | 1% | 0.9%
4.43 | 1% | 0.9%
4.44 | 1% | 0.9%
4.45 | 1% | 0.9%
4.46 | 1% | 0.9%
4.47 | 1% | 0.9%
4.48 | 1% | 0.9%
4.49 | 1% | 0.9%
4.5 | 1% | 0.9%
4.51 | 0.9% | 0.9%
4.52 | 0.9% | 0.9%
4.53 | 0.9% | 0.9%
4.54 | 0.9% | 0.9%
4.55 | 0.9% | 0.9%
4.56 | 0.9% | 0.9%
4.57 | 0.9% | 0.9%
4.58 | 0.9% | 0.9%
4.59 | 0.9% | 0.9%
4.6 | 0.9% | 0.9%
4.61 | 0.9% | 0.8%
4.62 | 0.9% | 0.8%
4.63 | 0.9% | 0.8%
4.64 | 0.9% | 0.8%
4.65 | 0.9% | 0.8%
4.66 | 0.9% | 0.8%
4.67 | 0.9% | 0.8%
4.68 | 0.9% | 0.8%
4.69 | 0.9% | 0.8%
4.7 | 0.9% | 0.8%
4.71 | 0.9% | 0.8%
4.72 | 0.9% | 0.8%
4.73 | 0.9% | 0.8%
4.74 | 0.9% | 0.8%
4.75 | 0.9% | 0.8%
4.76 | 0.9% | 0.8%
4.77 | 0.9% | 0.8%
4.78 | 0.9% | 0.8%
4.79 | 0.9% | 0.8%
4.8 | 0.9% | 0.8%
4.81 | 0.9% | 0.8%
4.82 | 0.9% | 0.8%
4.83 | 0.9% | 0.8%
4.84 | 0.9% | 0.8%
4.85 | 0.9% | 0.8%
4.86 | 0.9% | 0.8%
4.87 | 0.9% | 0.8%
4.88 | 0.9% | 0.8%
4.89 | 0.9% | 0.8%
4.9 | 0.9% | 0.8%
4.91 | 0.9% | 0.8%
4.92 | 0.9% | 0.8%
4.93 | 0.9% | 0.8%
4.94 | 0.9% | 0.8%
4.95 | 0.9% | 0.8%
4.96 | 0.9% | 0.8%
4.97 | 0.9% | 0.8%
4.98 | 0.9% | 0.8%
4.99 | 0.9% | 0.8%
5 | 0.9% | 0.8%
5.01 | 0.9% | 0.8%
5.02 | 0.9% | 0.8%
5.03 | 0.9% | 0.8%
5.04 | 0.9% | 0.8%
5.05 | 0.9% | 0.8%
5.06 | 0.9% | 0.8%
5.07 | 0.9% | 0.8%
5.08 | 0.9% | 0.8%
5.09 | 0.9% | 0.8%
5.1 | 0.9% | 0.8%
5.11 | 0.9% | 0.8%
5.12 | 0.9% | 0.8%
5.13 | 0.9% | 0.8%
5.14 | 0.9% | 0.8%
5.15 | 0.9% | 0.8%
5.16 | 0.9% | 0.8%
5.17 | 0.9% | 0.8%
5.18 | 0.9% | 0.8%
5.19 | 0.9% | 0.8%
5.2 | 0.9% | 0.8%
5.21 | 0.9% | 0.8%
5.22 | 0.9% | 0.8%
5.23 | 0.9% | 0.8%
5.24 | 0.9% | 0.8%
5.25 | 0.9% | 0.8%
5.26 | 0.9% | 0.8%
5.27 | 0.9% | 0.8%
5.28 | 0.9% | 0.8%
5.29 | 0.9% | 0.8%
5.3 | 0.9% | 0.8%
5.31 | 0.9% | 0.8%
5.32 | 0.9% | 0.8%
5.33 | 0.9% | 0.8%
5.34 | 0.9% | 0.8%
5.35 | 0.9% | 0.8%
5.36 | 0.9% | 0.8%
5.37 | 0.9% | 0.8%
5.38 | 0.9% | 0.8%
5.39 | 0.9% | 0.8%
5.4 | 0.9% | 0.8%
5.41 | 0.9% | 0.8%
5.42 | 0.9% | 0.8%
5.43 | 0.9% | 0.8%
5.44 | 0.9% | 0.8%
5.45 | 0.9% | 0.8%
5.46 | 0.9% | 0.8%
5.47 | 0.9% | 0.8%
5.48 | 0.9% | 0.8%
5.49 | 0.9% | 0.8%
5.5 | 0.9% | 0.8%
5.51 | 0.9% | 0.8%
5.52 | 0.9% | 0.8%
5.53 | 0.9% | 0.8%
5.54 | 0.9% | 0.8%
5.55 | 0.9% | 0.8%
5.56 | 0.9% | 0.8%
5.57 | 0.9% | 0.8%
5.58 | 0.9% | 0.8%
5.59 | 0.9% | 0.8%
5.6 | 0.9% | 0.8%
5.61 | 0.9% | 0.8%
5.62 | 0.9% | 0.8%
5.63 | 0.9% | 0.8%
5.64 | 0.8% | 0.8%
5.65 | 0.8% | 0.8%
5.66 | 0.8% | 0.8%
5.67 | 0.8% | 0.8%
5.68 | 0.8% | 0.8%
5.69 | 0.8% | 0.8%
5.7 | 0.8% | 0.8%
5.71 | 0.8% | 0.8%
5.72 | 0.8% | 0.8%
5.73 | 0.8% | 0.8%
5.74 | 0.8% | 0.8%
5.75 | 0.8% | 0.8%
5.76 | 0.8% | 0.8%
5.77 | 0.8% | 0.7%
5.78 | 0.8% | 0.7%
5.79 | 0.8% | 0.7%
5.8 | 0.8% | 0.7%
5.81 | 0.8% | 0.7%
5.82 | 0.8% | 0.7%
5.83 | 0.8% | 0.7%
5.84 | 0.8% | 0.7%
5.85 | 0.8% | 0.7%
5.86 | 0.8% | 0.7%
5.87 | 0.8% | 0.7%
5.88 | 0.8% | 0.7%
5.89 | 0.8% | 0.7%
5.9 | 0.8% | 0.7%
5.91 | 0.8% | 0.7%
5.92 | 0.8% | 0.7%
5.93 | 0.8% | 0.7%
5.94 | 0.8% | 0.7%
5.95 | 0.8% | 0.7%
5.96 | 0.8% | 0.7%
5.97 | 0.8% | 0.7%
5.98 | 0.8% | 0.7%
5.99 | 0.8% | 0.7%
6 | 0.8% | 0.7%
6.01 | 0.8% | 0.7%
6.02 | 0.8% | 0.7%
6.03 | 0.8% | 0.7%
6.04 | 0.8% | 0.7%
6.05 | 0.8% | 0.7%
6.06 | 0.8% | 0.7%
6.07 | 0.8% | 0.7%
6.08 | 0.8% | 0.7%
6.09 | 0.8% | 0.7%
6.1 | 0.8% | 0.7%
6.11 | 0.8% | 0.7%
6.12 | 0.8% | 0.7%
6.13 | 0.8% | 0.7%
6.14 | 0.8% | 0.7%
6.15 | 0.8% | 0.7%
6.16 | 0.8% | 0.7%
6.17 | 0.8% | 0.7%
6.18 | 0.8% | 0.7%
6.19 | 0.8% | 0.7%
6.2 | 0.8% | 0.7%
6.21 | 0.8% | 0.7%
6.22 | 0.8% | 0.7%
6.23 | 0.8% | 0.7%
6.24 | 0.8% | 0.7%
6.25 | 0.8% | 0.7%
6.26 | 0.8% | 0.7%
6.27 | 0.8% | 0.7%
6.28 | 0.8% | 0.7%
6.29 | 0.8% | 0.7%
6.3 | 0.8% | 0.7%
6.31 | 0.8% | 0.7%
6.32 | 0.8% | 0.7%
6.33 | 0.8% | 0.7%
6.34 | 0.8% | 0.7%
6.35 | 0.8% | 0.7%
6.36 | 0.8% | 0.7%
6.37 | 0.8% | 0.7%
6.38 | 0.8% | 0.7%
6.39 | 0.8% | 0.7%
6.4 | 0.8% | 0.7%
6.41 | 0.8% | 0.7%
6.42 | 0.8% | 0.7%
6.43 | 0.8% | 0.7%
6.44 | 0.8% | 0.7%
6.45 | 0.8% | 0.7%
6.46 | 0.8% | 0.7%
6.47 | 0.8% | 0.7%
6.48 | 0.8% | 0.7%
6.49 | 0.8% | 0.7%
6.5 | 0.8% | 0.7%
6.51 | 0.8% | 0.7%
6.52 | 0.8% | 0.7%
6.53 | 0.8% | 0.7%
6.54 | 0.8% | 0.7%
6.55 | 0.8% | 0.7%
6.56 | 0.8% | 0.7%
6.57 | 0.8% | 0.7%
6.58 | 0.8% | 0.7%
6.59 | 0.8% | 0.7%
6.6 | 0.8% | 0.7%
6.61 | 0.8% | 0.7%
6.62 | 0.8% | 0.7%
6.63 | 0.8% | 0.7%
6.64 | 0.8% | 0.7%
6.65 | 0.8% | 0.7%
6.66 | 0.8% | 0.7%
6.67 | 0.8% | 0.7%
6.68 | 0.8% | 0.7%
6.69 | 0.8% | 0.7%
6.7 | 0.8% | 0.7%
6.71 | 0.8% | 0.7%
6.72 | 0.8% | 0.7%
6.73 | 0.8% | 0.7%
6.74 | 0.8% | 0.7%
6.75 | 0.8% | 0.7%
6.76 | 0.8% | 0.7%
6.77 | 0.8% | 0.7%
6.78 | 0.8% | 0.7%
6.79 | 0.8% | 0.7%
6.8 | 0.8% | 0.7%
6.81 | 0.8% | 0.7%
6.82 | 0.8% | 0.7%
6.83 | 0.8% | 0.7%
6.84 | 0.8% | 0.7%
6.85 | 0.8% | 0.7%
6.86 | 0.8% | 0.7%
6.87 | 0.8% | 0.7%
6.88 | 0.8% | 0.7%
6.89 | 0.8% | 0.7%
6.9 | 0.8% | 0.7%
6.91 | 0.8% | 0.7%
6.92 | 0.8% | 0.7%
6.93 | 0.8% | 0.7%
6.94 | 0.8% | 0.7%
6.95 | 0.8% | 0.7%
6.96 | 0.8% | 0.7%
6.97 | 0.8% | 0.7%
6.98 | 0.8% | 0.7%
6.99 | 0.8% | 0.7%
7 | 0.8% | 0.7%
7.01 | 0.8% | 0.7%
7.02 | 0.8% | 0.7%
7.03 | 0.8% | 0.7%
7.04 | 0.8% | 0.7%
7.05 | 0.8% | 0.7%
7.06 | 0.8% | 0.7%
7.07 | 0.8% | 0.7%
7.08 | 0.8% | 0.7%
7.09 | 0.8% | 0.7%
7.1 | 0.8% | 0.7%
7.11 | 0.8% | 0.7%
7.12 | 0.8% | 0.7%
7.13 | 0.8% | 0.7%
7.14 | 0.8% | 0.7%
7.15 | 0.8% | 0.7%
7.16 | 0.8% | 0.7%
7.17 | 0.8% | 0.7%
7.18 | 0.8% | 0.7%
7.19 | 0.8% | 0.7%
7.2 | 0.8% | 0.7%
7.21 | 0.8% | 0.7%
7.22 | 0.8% | 0.7%
7.23 | 0.8% | 0.7%
7.24 | 0.8% | 0.7%
7.25 | 0.8% | 0.7%
7.26 | 0.7% | 0.7%
7.27 | 0.7% | 0.7%
7.28 | 0.7% | 0.7%
7.29 | 0.7% | 0.6%
7.3 | 0.7% | 0.6%
7.31 | 0.7% | 0.6%
7.32 | 0.7% | 0.6%
7.33 | 0.7% | 0.6%
7.34 | 0.7% | 0.6%
7.35 | 0.7% | 0.6%
7.36 | 0.7% | 0.6%
7.37 | 0.7% | 0.6%
7.38 | 0.7% | 0.6%
7.39 | 0.7% | 0.6%
7.4 | 0.7% | 0.6%
7.41 | 0.7% | 0.6%
7.42 | 0.7% | 0.6%
7.43 | 0.7% | 0.6%
7.44 | 0.7% | 0.6%
7.45 | 0.7% | 0.6%
7.46 | 0.7% | 0.6%
7.47 | 0.7% | 0.6%
7.48 | 0.7% | 0.6%
7.49 | 0.7% | 0.6%
7.5 | 0.7% | 0.6%
7.51 | 0.7% | 0.6%
7.52 | 0.7% | 0.6%
7.53 | 0.7% | 0.6%
7.54 | 0.7% | 0.6%
7.55 | 0.7% | 0.6%
7.56 | 0.7% | 0.6%
7.57 | 0.7% | 0.6%
7.58 | 0.7% | 0.6%
7.59 | 0.7% | 0.6%
7.6 | 0.7% | 0.6%
7.61 | 0.7% | 0.6%
7.62 | 0.7% | 0.6%
7.63 | 0.7% | 0.6%
7.64 | 0.7% | 0.6%
7.65 | 0.7% | 0.6%
7.66 | 0.7% | 0.6%
7.67 | 0.7% | 0.6%
7.68 | 0.7% | 0.6%
7.69 | 0.7% | 0.6%
7.7 | 0.7% | 0.6%
7.71 | 0.7% | 0.6%
7.72 | 0.7% | 0.6%
7.73 | 0.7% | 0.6%
7.74 | 0.7% | 0.6%
7.75 | 0.7% | 0.6%
7.76 | 0.7% | 0.6%
7.77 | 0.7% | 0.6%
7.78 | 0.7% | 0.6%
7.79 | 0.7% | 0.6%
7.8 | 0.7% | 0.6%
0.6%
0.6%
0.6%
0.6%
0.6%
0.6%
0.6%
0.6%
0.6%
0.6%

## 9 MeV PDDs
Energy: | 9 MeV
SSD: | 100 cm
| PDD | PDD | PDD
Depth | 6x6+ | 4x4 | 3x3 | Energy | SSD | Field Size | R100 | R50 | R95 | R90 | R80 | Rp | Ds | PDD(2.13) | PDD(2.1)
0 |  | 83.97% |  | 9 MeV | 100.00 cm | 3x3 | 1.62 cm | 3.49 cm | 2.25 cm | 2.48 cm | 2.80 cm | 4.46 cm | 87.7% | 96.8% | 97.3%
0.01 |  | 83.97% | 87.4% | 9 MeV | 100.00 cm | 4x4 | 1.97 cm | 3.61 cm | 2.52 cm | 2.73 cm | 3.02 cm | 4.43 cm | 83.8% | 99.7% | 99.8%
0.02 | 81.78% | 83.97% | 87.5% | 9 MeV | 100.00 cm | 6x6+ | 2.13 cm | 3.63 cm | 2.64 cm | 2.83 cm | 3.10 cm | 4.41 cm | 81.9% | 100% | 100%
0.03 | 81.78% | 83.97% | 87.5%
0.04 | 81.88% | 83.97% | 87.6%
0.05 | 81.98% | 83.97% | 87.7%
0.06 | 81.98% | 83.97% | 87.7%
0.07 | 82.08% | 83.97% | 87.8%
0.08 | 82.08% | 84.07% | 87.8%
0.09 | 82.18% | 84.17% | 87.9%
0.1 | 82.18% | 84.17% | 88%
0.11 | 82.28% | 84.27% | 88%
0.12 | 82.38% | 84.27% | 88.1%
0.13 | 82.48% | 84.37% | 88.1%
0.14 | 82.48% | 84.47% | 88.2%
0.15 | 82.58% | 84.47% | 88.3%
0.16 | 82.68% | 84.57% | 88.3%
0.17 | 82.78% | 84.67% | 88.4%
0.18 | 82.78% | 84.77% | 88.5%
0.19 | 82.88% | 84.87% | 88.6%
0.2 | 82.98% | 84.97% | 88.7%
0.21 | 83.08% | 85.07% | 88.8%
0.22 | 83.18% | 85.17% | 89%
0.23 | 83.28% | 85.27% | 89.1%
0.24 | 83.38% | 85.37% | 89.2%
0.25 | 83.48% | 85.57% | 89.3%
0.26 | 83.58% | 85.67% | 89.5%
0.27 | 83.78% | 85.77% | 89.6%
0.28 | 83.88% | 85.87% | 89.7%
0.29 | 83.98% | 85.97% | 89.9%
0.3 | 84.18% | 86.17% | 90%
0.31 | 84.28% | 86.27% | 90.1%
0.32 | 84.38% | 86.37% | 90.3%
0.33 | 84.58% | 86.47% | 90.4%
0.34 | 84.68% | 86.67% | 90.5%
0.35 | 84.78% | 86.77% | 90.7%
0.36 | 84.88% | 86.87% | 90.8%
0.37 | 85.09% | 86.97% | 90.9%
0.38 | 85.19% | 87.17% | 91.1%
0.39 | 85.29% | 87.27% | 91.2%
0.4 | 85.39% | 87.37% | 91.3%
0.41 | 85.49% | 87.47% | 91.5%
0.42 | 85.59% | 87.68% | 91.6%
0.43 | 85.69% | 87.78% | 91.7%
0.44 | 85.79% | 87.88% | 91.8%
0.45 | 85.89% | 87.98% | 91.9%
0.46 | 85.99% | 88.18% | 92%
0.47 | 86.19% | 88.28% | 92.1%
0.48 | 86.29% | 88.38% | 92.2%
0.49 | 86.39% | 88.48% | 92.3%
0.5 | 86.49% | 88.58% | 92.4%
0.51 | 86.59% | 88.78% | 92.5%
0.52 | 86.69% | 88.88% | 92.6%
0.53 | 86.79% | 88.98% | 92.7%
0.54 | 86.89% | 89.08% | 92.8%
0.55 | 86.99% | 89.18% | 92.9%
0.56 | 87.09% | 89.38% | 93%
0.57 | 87.09% | 89.48% | 93.1%
0.58 | 87.19% | 89.58% | 93.2%
0.59 | 87.29% | 89.68% | 93.3%
0.6 | 87.39% | 89.78% | 93.4%
0.61 | 87.49% | 89.88% | 93.5%
0.62 | 87.59% | 89.98% | 93.6%
0.63 | 87.69% | 90.08% | 93.7%
0.64 | 87.79% | 90.28% | 93.8%
0.65 | 87.89% | 90.38% | 93.9%
0.66 | 87.99% | 90.48% | 94%
0.67 | 88.09% | 90.58% | 94.1%
0.68 | 88.19% | 90.68% | 94.2%
0.69 | 88.29% | 90.78% | 94.3%
0.7 | 88.39% | 90.88% | 94.3%
0.71 | 88.49% | 90.98% | 94.4%
0.72 | 88.59% | 91.08% | 94.5%
0.73 | 88.69% | 91.08% | 94.6%
0.74 | 88.79% | 91.18% | 94.7%
0.75 | 88.89% | 91.28% | 94.8%
0.76 | 88.99% | 91.38% | 94.9%
0.77 | 89.09% | 91.48% | 95%
0.78 | 89.19% | 91.58% | 95.1%
0.79 | 89.29% | 91.68% | 95.1%
0.8 | 89.39% | 91.78% | 95.2%
0.81 | 89.49% | 91.88% | 95.3%
0.82 | 89.59% | 91.98% | 95.4%
0.83 | 89.69% | 92.08% | 95.4%
0.84 | 89.79% | 92.18% | 95.5%
0.85 | 89.89% | 92.18% | 95.6%
0.86 | 89.99% | 92.28% | 95.6%
0.87 | 90.09% | 92.38% | 95.7%
0.88 | 90.19% | 92.48% | 95.8%
0.89 | 90.29% | 92.59% | 95.9%
0.9 | 90.39% | 92.69% | 95.9%
0.91 | 90.49% | 92.79% | 96%
0.92 | 90.59% | 92.89% | 96.1%
0.93 | 90.69% | 92.99% | 96.2%
0.94 | 90.79% | 93.09% | 96.3%
0.95 | 90.89% | 93.19% | 96.4%
0.96 | 90.99% | 93.29% | 96.5%
0.97 | 91.09% | 93.29% | 96.6%
0.98 | 91.19% | 93.39% | 96.7%
0.99 | 91.29% | 93.49% | 96.8%
1 | 91.29% | 93.59% | 96.8%
1.01 | 91.39% | 93.69% | 96.9%
1.02 | 91.49% | 93.79% | 97%
1.03 | 91.59% | 93.89% | 97.1%
1.04 | 91.69% | 93.99% | 97.2%
1.05 | 91.79% | 94.09% | 97.2%
1.06 | 91.89% | 94.19% | 97.3%
1.07 | 91.99% | 94.29% | 97.4%
1.08 | 92.09% | 94.39% | 97.5%
1.09 | 92.19% | 94.39% | 97.5%
1.1 | 92.19% | 94.49% | 97.6%
1.11 | 92.29% | 94.59% | 97.7%
1.12 | 92.39% | 94.69% | 97.8%
1.13 | 92.49% | 94.79% | 97.8%
1.14 | 92.59% | 94.89% | 97.9%
1.15 | 92.69% | 94.99% | 97.9%
1.16 | 92.79% | 95.09% | 98%
1.17 | 92.89% | 95.19% | 98.1%
1.18 | 92.99% | 95.19% | 98.1%
1.19 | 93.09% | 95.29% | 98.2%
1.2 | 93.19% | 95.39% | 98.2%
1.21 | 93.29% | 95.49% | 98.3%
1.22 | 93.39% | 95.59% | 98.3%
1.23 | 93.49% | 95.69% | 98.4%
1.24 | 93.59% | 95.69% | 98.4%
1.25 | 93.69% | 95.79% | 98.5%
1.26 | 93.79% | 95.89% | 98.6%
1.27 | 93.89% | 95.99% | 98.6%
1.28 | 93.99% | 96.09% | 98.7%
1.29 | 94.09% | 96.19% | 98.7%
1.3 | 94.19% | 96.19% | 98.8%
1.31 | 94.29% | 96.29% | 98.9%
1.32 | 94.39% | 96.39% | 98.9%
1.33 | 94.39% | 96.49% | 99%
1.34 | 94.49% | 96.59% | 99.1%
1.35 | 94.59% | 96.69% | 99.1%
1.36 | 94.69% | 96.69% | 99.2%
1.37 | 94.79% | 96.79% | 99.2%
1.38 | 94.89% | 96.89% | 99.3%
1.39 | 94.99% | 96.99% | 99.3%
1.4 | 95.1% | 97.09% | 99.3%
1.41 | 95.2% | 97.09% | 99.4%
1.42 | 95.3% | 97.19% | 99.4%
1.43 | 95.4% | 97.29% | 99.4%
1.44 | 95.5% | 97.39% | 99.5%
1.45 | 95.6% | 97.39% | 99.5%
1.46 | 95.7% | 97.49% | 99.5%
1.47 | 95.8% | 97.6% | 99.6%
1.48 | 95.9% | 97.7% | 99.6%
1.49 | 96% | 97.7% | 99.6%
1.5 | 96.1% | 97.8% | 99.7%
1.51 | 96.2% | 97.9% | 99.7%
1.52 | 96.3% | 98% | 99.7%
1.53 | 96.5% | 98% | 99.7%
1.54 | 96.5% | 98.1% | 99.8%
1.55 | 96.6% | 98.2% | 99.8%
1.56 | 96.7% | 98.2% | 99.8%
1.57 | 96.8% | 98.3% | 99.9%
1.58 | 96.9% | 98.4% | 99.9%
1.59 | 97% | 98.5% | 99.9%
1.6 | 97.1% | 98.5% | 99.9%
1.61 | 97.2% | 98.6% | 100%
1.62 | 97.2% | 98.7% | 100%
1.63 | 97.3% | 98.7% | 100%
1.64 | 97.4% | 98.8% | 100%
1.65 | 97.5% | 98.9% | 100%
1.66 | 97.6% | 98.9% | 99.9%
1.67 | 97.7% | 99% | 99.9%
1.68 | 97.7% | 99% | 99.9%
1.69 | 97.8% | 99.1% | 99.9%
1.7 | 97.9% | 99.2% | 99.9%
1.71 | 98% | 99.2% | 99.9%
1.72 | 98.1% | 99.3% | 99.9%
1.73 | 98.1% | 99.3% | 99.8%
1.74 | 98.2% | 99.4% | 99.8%
1.75 | 98.3% | 99.4% | 99.8%
1.76 | 98.4% | 99.5% | 99.8%
1.77 | 98.4% | 99.5% | 99.7%
1.78 | 98.5% | 99.5% | 99.7%
1.79 | 98.6% | 99.6% | 99.7%
1.8 | 98.7% | 99.6% | 99.6%
1.81 | 98.8% | 99.7% | 99.6%
1.82 | 98.8% | 99.7% | 99.5%
1.83 | 98.9% | 99.7% | 99.5%
1.84 | 99% | 99.8% | 99.5%
1.85 | 99% | 99.8% | 99.4%
1.86 | 99.1% | 99.8% | 99.4%
1.87 | 99.2% | 99.9% | 99.3%
1.88 | 99.2% | 99.9% | 99.2%
1.89 | 99.3% | 99.9% | 99.2%
1.9 | 99.3% | 99.9% | 99.1%
1.91 | 99.4% | 100% | 99%
1.92 | 99.4% | 100% | 98.9%
1.93 | 99.5% | 100% | 98.8%
1.94 | 99.5% | 100% | 98.8%
1.95 | 99.6% | 100% | 98.7%
1.96 | 99.6% | 100% | 98.6%
1.97 | 99.7% | 100% | 98.5%
1.98 | 99.7% | 100% | 98.4%
1.99 | 99.8% | 100% | 98.4%
2 | 99.8% | 100% | 98.3%
2.01 | 99.8% | 100% | 98.2%
2.02 | 99.9% | 100% | 98.1%
2.03 | 99.9% | 100% | 98%
2.04 | 99.9% | 100% | 97.9%
2.05 | 100% | 100% | 97.8%
2.06 | 100% | 100% | 97.7%
2.07 | 100% | 99.9% | 97.6%
2.08 | 100% | 99.9% | 97.5%
2.09 | 100% | 99.9% | 97.4%
2.1 | 100% | 99.8% | 97.3%
2.11 | 100% | 99.8% | 97.1%
2.12 | 100% | 99.8% | 97%
2.13 | 100% | 99.7% | 96.8%
2.14 | 100% | 99.7% | 96.7%
2.15 | 100% | 99.6% | 96.5%
2.16 | 100% | 99.6% | 96.4%
2.17 | 100% | 99.5% | 96.3%
2.18 | 100% | 99.5% | 96.1%
2.19 | 100% | 99.4% | 96%
2.2 | 100% | 99.3% | 95.8%
2.21 | 99.9% | 99.3% | 95.7%
2.22 | 99.9% | 99.2% | 95.5%
2.23 | 99.9% | 99.1% | 95.4%
2.24 | 99.8% | 99% | 95.2%
2.25 | 99.8% | 99% | 95%
2.26 | 99.7% | 98.9% | 94.8%
2.27 | 99.7% | 98.8% | 94.6%
2.28 | 99.7% | 98.7% | 94.4%
2.29 | 99.6% | 98.6% | 94.2%
2.3 | 99.5% | 98.5% | 94%
2.31 | 99.5% | 98.4% | 93.8%
2.32 | 99.4% | 98.2% | 93.6%
2.33 | 99.3% | 98.1% | 93.4%
2.34 | 99.3% | 98% | 93.2%
2.35 | 99.2% | 97.9% | 92.9%
2.36 | 99.1% | 97.8% | 92.7%
2.37 | 99% | 97.6% | 92.5%
2.38 | 98.9% | 97.49% | 92.3%
2.39 | 98.9% | 97.29% | 92%
2.4 | 98.8% | 97.19% | 91.8%
2.41 | 98.7% | 96.99% | 91.6%
2.42 | 98.6% | 96.89% | 91.4%
2.43 | 98.5% | 96.69% | 91.1%
2.44 | 98.4% | 96.49% | 90.9%
2.45 | 98.3% | 96.39% | 90.7%
2.46 | 98.1% | 96.19% | 90.4%
2.47 | 98% | 95.99% | 90.2%
2.48 | 97.9% | 95.79% | 89.9%
2.49 | 97.8% | 95.59% | 89.6%
2.5 | 97.6% | 95.39% | 89.4%
2.51 | 97.5% | 95.19% | 89.1%
2.52 | 97.4% | 94.99% | 88.8%
2.53 | 97.2% | 94.79% | 88.6%
2.54 | 97% | 94.59% | 88.3%
2.55 | 96.9% | 94.39% | 88%
2.56 | 96.7% | 94.19% | 87.8%
2.57 | 96.5% | 93.89% | 87.5%
2.58 | 96.3% | 93.69% | 87.2%
2.59 | 96.1% | 93.49% | 86.9%
2.6 | 95.9% | 93.29% | 86.6%
2.61 | 95.7% | 92.99% | 86.3%
2.62 | 95.5% | 92.79% | 86%
2.63 | 95.3% | 92.48% | 85.7%
2.64 | 95.1% | 92.28% | 85.4%
2.65 | 94.89% | 92.08% | 85.1%
2.66 | 94.69% | 91.78% | 84.8%
2.67 | 94.49% | 91.48% | 84.5%
2.68 | 94.29% | 91.28% | 84.1%
2.69 | 93.99% | 90.98% | 83.8%
2.7 | 93.79% | 90.68% | 83.5%
2.71 | 93.49% | 90.48% | 83.2%
2.72 | 93.29% | 90.18% | 82.8%
2.73 | 92.99% | 89.88% | 82.5%
2.74 | 92.79% | 89.58% | 82.1%
2.75 | 92.49% | 89.28% | 81.8%
2.76 | 92.19% | 89.08% | 81.4%
2.77 | 91.89% | 88.78% | 81%
2.78 | 91.69% | 88.38% | 80.7%
2.79 | 91.39% | 88.08% | 80.3%
2.8 | 91.09% | 87.78% | 79.9%
2.81 | 90.79% | 87.47% | 79.6%
2.82 | 90.49% | 87.17% | 79.2%
2.83 | 90.19% | 86.87% | 78.9%
2.84 | 89.89% | 86.57% | 78.5%
2.85 | 89.49% | 86.17% | 78.1%
2.86 | 89.19% | 85.87% | 77.8%
2.87 | 88.89% | 85.57% | 77.4%
2.88 | 88.59% | 85.17% | 77%
2.89 | 88.29% | 84.87% | 76.7%
2.9 | 87.99% | 84.57% | 76.3%
2.91 | 87.59% | 84.17% | 75.9%
2.92 | 87.29% | 83.87% | 75.6%
2.93 | 86.89% | 83.47% | 75.2%
2.94 | 86.49% | 83.17% | 74.8%
2.95 | 86.09% | 82.77% | 74.5%
2.96 | 85.79% | 82.46% | 74.1%
2.97 | 85.39% | 82.06% | 73.6%
2.98 | 84.98% | 81.66% | 73.2%
2.99 | 84.58% | 81.36% | 72.8%
3 | 84.18% | 80.96% | 72.4%
3.01 | 83.78% | 80.56% | 72%
3.02 | 83.38% | 80.16% | 71.6%
3.03 | 82.98% | 79.76% | 71.2%
3.04 | 82.48% | 79.36% | 70.8%
3.05 | 82.08% | 78.96% | 70.4%
3.06 | 81.68% | 78.56% | 70%
3.07 | 81.18% | 78.16% | 69.6%
3.08 | 80.78% | 77.76% | 69.2%
3.09 | 80.28% | 77.35% | 68.8%
3.1 | 79.88% | 76.95% | 68.3%
3.11 | 79.38% | 76.45% | 67.9%
3.12 | 78.98% | 76.05% | 67.5%
3.13 | 78.58% | 75.55% | 67.1%
3.14 | 78.08% | 75.15% | 66.7%
3.15 | 77.68% | 74.75% | 66.2%
3.16 | 77.18% | 74.25% | 65.8%
3.17 | 76.68% | 73.75% | 65.4%
3.18 | 76.28% | 73.35% | 65%
3.19 | 75.78% | 72.85% | 64.6%
3.2 | 75.38% | 72.34% | 64.1%
3.21 | 74.87% | 71.94% | 63.7%
3.22 | 74.37% | 71.44% | 63.3%
3.23 | 73.97% | 70.94% | 62.8%
3.24 | 73.47% | 70.44% | 62.4%
3.25 | 72.97% | 69.94% | 62%
3.26 | 72.47% | 69.44% | 61.5%
3.27 | 71.97% | 68.94% | 61.1%
3.28 | 71.47% | 68.44% | 60.6%
3.29 | 70.97% | 67.94% | 60.2%
3.3 | 70.37% | 67.33% | 59.7%
3.31 | 69.77% | 66.83% | 59.3%
3.32 | 69.17% | 66.33% | 58.8%
3.33 | 68.67% | 65.83% | 58.3%
3.34 | 68.07% | 65.23% | 57.8%
3.35 | 67.47% | 64.73% | 57.3%
3.36 | 66.87% | 64.23% | 56.8%
3.37 | 66.27% | 63.63% | 56.3%
3.38 | 65.67% | 63.13% | 55.8%
3.39 | 65.07% | 62.53% | 55.3%
3.4 | 64.46% | 62.02% | 54.8%
3.41 | 63.86% | 61.42% | 54.3%
3.42 | 63.26% | 60.92% | 53.8%
3.43 | 62.66% | 60.32% | 53.2%
3.44 | 62.16% | 59.82% | 52.7%
3.45 | 61.56% | 59.22% | 52.2%
3.46 | 60.96% | 58.62% | 51.7%
3.47 | 60.36% | 58.02% | 51.2%
3.48 | 59.76% | 57.52% | 50.7%
3.49 | 59.16% | 56.91% | 50.2%
3.5 | 58.46% | 56.31% | 49.6%
3.51 | 57.86% | 55.71% | 49.1%
3.52 | 57.26% | 55.11% | 48.6%
3.53 | 56.66% | 54.51% | 48.1%
3.54 | 55.96% | 53.91% | 47.6%
3.55 | 55.36% | 53.31% | 47.1%
3.56 | 54.75% | 52.81% | 46.6%
3.57 | 54.05% | 52.2% | 46.1%
3.58 | 53.45% | 51.5% | 45.6%
3.59 | 52.85% | 50.9% | 45.1%
3.6 | 52.15% | 50.3% | 44.6%
3.61 | 51.55% | 49.7% | 44%
3.62 | 50.95% | 49.1% | 43.5%
3.63 | 50.25% | 48.5% | 43%
3.64 | 49.65% | 47.9% | 42.5%
3.65 | 49.05% | 47.29% | 42%
3.66 | 48.35% | 46.69% | 41.5%
3.67 | 47.75% | 46.09% | 41%
3.68 | 47.15% | 45.49% | 40.5%
3.69 | 46.45% | 44.89% | 40%
3.7 | 45.85% | 44.29% | 39.4%
3.71 | 45.25% | 43.69% | 38.9%
3.72 | 44.64% | 43.09% | 38.4%
3.73 | 43.94% | 42.48% | 37.9%
3.74 | 43.34% | 41.98% | 37.4%
3.75 | 42.74% | 41.38% | 36.9%
3.76 | 42.14% | 40.78% | 36.4%
3.77 | 41.44% | 40.18% | 35.8%
3.78 | 40.84% | 39.58% | 35.3%
3.79 | 40.24% | 38.98% | 34.8%
3.8 | 39.54% | 38.38% | 34.2%
3.81 | 38.94% | 37.78% | 33.7%
3.82 | 38.34% | 37.27% | 33.2%
3.83 | 37.64% | 36.67% | 32.6%
3.84 | 37.04% | 36.07% | 32.1%
3.85 | 36.44% | 35.47% | 31.6%
3.86 | 35.84% | 34.87% | 31%
3.87 | 35.14% | 34.37% | 30.5%
3.88 | 34.53% | 33.77% | 30%
3.89 | 33.93% | 33.17% | 29.4%
3.9 | 33.33% | 32.57% | 28.9%
3.91 | 32.73% | 32.06% | 28.4%
3.92 | 32.13% | 31.46% | 27.9%
3.93 | 31.53% | 30.96% | 27.4%
3.94 | 30.93% | 30.36% | 27%
3.95 | 30.33% | 29.76% | 26.5%
3.96 | 29.73% | 29.26% | 26%
3.97 | 29.13% | 28.66% | 25.5%
3.98 | 28.63% | 28.16% | 25.1%
3.99 | 28.03% | 27.56% | 24.6%
4 | 27.43% | 27.05% | 24.1%
4.01 | 26.93% | 26.55% | 23.7%
4.02 | 26.43% | 25.95% | 23.2%
4.03 | 25.83% | 25.45% | 22.7%
4.04 | 25.33% | 24.95% | 22.3%
4.05 | 24.82% | 24.45% | 21.8%
4.06 | 24.32% | 23.85% | 21.4%
4.07 | 23.72% | 23.35% | 21%
4.08 | 23.22% | 22.85% | 20.5%
4.09 | 22.72% | 22.34% | 20.1%
4.1 | 22.22% | 21.84% | 19.7%
4.11 | 21.72% | 21.34% | 19.2%
4.12 | 21.22% | 20.94% | 18.8%
4.13 | 20.72% | 20.44% | 18.4%
4.14 | 20.22% | 19.94% | 17.9%
4.15 | 19.72% | 19.54% | 17.5%
4.16 | 19.22% | 19.04% | 17.1%
4.17 | 18.72% | 18.64% | 16.7%
4.18 | 18.32% | 18.14% | 16.3%
4.19 | 17.82% | 17.74% | 15.9%
4.2 | 17.42% | 17.23% | 15.5%
4.21 | 16.92% | 16.83% | 15.1%
4.22 | 16.42% | 16.43% | 14.8%
4.23 | 16.02% | 16.03% | 14.4%
4.24 | 15.62% | 15.53% | 14%
4.25 | 15.22% | 15.13% | 13.6%
4.26 | 14.81% | 14.73% | 13.2%
4.27 | 14.41% | 14.33% | 12.9%
4.28 | 14.01% | 13.93% | 12.5%
4.29 | 13.61% | 13.63% | 12.2%
4.3 | 13.21% | 13.23% | 11.9%
4.31 | 12.81% | 12.83% | 11.6%
4.32 | 12.41% | 12.42% | 11.3%
4.33 | 12.11% | 12.12% | 11%
4.34 | 11.71% | 11.72% | 10.7%
4.35 | 11.31% | 11.42% | 10.3%
4.36 | 11.01% | 11.02% | 10%
4.37 | 10.71% | 10.72% | 9.7%
4.38 | 10.41% | 10.42% | 9.4%
4.39 | 10.11% | 10.12% | 9.1%
4.4 | 9.71% | 9.719% | 8.8%
4.41 | 9.409% | 9.419% | 8.6%
4.42 | 9.209% | 9.118% | 8.3%
4.43 | 8.909% | 8.918% | 8.1%
4.44 | 8.609% | 8.617% | 7.9%
4.45 | 8.308% | 8.317% | 7.6%
4.46 | 8.008% | 8.016% | 7.4%
4.47 | 7.808% | 7.816% | 7.1%
4.48 | 7.508% | 7.515% | 6.9%
4.49 | 7.307% | 7.315% | 6.7%
4.5 | 7.107% | 7.014% | 6.4%
4.51 | 6.807% | 6.814% | 6.2%
4.52 | 6.607% | 6.613% | 6%
4.53 | 6.406% | 6.413% | 5.9%
4.54 | 6.206% | 6.212% | 5.7%
4.55 | 6.006% | 5.912% | 5.5%
4.56 | 5.806% | 5.711% | 5.4%
4.57 | 5.606% | 5.611% | 5.2%
4.58 | 5.405% | 5.411% | 5%
4.59 | 5.305% | 5.21% | 4.9%
4.6 | 5.105% | 5.01% | 4.7%
4.61 | 4.905% | 4.81% | 4.5%
4.62 | 4.805% | 4.709% | 4.4%
4.63 | 4.605% | 4.509% | 4.2%
4.64 | 4.505% | 4.409% | 4.1%
4.65 | 4.304% | 4.208% | 4%
4.66 | 4.204% | 4.108% | 3.9%
4.67 | 4.104% | 3.908% | 3.7%
4.68 | 3.904% | 3.808% | 3.6%
4.69 | 3.804% | 3.707% | 3.5%
4.7 | 3.704% | 3.607% | 3.4%
4.71 | 3.504% | 3.407% | 3.3%
4.72 | 3.403% | 3.307% | 3.2%
4.73 | 3.303% | 3.206% | 3.1%
4.74 | 3.303% | 3.106% | 3%
4.75 | 3.203% | 3.006% | 2.9%
4.76 | 3.103% | 2.906% | 2.8%
4.77 | 3.003% | 2.906% | 2.7%
4.78 | 2.903% | 2.806% | 2.7%
4.79 | 2.803% | 2.705% | 2.6%
4.8 | 2.803% | 2.605% | 2.5%
4.81 | 2.703% | 2.505% | 2.5%
4.82 | 2.603% | 2.505% | 2.4%
4.83 | 2.503% | 2.405% | 2.4%
4.84 | 2.503% | 2.405% | 2.3%
4.85 | 2.402% | 2.305% | 2.2%
4.86 | 2.402% | 2.204% | 2.2%
4.87 | 2.302% | 2.204% | 2.1%
4.88 | 2.302% | 2.104% | 2.1%
4.89 | 2.202% | 2.104% | 2%
4.9 | 2.202% | 2.004% | 2%
4.91 | 2.102% | 2.004% | 2%
4.92 | 2.102% | 2.004% | 1.9%
4.93 | 2.102% | 1.904% | 1.9%
4.94 | 2.002% | 1.904% | 1.9%
4.95 | 2.002% | 1.904% | 1.8%
4.96 | 2.002% | 1.804% | 1.8%
4.97 | 1.902% | 1.804% | 1.8%
4.98 | 1.902% | 1.804% | 1.8%
4.99 | 1.902% | 1.703% | 1.7%
5 | 1.902% | 1.703% | 1.7%
5.01 | 1.802% | 1.703% | 1.7%
5.02 | 1.802% | 1.703% | 1.7%
5.03 | 1.802% | 1.703% | 1.7%
5.04 | 1.802% | 1.603% | 1.7%
5.05 | 1.802% | 1.603% | 1.6%
5.06 | 1.802% | 1.603% | 1.6%
5.07 | 1.702% | 1.603% | 1.6%
5.08 | 1.702% | 1.603% | 1.6%
5.09 | 1.702% | 1.603% | 1.6%
5.1 | 1.702% | 1.603% | 1.6%
5.11 | 1.702% | 1.503% | 1.6%
5.12 | 1.702% | 1.503% | 1.5%
5.13 | 1.702% | 1.503% | 1.5%
5.14 | 1.702% | 1.503% | 1.5%
5.15 | 1.702% | 1.503% | 1.5%
5.16 | 1.702% | 1.503% | 1.5%
5.17 | 1.602% | 1.503% | 1.5%
5.18 | 1.602% | 1.503% | 1.5%
5.19 | 1.602% | 1.503% | 1.5%
5.2 | 1.602% | 1.503% | 1.5%
5.21 | 1.602% | 1.503% | 1.5%
5.22 | 1.602% | 1.503% | 1.5%
5.23 | 1.602% | 1.503% | 1.5%
5.24 | 1.602% | 1.403% | 1.5%
5.25 | 1.602% | 1.403% | 1.5%
5.26 | 1.602% | 1.403% | 1.5%
5.27 | 1.602% | 1.403% | 1.5%
5.28 | 1.602% | 1.403% | 1.5%
5.29 | 1.602% | 1.403% | 1.5%
5.3 | 1.602% | 1.403% | 1.4%
5.31 | 1.602% | 1.403% | 1.4%
5.32 | 1.602% | 1.403% | 1.4%
5.33 | 1.602% | 1.403% | 1.4%
5.34 | 1.602% | 1.403% | 1.4%
5.35 | 1.602% | 1.403% | 1.4%
5.36 | 1.602% | 1.403% | 1.4%
5.37 | 1.602% | 1.403% | 1.4%
5.38 | 1.602% | 1.403% | 1.4%
5.39 | 1.602% | 1.403% | 1.4%
5.4 | 1.602% | 1.403% | 1.4%
5.41 | 1.602% | 1.403% | 1.4%
5.42 | 1.602% | 1.403% | 1.4%
5.43 | 1.502% | 1.403% | 1.4%
5.44 | 1.502% | 1.403% | 1.4%
5.45 | 1.502% | 1.403% | 1.4%
5.46 | 1.502% | 1.403% | 1.4%
5.47 | 1.502% | 1.403% | 1.4%
5.48 | 1.502% | 1.403% | 1.4%
5.49 | 1.502% | 1.403% | 1.4%
5.5 | 1.502% | 1.403% | 1.4%
5.51 | 1.502% | 1.403% | 1.4%
5.52 | 1.502% | 1.403% | 1.4%
5.53 | 1.502% | 1.403% | 1.4%
5.54 | 1.502% | 1.403% | 1.4%
5.55 | 1.502% | 1.403% | 1.4%
5.56 | 1.502% | 1.403% | 1.4%
5.57 | 1.502% | 1.403% | 1.4%
5.58 | 1.502% | 1.403% | 1.4%
5.59 | 1.502% | 1.403% | 1.4%
5.6 | 1.502% | 1.403% | 1.4%
5.61 | 1.502% | 1.403% | 1.4%
5.62 | 1.502% | 1.403% | 1.4%
5.63 | 1.502% | 1.403% | 1.4%
5.64 | 1.502% | 1.403% | 1.4%
5.65 | 1.502% | 1.403% | 1.4%
5.66 | 1.502% | 1.403% | 1.4%
5.67 | 1.502% | 1.403% | 1.4%
5.68 | 1.502% | 1.403% | 1.4%
5.69 | 1.502% | 1.403% | 1.4%
5.7 | 1.502% | 1.403% | 1.4%
5.71 | 1.502% | 1.403% | 1.4%
5.72 | 1.502% | 1.403% | 1.4%
5.73 | 1.502% | 1.403% | 1.4%
5.74 | 1.502% | 1.403% | 1.4%
5.75 | 1.502% | 1.403% | 1.4%
5.76 | 1.502% | 1.403% | 1.3%
5.77 | 1.502% | 1.403% | 1.3%
5.78 | 1.502% | 1.403% | 1.3%
5.79 | 1.502% | 1.403% | 1.3%
5.8 | 1.502% | 1.403% | 1.3%
5.81 | 1.502% | 1.303% | 1.3%
5.82 | 1.502% | 1.303% | 1.3%
5.83 | 1.502% | 1.303% | 1.3%
5.84 | 1.502% | 1.303% | 1.3%
5.85 | 1.502% | 1.303% | 1.3%
5.86 | 1.502% | 1.303% | 1.3%
5.87 | 1.502% | 1.303% | 1.3%
5.88 | 1.502% | 1.303% | 1.3%
5.89 | 1.502% | 1.303% | 1.3%
5.9 | 1.502% | 1.303% | 1.3%
5.91 | 1.502% | 1.303% | 1.3%
5.92 | 1.502% | 1.303% | 1.3%
5.93 | 1.502% | 1.303% | 1.3%
5.94 | 1.502% | 1.303% | 1.3%
5.95 | 1.502% | 1.303% | 1.3%
5.96 | 1.502% | 1.303% | 1.3%
5.97 | 1.502% | 1.303% | 1.3%
5.98 | 1.502% | 1.303% | 1.3%
5.99 | 1.502% | 1.303% | 1.3%
6 | 1.502% | 1.303% | 1.3%
6.01 | 1.502% | 1.303% | 1.3%
6.02 | 1.502% | 1.303% | 1.3%
6.03 | 1.502% | 1.303% | 1.3%
6.04 | 1.502% | 1.303% | 1.3%
6.05 | 1.502% | 1.303% | 1.3%
6.06 | 1.502% | 1.303% | 1.3%
6.07 | 1.502% | 1.303% | 1.3%
6.08 | 1.502% | 1.303% | 1.3%
6.09 | 1.502% | 1.303% | 1.3%
6.1 | 1.502% | 1.303% | 1.3%
6.11 | 1.502% | 1.303% | 1.3%
6.12 | 1.502% | 1.303% | 1.3%
6.13 | 1.502% | 1.303% | 1.3%
6.14 | 1.502% | 1.303% | 1.3%
6.15 | 1.401% | 1.303% | 1.3%
6.16 | 1.401% | 1.303% | 1.3%
6.17 | 1.401% | 1.303% | 1.3%
6.18 | 1.401% | 1.303% | 1.3%
6.19 | 1.401% | 1.303% | 1.3%
6.2 | 1.401% | 1.303% | 1.3%
6.21 | 1.401% | 1.303% | 1.3%
6.22 | 1.401% | 1.303% | 1.3%
6.23 | 1.401% | 1.303% | 1.3%
6.24 | 1.401% | 1.303% | 1.3%
6.25 | 1.401% | 1.303% | 1.3%
6.26 | 1.401% | 1.303% | 1.3%
6.27 | 1.401% | 1.303% | 1.3%
6.28 | 1.401% | 1.303% | 1.3%
6.29 | 1.401% | 1.303% | 1.3%
6.3 | 1.401% | 1.303% | 1.3%
6.31 | 1.401% | 1.303% | 1.3%
6.32 | 1.401% | 1.303% | 1.3%
6.33 | 1.401% | 1.303% | 1.3%
6.34 | 1.401% | 1.303% | 1.3%
6.35 | 1.401% | 1.303% | 1.3%
6.36 | 1.401% | 1.303% | 1.3%
6.37 | 1.401% | 1.303% | 1.3%
6.38 | 1.401% | 1.303% | 1.3%
6.39 | 1.401% | 1.303% | 1.3%
6.4 | 1.401% | 1.303% | 1.2%
6.41 | 1.401% | 1.303% | 1.2%
6.42 | 1.401% | 1.303% | 1.2%
6.43 | 1.401% | 1.303% | 1.2%
6.44 | 1.401% | 1.303% | 1.2%
6.45 | 1.401% | 1.303% | 1.2%
6.46 | 1.401% | 1.303% | 1.2%
6.47 | 1.401% | 1.303% | 1.2%
6.48 | 1.401% | 1.303% | 1.2%
6.49 | 1.401% | 1.303% | 1.2%
6.5 | 1.401% | 1.303% | 1.2%
6.51 | 1.401% | 1.303% | 1.2%
6.52 | 1.401% | 1.303% | 1.2%
6.53 | 1.401% | 1.303% | 1.2%
6.54 | 1.401% | 1.303% | 1.2%
6.55 | 1.401% | 1.303% | 1.2%
6.56 | 1.401% | 1.303% | 1.2%
6.57 | 1.401% | 1.303% | 1.2%
6.58 | 1.401% | 1.303% | 1.2%
6.59 | 1.401% | 1.303% | 1.2%
6.6 | 1.401% | 1.202% | 1.2%
6.61 | 1.401% | 1.202% | 1.2%
6.62 | 1.401% | 1.202% | 1.2%
6.63 | 1.401% | 1.202% | 1.2%
6.64 | 1.401% | 1.202% | 1.2%
6.65 | 1.401% | 1.202% | 1.2%
6.66 | 1.401% | 1.202% | 1.2%
6.67 | 1.401% | 1.202% | 1.2%
6.68 | 1.401% | 1.202% | 1.2%
6.69 | 1.401% | 1.202% | 1.2%
6.7 | 1.401% | 1.202% | 1.2%
6.71 | 1.401% | 1.202% | 1.2%
6.72 | 1.401% | 1.202% | 1.2%
6.73 | 1.401% | 1.202% | 1.2%
6.74 | 1.401% | 1.202% | 1.2%
6.75 | 1.401% | 1.202% | 1.2%
6.76 | 1.401% | 1.202% | 1.2%
6.77 | 1.401% | 1.202% | 1.2%
6.78 | 1.401% | 1.202% | 1.2%
6.79 | 1.401% | 1.202% | 1.2%
6.8 | 1.401% | 1.202% | 1.2%
6.81 | 1.401% | 1.202% | 1.2%
6.82 | 1.401% | 1.202% | 1.2%
6.83 | 1.401% | 1.202% | 1.2%
6.84 | 1.401% | 1.202% | 1.2%
6.85 | 1.401% | 1.202% | 1.2%
6.86 | 1.401% | 1.202% | 1.2%
6.87 | 1.401% | 1.202% | 1.2%
6.88 | 1.401% | 1.202% | 1.2%
6.89 | 1.401% | 1.202% | 1.2%
6.9 | 1.401% | 1.202% | 1.2%
6.91 | 1.401% | 1.202% | 1.2%
6.92 | 1.401% | 1.202% | 1.2%
6.93 | 1.401% | 1.202% | 1.2%
6.94 | 1.401% | 1.202% | 1.2%
6.95 | 1.301% | 1.202% | 1.2%
6.96 | 1.301% | 1.202% | 1.2%
6.97 | 1.301% | 1.202% | 1.2%
6.98 | 1.301% | 1.202% | 1.2%
6.99 | 1.301% | 1.202% | 1.2%
7 | 1.301% | 1.202% | 1.2%
7.01 | 1.301% | 1.202% | 1.2%
7.02 | 1.301% | 1.202% | 1.2%
7.03 | 1.301% | 1.202% | 1.2%
7.04 | 1.301% | 1.202% | 1.2%
7.05 | 1.301% | 1.202% | 1.2%
7.06 | 1.301% | 1.202% | 1.2%
7.07 | 1.301% | 1.202% | 1.2%
7.08 | 1.301% | 1.202% | 1.2%
7.09 | 1.301% | 1.202% | 1.2%
7.1 | 1.301% | 1.202% | 1.2%
7.11 | 1.301% | 1.202% | 1.2%
7.12 | 1.301% | 1.202% | 1.2%
7.13 | 1.301% | 1.202% | 1.2%
7.14 | 1.301% | 1.202% | 1.2%
7.15 | 1.301% | 1.202% | 1.2%
7.16 | 1.301% | 1.202% | 1.2%
7.17 | 1.301% | 1.202% | 1.2%
7.18 | 1.301% | 1.202% | 1.2%
7.19 | 1.301% | 1.202% | 1.2%
7.2 | 1.301% | 1.202% | 1.2%
7.21 | 1.301% | 1.202% | 1.2%
7.22 | 1.301% | 1.202% | 1.2%
7.23 | 1.301% | 1.202% | 1.2%
7.24 | 1.301% | 1.202% | 1.2%
7.25 | 1.301% | 1.202% | 1.2%
7.26 | 1.301% | 1.202% | 1.2%
7.27 | 1.301% | 1.202% | 1.2%
7.28 | 1.301% | 1.202% | 1.2%
7.29 | 1.301% | 1.202% | 1.2%
7.3 | 1.301% | 1.202% | 1.2%
7.31 | 1.301% | 1.202% | 1.2%
7.32 | 1.301% | 1.202% | 1.2%
7.33 | 1.301% | 1.202% | 1.2%
7.34 | 1.301% | 1.202% | 1.2%
7.35 | 1.301% | 1.202% | 1.2%
7.36 | 1.301% | 1.102% | 1.2%
7.37 | 1.301% | 1.102% | 1.2%
7.38 | 1.301% | 1.102% | 1.2%
7.39 | 1.301% | 1.102% | 1.2%
7.4 | 1.301% | 1.102% | 1.2%
7.41 | 1.301% | 1.102% | 1.2%
7.42 | 1.301% | 1.102% | 1.2%
7.43 | 1.301% | 1.102% | 1.1%
7.44 | 1.301% | 1.102% | 1.1%
7.45 | 1.301% | 1.102% | 1.1%
7.46 | 1.301% | 1.102% | 1.1%
7.47 | 1.301% | 1.102% | 1.1%
7.48 | 1.301% | 1.102% | 1.1%
7.49 | 1.301% | 1.102% | 1.1%
7.5 | 1.301% | 1.102% | 1.1%
7.51 | 1.301% | 1.102% | 1.1%
7.52 | 1.301% | 1.102% | 1.1%
7.53 | 1.301% | 1.102% | 1.1%
7.54 | 1.301% | 1.102% | 1.1%
7.55 | 1.301% | 1.102% | 1.1%
7.56 | 1.301% | 1.102% | 1.1%
7.57 | 1.301% | 1.102% | 1.1%
7.58 | 1.301% | 1.102% | 1.1%
7.59 | 1.301% | 1.102% | 1.1%
7.6 | 1.301% | 1.102% | 1.1%
7.61 | 1.301% | 1.102% | 1.1%
7.62 | 1.301% | 1.102% | 1.1%
7.63 | 1.301% | 1.102% | 1.1%
7.64 | 1.301% | 1.102% | 1.1%
7.65 | 1.301% | 1.102% | 1.1%
7.66 | 1.301% | 1.102% | 1.1%
7.67 | 1.301% | 1.102% | 1.1%
7.68 | 1.301% | 1.102% | 1.1%
7.69 | 1.301% | 1.102% | 1.1%
7.7 | 1.301% | 1.102% | 1.1%
7.71 | 1.301% | 1.102% | 1.1%
7.72 | 1.301% | 1.102% | 1.1%
7.73 | 1.301% | 1.102% | 1.1%
7.74 | 1.301% | 1.102% | 1.1%
7.75 | 1.301% | 1.102% | 1.1%
7.76 | 1.301% | 1.102% | 1.1%
7.77 | 1.301% | 1.102% | 1.1%
7.78 | 1.301% | 1.102% | 1.1%
7.79 | 1.301% | 1.102% | 1.1%
7.8 | 1.301% | 1.102% | 1.1%
7.81 | 1.301% | 1.102% | 1.1%
7.82 | 1.301% | 1.102% | 1.1%
7.83 | 1.301% | 1.102% | 1.1%
7.84 | 1.301% | 1.102% | 1.1%
7.85 | 1.301% | 1.102% | 1.1%
7.86 | 1.301% | 1.102% | 1.1%
7.87 | 1.301% | 1.102% | 1.1%
7.88 | 1.301% | 1.102% | 1.1%
7.89 | 1.301% | 1.102% | 1.1%
7.9 | 1.301% | 1.102% | 1.1%
7.91 | 1.301% | 1.102% | 1.1%
7.92 | 1.301% | 1.102% | 1.1%
7.93 | 1.301% | 1.102% | 1.1%
7.94 | 1.301% | 1.102% | 1.1%
7.95 | 1.301% | 1.102% | 1.1%
7.96 | 1.301% | 1.102% | 1.1%
7.97 | 1.301% | 1.102% | 1.1%
7.98 | 1.301% | 1.102% | 1.1%
7.99 | 1.301% | 1.102% | 1.1%
8 | 1.301% | 1.102% | 1.1%
8.01 | 1.201% | 1.102% | 1.1%
8.02 | 1.201% | 1.102% | 1.1%
8.03 | 1.201% | 1.102% | 1.1%
8.04 | 1.201% | 1.102% | 1.1%
8.05 | 1.201% | 1.102% | 1.1%
8.06 | 1.201% | 1.102% | 1.1%
8.07 | 1.201% | 1.102% | 1.1%
8.08 | 1.201% | 1.102% | 1.1%
8.09 | 1.201% | 1.102% | 1.1%
8.1 | 1.201% | 1.102% | 1.1%
8.11 | 1.201% | 1.102% | 1.1%
8.12 | 1.201% | 1.102% | 1.1%
8.13 | 1.201% | 1.102% | 1.1%
8.14 | 1.201% | 1.102% | 1.1%
8.15 | 1.201% | 1.102% | 1.1%
8.16 | 1.201% | 1.102% | 1.1%
8.17 | 1.201% | 1.102% | 1.1%
8.18 | 1.201% | 1.102% | 1.1%
8.19 | 1.201% | 1.102% | 1.1%
8.2 | 1.201% | 1.102% | 1.1%
8.21 | 1.201% | 1.102% | 1.1%
8.22 | 1.201% | 1.102% | 1.1%
8.23 | 1.201% | 1.102% | 1.1%
8.24 | 1.201% | 1.102% | 1.1%
8.25 | 1.201% | 1.102% | 1.1%
8.26 | 1.201% | 1.102% | 1.1%
8.27 | 1.201% | 1.102% | 1.1%
8.28 | 1.201% | 1.102% | 1.1%
8.29 | 1.201% | 1.102% | 1.1%
8.3 | 1.201% | 1.102% | 1.1%
8.31 | 1.201% | 1.102% | 1.1%
8.32 | 1.201% | 1.102% | 1.1%
8.33 | 1.201% | 1.102% | 1.1%
8.34 | 1.201% | 1.102% | 1.1%
8.35 | 1.201% | 1.102% | 1.1%
8.36 | 1.201% | 1.102% | 1.1%
8.37 | 1.201% | 1.002% | 1.1%
8.38 | 1.201% | 1.002% | 1%
8.39 | 1.201% | 1.002% | 1%
8.4 | 1.201% | 1.002% | 1%
8.41 | 1.201% | 1.002% | 1%
8.42 | 1.201% | 1.002% | 1%
8.43 | 1.201% | 1.002% | 1%
8.44 | 1.201% | 1.002% | 1%
8.45 | 1.201% | 1.002% | 1%
8.46 | 1.201% | 1.002% | 1%
8.47 | 1.201% | 1.002% | 1%
8.48 | 1.201% | 1.002% | 1%
8.49 | 1.201% | 1.002% | 1%
8.5 | 1.201% | 1.002% | 1%
8.51 | 1.201% | 1.002% | 1%
8.52 | 1.201% | 1.002% | 1%
8.53 | 1.201% | 1.002% | 1%
8.54 | 1.201% | 1.002% | 1%
8.55 | 1.201% | 1.002% | 1%
8.56 | 1.201% | 1.002% | 1%
8.57 | 1.201% | 1.002% | 1%
8.58 | 1.201% | 1.002% | 1%
8.59 | 1.201% | 1.002% | 1%
8.6 | 1.201% | 1.002% | 1%
8.61 | 1.201% | 1.002% | 1%
8.62 | 1.201% | 1.002% | 1%
8.63 | 1.201% | 1.002% | 1%
8.64 | 1.201% | 1.002% | 1%
8.65 | 1.201% | 1.002% | 1%
8.66 | 1.201% | 1.002% | 1%
8.67 | 1.201% | 1.002% | 1%
8.68 | 1.201% | 1.002% | 1%
8.69 | 1.201% | 1.002% | 1%
8.7 | 1.201% | 1.002% | 1%
8.71 | 1.201% | 1.002% | 1%
8.72 | 1.201% | 1.002% | 1%
8.73 | 1.201% | 1.002% | 1%
8.74 | 1.201% | 1.002% | 1%
8.75 | 1.201% | 1.002% | 1%
8.76 | 1.201% | 1.002% | 1%
8.77 | 1.201% | 1.002% | 1%
8.78 | 1.201% | 1.002% | 1%
8.79 | 1.201% | 1.002% | 1%
8.8 | 1.201% | 1.002% | 1%
8.81 | 1.201% | 1.002% | 1%
8.82 | 1.201% | 1.002% | 1%
8.83 | 1.201% | 1.002% | 1%
|  | 1.002% | 1%
|  | 1.002% | 1%
|  | 1.002% | 1%
|  | 1.002% | 1%
|  | 1.002% | 1%
|  | 1.002% | 1%
|  | 1.002% | 1%
1.002%
1.002%
1.002%

## 12 MeV PDDs
Energy: | 12 MeV
SSD: | 100 cm
PDD
Depth | 6x6+ | 4x4 | 3x3
0 |  | 89.49%
0.01 |  | 89.49% |  | Energy | SSD | Field size | R100 | R50 | R95 | R90 | R80 | Rp | Ds | PDD(2.78) | PDD(2.9)
0.02 |  | 89.49% |  | 12 MeV | 100.00 cm | 3x3 | 1.96 cm | 4.56 cm | 2.76 cm | 3.08 cm | 3.54 cm | 6.08 cm | 91.2% | 94.8% | 92.9%
0.03 | 87.09% | 89.49% |  | 12 MeV | 100.00 cm | 4x4 | 2.17 cm | 4.86 cm | 3.21 cm | 3.54 cm | 3.97 cm | 6.11 cm | 89.4% | 99% | 98.1%
0.04 | 87.09% | 89.49% | 91.2% | 12 MeV | 100.00 cm | 6x6+ | 2.78 cm | 5.02 cm | 3.66 cm | 3.93 cm | 4.30 cm | 6.05 cm | 87% | 100% | 100%
0.05 | 87.09% | 89.49% | 91.3%
0.06 | 87.19% | 89.49% | 91.3%
0.07 | 87.19% | 89.49% | 91.4%
0.08 | 87.29% | 89.49% | 91.4%
0.09 | 87.29% | 89.49% | 91.5%
0.1 | 87.39% | 89.59% | 91.5%
0.11 | 87.39% | 89.59% | 91.6%
0.12 | 87.49% | 89.59% | 91.7%
0.13 | 87.49% | 89.69% | 91.7%
0.14 | 87.59% | 89.69% | 91.8%
0.15 | 87.59% | 89.79% | 91.9%
0.16 | 87.69% | 89.79% | 92.1%
0.17 | 87.79% | 89.89% | 92.2%
0.18 | 87.89% | 89.99% | 92.3%
0.19 | 87.89% | 90.09% | 92.4%
0.2 | 87.99% | 90.19% | 92.5%
0.21 | 88.09% | 90.29% | 92.6%
0.22 | 88.19% | 90.29% | 92.7%
0.23 | 88.29% | 90.39% | 92.8%
0.24 | 88.39% | 90.59% | 92.9%
0.25 | 88.49% | 90.69% | 93%
0.26 | 88.59% | 90.79% | 93.1%
0.27 | 88.69% | 90.89% | 93.2%
0.28 | 88.79% | 90.99% | 93.3%
0.29 | 88.99% | 91.09% | 93.4%
0.3 | 89.09% | 91.19% | 93.5%
0.31 | 89.19% | 91.29% | 93.6%
0.32 | 89.29% | 91.39% | 93.7%
0.33 | 89.39% | 91.49% | 93.8%
0.34 | 89.49% | 91.59% | 93.9%
0.35 | 89.59% | 91.69% | 94%
0.36 | 89.79% | 91.79% | 94.1%
0.37 | 89.89% | 91.89% | 94.2%
0.38 | 89.99% | 91.99% | 94.3%
0.39 | 90.09% | 92.09% | 94.4%
0.4 | 90.09% | 92.19% | 94.5%
0.41 | 90.19% | 92.29% | 94.6%
0.42 | 90.29% | 92.39% | 94.7%
0.43 | 90.39% | 92.49% | 94.8%
0.44 | 90.49% | 92.59% | 94.9%
0.45 | 90.59% | 92.69% | 95%
0.46 | 90.69% | 92.79% | 95.1%
0.47 | 90.69% | 92.89% | 95.2%
0.48 | 90.79% | 92.99% | 95.2%
0.49 | 90.89% | 93.09% | 95.3%
0.5 | 90.99% | 93.19% | 95.4%
0.51 | 90.99% | 93.29% | 95.5%
0.52 | 91.09% | 93.39% | 95.6%
0.53 | 91.19% | 93.49% | 95.7%
0.54 | 91.29% | 93.59% | 95.8%
0.55 | 91.29% | 93.69% | 95.8%
0.56 | 91.39% | 93.79% | 95.9%
0.57 | 91.49% | 93.79% | 96%
0.58 | 91.59% | 93.89% | 96.1%
0.59 | 91.69% | 93.99% | 96.2%
0.6 | 91.69% | 94.09% | 96.3%
0.61 | 91.79% | 94.19% | 96.3%
0.62 | 91.89% | 94.19% | 96.4%
0.63 | 91.89% | 94.29% | 96.5%
0.64 | 91.99% | 94.39% | 96.5%
0.65 | 92.09% | 94.39% | 96.6%
0.66 | 92.09% | 94.49% | 96.6%
0.67 | 92.19% | 94.59% | 96.7%
0.68 | 92.19% | 94.59% | 96.7%
0.69 | 92.29% | 94.69% | 96.8%
0.7 | 92.29% | 94.79% | 96.8%
0.71 | 92.39% | 94.79% | 96.9%
0.72 | 92.49% | 94.89% | 96.9%
0.73 | 92.49% | 94.89% | 97%
0.74 | 92.59% | 94.99% | 97%
0.75 | 92.59% | 95.1% | 97%
0.76 | 92.69% | 95.1% | 97.1%
0.77 | 92.69% | 95.2% | 97.1%
0.78 | 92.79% | 95.2% | 97.2%
0.79 | 92.89% | 95.3% | 97.2%
0.8 | 92.89% | 95.3% | 97.3%
0.81 | 92.99% | 95.4% | 97.3%
0.82 | 92.99% | 95.5% | 97.4%
0.83 | 93.09% | 95.5% | 97.5%
0.84 | 93.09% | 95.6% | 97.5%
0.85 | 93.19% | 95.6% | 97.6%
0.86 | 93.29% | 95.7% | 97.6%
0.87 | 93.29% | 95.7% | 97.7%
0.88 | 93.39% | 95.8% | 97.7%
0.89 | 93.39% | 95.8% | 97.8%
0.9 | 93.49% | 95.9% | 97.8%
0.91 | 93.49% | 95.9% | 97.9%
0.92 | 93.59% | 96% | 97.9%
0.93 | 93.59% | 96% | 97.9%
0.94 | 93.69% | 96.1% | 98%
0.95 | 93.69% | 96.1% | 98%
0.96 | 93.79% | 96.2% | 98.1%
0.97 | 93.89% | 96.2% | 98.1%
0.98 | 93.89% | 96.3% | 98.1%
0.99 | 93.99% | 96.3% | 98.2%
1 | 93.99% | 96.4% | 98.2%
1.01 | 93.99% | 96.4% | 98.2%
1.02 | 94.09% | 96.5% | 98.3%
1.03 | 94.09% | 96.5% | 98.3%
1.04 | 94.19% | 96.6% | 98.4%
1.05 | 94.19% | 96.6% | 98.4%
1.06 | 94.29% | 96.7% | 98.5%
1.07 | 94.29% | 96.7% | 98.5%
1.08 | 94.29% | 96.7% | 98.6%
1.09 | 94.39% | 96.8% | 98.6%
1.1 | 94.39% | 96.8% | 98.6%
1.11 | 94.39% | 96.9% | 98.7%
1.12 | 94.49% | 96.9% | 98.7%
1.13 | 94.49% | 97% | 98.7%
1.14 | 94.59% | 97% | 98.7%
1.15 | 94.59% | 97.1% | 98.7%
1.16 | 94.59% | 97.1% | 98.8%
1.17 | 94.69% | 97.2% | 98.8%
1.18 | 94.69% | 97.2% | 98.8%
1.19 | 94.79% | 97.3% | 98.8%
1.2 | 94.79% | 97.3% | 98.8%
1.21 | 94.89% | 97.3% | 98.8%
1.22 | 94.89% | 97.4% | 98.8%
1.23 | 94.89% | 97.4% | 98.9%
1.24 | 94.99% | 97.5% | 98.9%
1.25 | 94.99% | 97.5% | 99%
1.26 | 94.99% | 97.6% | 99%
1.27 | 95.1% | 97.6% | 99%
1.28 | 95.1% | 97.6% | 99.1%
1.29 | 95.1% | 97.7% | 99.1%
1.3 | 95.2% | 97.7% | 99.1%
1.31 | 95.2% | 97.8% | 99.2%
1.32 | 95.3% | 97.8% | 99.2%
1.33 | 95.3% | 97.9% | 99.3%
1.34 | 95.3% | 97.9% | 99.3%
1.35 | 95.4% | 97.9% | 99.3%
1.36 | 95.4% | 98% | 99.3%
1.37 | 95.5% | 98% | 99.3%
1.38 | 95.5% | 98.1% | 99.4%
1.39 | 95.6% | 98.1% | 99.4%
1.4 | 95.6% | 98.2% | 99.4%
1.41 | 95.6% | 98.2% | 99.4%
1.42 | 95.7% | 98.2% | 99.5%
1.43 | 95.7% | 98.3% | 99.5%
1.44 | 95.8% | 98.3% | 99.5%
1.45 | 95.8% | 98.4% | 99.5%
1.46 | 95.8% | 98.4% | 99.6%
1.47 | 95.9% | 98.4% | 99.6%
1.48 | 95.9% | 98.5% | 99.6%
1.49 | 96% | 98.5% | 99.6%
1.5 | 96% | 98.6% | 99.6%
1.51 | 96% | 98.6% | 99.6%
1.52 | 96.1% | 98.6% | 99.6%
1.53 | 96.1% | 98.7% | 99.7%
1.54 | 96.2% | 98.7% | 99.7%
1.55 | 96.2% | 98.7% | 99.7%
1.56 | 96.2% | 98.8% | 99.7%
1.57 | 96.3% | 98.8% | 99.7%
1.58 | 96.3% | 98.9% | 99.7%
1.59 | 96.4% | 98.9% | 99.7%
1.6 | 96.4% | 98.9% | 99.7%
1.61 | 96.5% | 99% | 99.7%
1.62 | 96.5% | 99% | 99.8%
1.63 | 96.5% | 99% | 99.8%
1.64 | 96.6% | 99.1% | 99.8%
1.65 | 96.6% | 99.1% | 99.8%
1.66 | 96.7% | 99.1% | 99.9%
1.67 | 96.7% | 99.2% | 99.9%
1.68 | 96.8% | 99.2% | 99.9%
1.69 | 96.8% | 99.2% | 99.9%
1.7 | 96.8% | 99.3% | 99.9%
1.71 | 96.9% | 99.3% | 99.9%
1.72 | 96.9% | 99.3% | 99.9%
1.73 | 96.9% | 99.3% | 99.9%
1.74 | 97% | 99.4% | 99.9%
1.75 | 97% | 99.4% | 99.9%
1.76 | 97% | 99.4% | 99.9%
1.77 | 97.1% | 99.4% | 99.9%
1.78 | 97.1% | 99.5% | 99.9%
1.79 | 97.1% | 99.5% | 99.9%
1.8 | 97.2% | 99.5% | 99.9%
1.81 | 97.2% | 99.5% | 99.9%
1.82 | 97.3% | 99.5% | 99.9%
1.83 | 97.3% | 99.6% | 99.9%
1.84 | 97.3% | 99.6% | 99.9%
1.85 | 97.4% | 99.6% | 99.9%
1.86 | 97.4% | 99.6% | 99.9%
1.87 | 97.4% | 99.6% | 99.9%
1.88 | 97.5% | 99.7% | 99.9%
1.89 | 97.5% | 99.7% | 99.9%
1.9 | 97.5% | 99.7% | 99.9%
1.91 | 97.6% | 99.7% | 100%
1.92 | 97.6% | 99.7% | 100%
1.93 | 97.6% | 99.7% | 100%
1.94 | 97.7% | 99.8% | 100%
1.95 | 97.7% | 99.8% | 100%
1.96 | 97.7% | 99.8% | 100%
1.97 | 97.8% | 99.8% | 100%
1.98 | 97.8% | 99.8% | 100%
1.99 | 97.8% | 99.8% | 99.9%
2 | 97.9% | 99.9% | 99.9%
2.01 | 97.9% | 99.9% | 99.9%
2.02 | 97.9% | 99.9% | 99.9%
2.03 | 98% | 99.9% | 99.8%
2.04 | 98% | 99.9% | 99.8%
2.05 | 98% | 99.9% | 99.8%
2.06 | 98.1% | 99.9% | 99.8%
2.07 | 98.1% | 99.9% | 99.7%
2.08 | 98.1% | 100% | 99.7%
2.09 | 98.2% | 100% | 99.7%
2.1 | 98.2% | 100% | 99.7%
2.11 | 98.3% | 100% | 99.7%
2.12 | 98.3% | 100% | 99.6%
2.13 | 98.3% | 100% | 99.6%
2.14 | 98.4% | 100% | 99.6%
2.15 | 98.4% | 100% | 99.6%
2.16 | 98.5% | 100% | 99.6%
2.17 | 98.5% | 100% | 99.6%
2.18 | 98.5% | 100% | 99.5%
2.19 | 98.6% | 100% | 99.5%
2.2 | 98.6% | 100% | 99.5%
2.21 | 98.6% | 100% | 99.5%
2.22 | 98.7% | 100% | 99.4%
2.23 | 98.7% | 100% | 99.3%
2.24 | 98.7% | 100% | 99.3%
2.25 | 98.8% | 100% | 99.2%
2.26 | 98.8% | 100% | 99.2%
2.27 | 98.8% | 100% | 99.1%
2.28 | 98.9% | 100% | 99.1%
2.29 | 98.9% | 100% | 99%
2.3 | 98.9% | 100% | 98.9%
2.31 | 99% | 100% | 98.9%
2.32 | 99% | 100% | 98.8%
2.33 | 99% | 100% | 98.7%
2.34 | 99% | 100% | 98.7%
2.35 | 99% | 100% | 98.6%
2.36 | 99.1% | 100% | 98.5%
2.37 | 99.1% | 100% | 98.5%
2.38 | 99.1% | 100% | 98.4%
2.39 | 99.1% | 100% | 98.3%
2.4 | 99.2% | 100% | 98.3%
2.41 | 99.2% | 100% | 98.2%
2.42 | 99.2% | 100% | 98.1%
2.43 | 99.3% | 100% | 98.1%
2.44 | 99.3% | 100% | 98%
2.45 | 99.3% | 100% | 97.9%
2.46 | 99.4% | 100% | 97.9%
2.47 | 99.4% | 99.9% | 97.8%
2.48 | 99.5% | 99.9% | 97.7%
2.49 | 99.5% | 99.9% | 97.6%
2.5 | 99.5% | 99.9% | 97.6%
2.51 | 99.6% | 99.9% | 97.5%
2.52 | 99.6% | 99.9% | 97.4%
2.53 | 99.6% | 99.8% | 97.3%
2.54 | 99.7% | 99.8% | 97.3%
2.55 | 99.7% | 99.8% | 97.2%
2.56 | 99.7% | 99.8% | 97.1%
2.57 | 99.7% | 99.8% | 97%
2.58 | 99.7% | 99.7% | 96.9%
2.59 | 99.7% | 99.7% | 96.8%
2.6 | 99.8% | 99.7% | 96.7%
2.61 | 99.8% | 99.7% | 96.6%
2.62 | 99.8% | 99.6% | 96.4%
2.63 | 99.8% | 99.6% | 96.3%
2.64 | 99.8% | 99.5% | 96.2%
2.65 | 99.9% | 99.5% | 96.1%
2.66 | 99.9% | 99.5% | 96%
2.67 | 99.9% | 99.4% | 95.9%
2.68 | 99.9% | 99.4% | 95.8%
2.69 | 99.9% | 99.3% | 95.7%
2.7 | 100% | 99.3% | 95.6%
2.71 | 100% | 99.3% | 95.5%
2.72 | 100% | 99.2% | 95.4%
2.73 | 100% | 99.2% | 95.2%
2.74 | 100% | 99.1% | 95.1%
2.75 | 100% | 99.1% | 95%
2.76 | 100% | 99% | 94.9%
2.77 | 100% | 99% | 94.8%
2.78 | 100% | 98.9% | 94.6%
2.79 | 100% | 98.8% | 94.5%
2.8 | 100% | 98.8% | 94.4%
2.81 | 100% | 98.7% | 94.2%
2.82 | 100% | 98.7% | 94.1%
2.83 | 100% | 98.6% | 94%
2.84 | 100% | 98.5% | 93.8%
2.85 | 100% | 98.5% | 93.7%
2.86 | 100% | 98.4% | 93.5%
2.87 | 100% | 98.3% | 93.4%
2.88 | 100% | 98.2% | 93.2%
2.89 | 100% | 98.2% | 93.1%
2.9 | 100% | 98.1% | 92.9%
2.91 | 100% | 98% | 92.7%
2.92 | 100% | 97.9% | 92.6%
2.93 | 100% | 97.9% | 92.4%
2.94 | 100% | 97.8% | 92.3%
2.95 | 99.9% | 97.7% | 92.1%
2.96 | 99.9% | 97.6% | 91.9%
2.97 | 99.9% | 97.5% | 91.8%
2.98 | 99.9% | 97.4% | 91.6%
2.99 | 99.9% | 97.3% | 91.4%
3 | 99.9% | 97.2% | 91.3%
3.01 | 99.9% | 97.2% | 91.1%
3.02 | 99.9% | 97.1% | 90.9%
3.03 | 99.8% | 97% | 90.7%
3.04 | 99.8% | 96.9% | 90.6%
3.05 | 99.8% | 96.8% | 90.4%
3.06 | 99.7% | 96.7% | 90.2%
3.07 | 99.7% | 96.6% | 90%
3.08 | 99.7% | 96.5% | 89.8%
3.09 | 99.7% | 96.4% | 89.7%
3.1 | 99.6% | 96.3% | 89.5%
3.11 | 99.6% | 96.2% | 89.3%
3.12 | 99.6% | 96.1% | 89.1%
3.13 | 99.5% | 96% | 89%
3.14 | 99.5% | 95.9% | 88.8%
3.15 | 99.5% | 95.7% | 88.6%
3.16 | 99.4% | 95.6% | 88.4%
3.17 | 99.4% | 95.5% | 88.2%
3.18 | 99.3% | 95.4% | 88%
3.19 | 99.3% | 95.3% | 87.8%
3.2 | 99.3% | 95.2% | 87.6%
3.21 | 99.2% | 95.1% | 87.3%
3.22 | 99.2% | 94.89% | 87.1%
3.23 | 99.1% | 94.79% | 86.9%
3.24 | 99.1% | 94.69% | 86.7%
3.25 | 99.1% | 94.59% | 86.5%
3.26 | 99% | 94.49% | 86.3%
3.27 | 98.9% | 94.29% | 86.1%
3.28 | 98.9% | 94.19% | 85.9%
3.29 | 98.8% | 94.09% | 85.7%
3.3 | 98.8% | 93.89% | 85.5%
3.31 | 98.7% | 93.79% | 85.2%
3.32 | 98.6% | 93.69% | 85%
3.33 | 98.6% | 93.49% | 84.8%
3.34 | 98.5% | 93.39% | 84.6%
3.35 | 98.4% | 93.19% | 84.3%
3.36 | 98.3% | 93.09% | 84.1%
3.37 | 98.3% | 92.99% | 83.9%
3.38 | 98.2% | 92.79% | 83.7%
3.39 | 98.1% | 92.59% | 83.4%
3.4 | 98% | 92.49% | 83.2%
3.41 | 97.9% | 92.29% | 83%
3.42 | 97.8% | 92.19% | 82.7%
3.43 | 97.7% | 91.99% | 82.5%
3.44 | 97.6% | 91.79% | 82.2%
3.45 | 97.5% | 91.69% | 82%
3.46 | 97.4% | 91.49% | 81.7%
3.47 | 97.3% | 91.29% | 81.5%
3.48 | 97.2% | 91.09% | 81.2%
3.49 | 97.1% | 90.99% | 81%
3.5 | 97% | 90.79% | 80.7%
3.51 | 96.9% | 90.59% | 80.5%
3.52 | 96.8% | 90.39% | 80.2%
3.53 | 96.7% | 90.19% | 79.9%
3.54 | 96.6% | 89.99% | 79.6%
3.55 | 96.5% | 89.79% | 79.4%
3.56 | 96.3% | 89.59% | 79.1%
3.57 | 96.2% | 89.39% | 78.8%
3.58 | 96.1% | 89.19% | 78.5%
3.59 | 96% | 88.99% | 78.3%
3.6 | 95.8% | 88.79% | 78%
3.61 | 95.7% | 88.59% | 77.7%
3.62 | 95.6% | 88.39% | 77.5%
3.63 | 95.4% | 88.19% | 77.2%
3.64 | 95.3% | 87.89% | 77%
3.65 | 95.1% | 87.69% | 76.7%
3.66 | 94.99% | 87.49% | 76.5%
3.67 | 94.79% | 87.29% | 76.2%
3.68 | 94.69% | 87.09% | 75.9%
3.69 | 94.59% | 86.79% | 75.7%
3.7 | 94.39% | 86.59% | 75.4%
3.71 | 94.19% | 86.39% | 75.1%
3.72 | 94.09% | 86.09% | 74.9%
3.73 | 93.89% | 85.89% | 74.6%
3.74 | 93.79% | 85.69% | 74.4%
3.75 | 93.59% | 85.39% | 74.1%
3.76 | 93.39% | 85.19% | 73.8%
3.77 | 93.29% | 84.98% | 73.5%
3.78 | 93.09% | 84.68% | 73.2%
3.79 | 92.89% | 84.48% | 72.9%
3.8 | 92.69% | 84.18% | 72.6%
3.81 | 92.49% | 83.98% | 72.3%
3.82 | 92.29% | 83.78% | 72%
3.83 | 92.19% | 83.48% | 71.7%
3.84 | 91.99% | 83.28% | 71.4%
3.85 | 91.79% | 82.98% | 71.1%
3.86 | 91.59% | 82.78% | 70.8%
3.87 | 91.39% | 82.48% | 70.5%
3.88 | 91.19% | 82.28% | 70.2%
3.89 | 90.89% | 81.98% | 69.9%
3.9 | 90.69% | 81.78% | 69.7%
3.91 | 90.49% | 81.48% | 69.4%
3.92 | 90.29% | 81.28% | 69.1%
3.93 | 90.09% | 80.98% | 68.8%
3.94 | 89.79% | 80.68% | 68.5%
3.95 | 89.59% | 80.48% | 68.2%
3.96 | 89.39% | 80.18% | 67.9%
3.97 | 89.19% | 79.88% | 67.6%
3.98 | 88.99% | 79.68% | 67.3%
3.99 | 88.69% | 79.38% | 67%
4 | 88.49% | 79.08% | 66.8%
4.01 | 88.29% | 78.88% | 66.5%
4.02 | 87.99% | 78.58% | 66.2%
4.03 | 87.79% | 78.28% | 65.9%
4.04 | 87.49% | 78.08% | 65.6%
4.05 | 87.29% | 77.78% | 65.3%
4.06 | 87.09% | 77.48% | 65%
4.07 | 86.79% | 77.18% | 64.7%
4.08 | 86.49% | 76.98% | 64.4%
4.09 | 86.29% | 76.68% | 64.1%
4.1 | 85.99% | 76.38% | 63.8%
4.11 | 85.79% | 76.08% | 63.5%
4.12 | 85.49% | 75.78% | 63.2%
4.13 | 85.19% | 75.58% | 62.9%
4.14 | 84.88% | 75.28% | 62.6%
4.15 | 84.58% | 74.97% | 62.3%
4.16 | 84.28% | 74.67% | 62%
4.17 | 83.98% | 74.37% | 61.8%
4.18 | 83.68% | 74.07% | 61.5%
4.19 | 83.38% | 73.77% | 61.2%
4.2 | 83.08% | 73.47% | 60.9%
4.21 | 82.78% | 73.17% | 60.6%
4.22 | 82.48% | 72.87% | 60.3%
4.23 | 82.18% | 72.57% | 60%
4.24 | 81.88% | 72.27% | 59.7%
4.25 | 81.58% | 71.97% | 59.4%
4.26 | 81.28% | 71.67% | 59.1%
4.27 | 80.88% | 71.37% | 58.8%
4.28 | 80.58% | 70.97% | 58.5%
4.29 | 80.28% | 70.67% | 58.2%
4.3 | 79.88% | 70.37% | 57.8%
4.31 | 79.58% | 70.07% | 57.5%
4.32 | 79.28% | 69.67% | 57.2%
4.33 | 78.88% | 69.37% | 56.9%
4.34 | 78.58% | 69.07% | 56.6%
4.35 | 78.18% | 68.67% | 56.2%
4.36 | 77.88% | 68.37% | 55.9%
4.37 | 77.48% | 68.07% | 55.6%
4.38 | 77.18% | 67.67% | 55.3%
4.39 | 76.78% | 67.37% | 55%
4.4 | 76.48% | 66.97% | 54.7%
4.41 | 76.08% | 66.67% | 54.4%
4.42 | 75.78% | 66.37% | 54.1%
4.43 | 75.38% | 65.97% | 53.8%
4.44 | 74.97% | 65.67% | 53.5%
4.45 | 74.67% | 65.27% | 53.2%
4.46 | 74.27% | 64.96% | 52.9%
4.47 | 73.97% | 64.56% | 52.6%
4.48 | 73.57% | 64.26% | 52.3%
4.49 | 73.17% | 63.86% | 51.9%
4.5 | 72.87% | 63.56% | 51.6%
4.51 | 72.47% | 63.16% | 51.3%
4.52 | 72.07% | 62.86% | 51%
4.53 | 71.67% | 62.46% | 50.7%
4.54 | 71.27% | 62.06% | 50.3%
4.55 | 70.97% | 61.76% | 50%
4.56 | 70.57% | 61.36% | 49.7%
4.57 | 70.17% | 61.06% | 49.4%
4.58 | 69.77% | 60.66% | 49%
4.59 | 69.37% | 60.26% | 48.7%
4.6 | 68.97% | 59.96% | 48.4%
4.61 | 68.47% | 59.56% | 48.1%
4.62 | 68.07% | 59.16% | 47.7%
4.63 | 67.67% | 58.86% | 47.4%
4.64 | 67.17% | 58.46% | 47.1%
4.65 | 66.77% | 58.06% | 46.8%
4.66 | 66.37% | 57.66% | 46.5%
4.67 | 65.87% | 57.36% | 46.2%
4.68 | 65.47% | 56.96% | 45.9%
4.69 | 65.07% | 56.56% | 45.6%
4.7 | 64.56% | 56.16% | 45.3%
4.71 | 64.16% | 55.76% | 44.9%
4.72 | 63.76% | 55.46% | 44.6%
4.73 | 63.26% | 55.06% | 44.3%
4.74 | 62.86% | 54.65% | 44%
4.75 | 62.46% | 54.25% | 43.7%
4.76 | 61.96% | 53.85% | 43.3%
4.77 | 61.56% | 53.45% | 43%
4.78 | 61.06% | 53.15% | 42.7%
4.79 | 60.66% | 52.75% | 42.4%
4.8 | 60.16% | 52.35% | 42%
4.81 | 59.76% | 51.95% | 41.7%
4.82 | 59.36% | 51.55% | 41.4%
4.83 | 58.86% | 51.15% | 41%
4.84 | 58.36% | 50.75% | 40.7%
4.85 | 57.96% | 50.35% | 40.4%
4.86 | 57.46% | 50.05% | 40%
4.87 | 57.06% | 49.65% | 39.7%
4.88 | 56.56% | 49.25% | 39.4%
4.89 | 56.06% | 48.85% | 39%
4.9 | 55.66% | 48.45% | 38.7%
4.91 | 55.16% | 48.05% | 38.4%
4.92 | 54.75% | 47.65% | 38%
4.93 | 54.25% | 47.25% | 37.7%
4.94 | 53.75% | 46.85% | 37.4%
4.95 | 53.35% | 46.45% | 37.1%
4.96 | 52.85% | 46.15% | 36.7%
4.97 | 52.35% | 45.75% | 36.4%
4.98 | 51.95% | 45.35% | 36.1%
4.99 | 51.45% | 44.94% | 35.8%
5 | 50.95% | 44.54% | 35.5%
5.01 | 50.55% | 44.14% | 35.2%
5.02 | 50.05% | 43.74% | 34.9%
5.03 | 49.55% | 43.34% | 34.6%
5.04 | 49.05% | 42.94% | 34.3%
5.05 | 48.65% | 42.54% | 34%
5.06 | 48.15% | 42.14% | 33.7%
5.07 | 47.65% | 41.74% | 33.4%
5.08 | 47.25% | 41.34% | 33.1%
5.09 | 46.75% | 40.94% | 32.7%
5.1 | 46.25% | 40.54% | 32.4%
5.11 | 45.75% | 40.14% | 32.1%
5.12 | 45.35% | 39.74% | 31.8%
5.13 | 44.84% | 39.34% | 31.5%
5.14 | 44.34% | 38.94% | 31.2%
5.15 | 43.94% | 38.64% | 30.9%
5.16 | 43.44% | 38.24% | 30.5%
5.17 | 42.94% | 37.84% | 30.2%
5.18 | 42.54% | 37.44% | 29.9%
5.19 | 42.04% | 37.04% | 29.6%
5.2 | 41.54% | 36.64% | 29.3%
5.21 | 41.04% | 36.24% | 29%
5.22 | 40.64% | 35.84% | 28.7%
5.23 | 40.14% | 35.44% | 28.4%
5.24 | 39.64% | 35.04% | 28.1%
5.25 | 39.14% | 34.63% | 27.8%
5.26 | 38.64% | 34.23% | 27.5%
5.27 | 38.24% | 33.83% | 27.2%
5.28 | 37.74% | 33.43% | 26.9%
5.29 | 37.24% | 33.03% | 26.6%
5.3 | 36.74% | 32.63% | 26.3%
5.31 | 36.34% | 32.23% | 25.9%
5.32 | 35.84% | 31.93% | 25.6%
5.33 | 35.44% | 31.53% | 25.3%
5.34 | 34.93% | 31.13% | 25%
5.35 | 34.53% | 30.73% | 24.7%
5.36 | 34.03% | 30.33% | 24.4%
5.37 | 33.63% | 29.93% | 24.1%
5.38 | 33.13% | 29.53% | 23.8%
5.39 | 32.73% | 29.13% | 23.5%
5.4 | 32.33% | 28.73% | 23.2%
5.41 | 31.93% | 28.43% | 22.9%
5.42 | 31.43% | 28.03% | 22.6%
5.43 | 31.03% | 27.63% | 22.3%
5.44 | 30.63% | 27.23% | 22%
5.45 | 30.13% | 26.83% | 21.7%
5.46 | 29.73% | 26.53% | 21.4%
5.47 | 29.33% | 26.13% | 21.1%
5.48 | 28.83% | 25.73% | 20.8%
5.49 | 28.43% | 25.33% | 20.5%
5.5 | 27.93% | 25.03% | 20.2%
5.51 | 27.53% | 24.62% | 19.9%
5.52 | 27.13% | 24.22% | 19.5%
5.53 | 26.63% | 23.92% | 19.2%
5.54 | 26.23% | 23.52% | 18.9%
5.55 | 25.83% | 23.22% | 18.6%
5.56 | 25.33% | 22.82% | 18.3%
5.57 | 24.92% | 22.42% | 18.1%
5.58 | 24.52% | 22.12% | 17.8%
5.59 | 24.12% | 21.72% | 17.5%
5.6 | 23.72% | 21.42% | 17.3%
5.61 | 23.32% | 21.02% | 17%
5.62 | 22.92% | 20.72% | 16.7%
5.63 | 22.52% | 20.42% | 16.4%
5.64 | 22.12% | 20.02% | 16.2%
5.65 | 21.72% | 19.72% | 15.9%
5.66 | 21.32% | 19.32% | 15.6%
5.67 | 21.02% | 19.02% | 15.4%
5.68 | 20.62% | 18.72% | 15.1%
5.69 | 20.32% | 18.42% | 14.9%
5.7 | 19.92% | 18.02% | 14.7%
5.71 | 19.62% | 17.72% | 14.4%
5.72 | 19.22% | 17.42% | 14.2%
5.73 | 18.92% | 17.12% | 14%
5.74 | 18.52% | 16.82% | 13.7%
5.75 | 18.22% | 16.52% | 13.5%
5.76 | 17.82% | 16.12% | 13.3%
5.77 | 17.52% | 15.82% | 13.1%
5.78 | 17.12% | 15.52% | 12.8%
5.79 | 16.82% | 15.22% | 12.6%
5.8 | 16.52% | 14.91% | 12.4%
5.81 | 16.22% | 14.71% | 12.1%
5.82 | 15.82% | 14.41% | 11.9%
5.83 | 15.52% | 14.11% | 11.7%
5.84 | 15.22% | 13.81% | 11.5%
5.85 | 14.91% | 13.51% | 11.3%
5.86 | 14.61% | 13.31% | 11.1%
5.87 | 14.31% | 13.01% | 10.8%
5.88 | 14.01% | 12.71% | 10.6%
5.89 | 13.71% | 12.51% | 10.4%
5.9 | 13.41% | 12.21% | 10.2%
5.91 | 13.11% | 11.91% | 10%
5.92 | 12.91% | 11.71% | 9.8%
5.93 | 12.61% | 11.51% | 9.6%
5.94 | 12.31% | 11.21% | 9.4%
5.95 | 12.11% | 11.01% | 9.3%
5.96 | 11.81% | 10.71% | 9.1%
5.97 | 11.61% | 10.51% | 8.9%
5.98 | 11.31% | 10.31% | 8.7%
5.99 | 11.11% | 10.11% | 8.5%
6 | 10.81% | 9.81% | 8.4%
6.01 | 10.61% | 9.61% | 8.2%
6.02 | 10.31% | 9.409% | 8%
6.03 | 10.11% | 9.209% | 7.8%
6.04 | 9.91% | 9.009% | 7.7%
6.05 | 9.71% | 8.809% | 7.5%
6.06 | 9.51% | 8.609% | 7.4%
6.07 | 9.309% | 8.408% | 7.2%
6.08 | 9.009% | 8.208% | 7.1%
6.09 | 8.809% | 8.008% | 6.9%
6.1 | 8.609% | 7.908% | 6.8%
6.11 | 8.509% | 7.708% | 6.7%
6.12 | 8.308% | 7.508% | 6.5%
6.13 | 8.108% | 7.307% | 6.4%
6.14 | 7.908% | 7.207% | 6.2%
6.15 | 7.708% | 7.007% | 6.1%
6.16 | 7.508% | 6.807% | 6%
6.17 | 7.407% | 6.707% | 5.9%
6.18 | 7.207% | 6.507% | 5.7%
6.19 | 7.107% | 6.406% | 5.6%
6.2 | 6.907% | 6.306% | 5.5%
6.21 | 6.807% | 6.106% | 5.4%
6.22 | 6.607% | 6.006% | 5.3%
6.23 | 6.507% | 5.806% | 5.2%
6.24 | 6.306% | 5.706% | 5.1%
6.25 | 6.206% | 5.606% | 5%
6.26 | 6.006% | 5.506% | 4.9%
6.27 | 5.906% | 5.405% | 4.7%
6.28 | 5.806% | 5.205% | 4.6%
6.29 | 5.706% | 5.105% | 4.6%
6.3 | 5.606% | 5.005% | 4.5%
6.31 | 5.405% | 4.905% | 4.4%
6.32 | 5.305% | 4.805% | 4.3%
6.33 | 5.205% | 4.705% | 4.2%
6.34 | 5.105% | 4.605% | 4.2%
6.35 | 5.005% | 4.505% | 4.1%
6.36 | 4.905% | 4.404% | 4%
6.37 | 4.805% | 4.304% | 3.9%
6.38 | 4.705% | 4.204% | 3.9%
6.39 | 4.605% | 4.204% | 3.8%
6.4 | 4.505% | 4.104% | 3.7%
6.41 | 4.404% | 4.004% | 3.7%
6.42 | 4.404% | 3.904% | 3.6%
6.43 | 4.304% | 3.904% | 3.6%
6.44 | 4.204% | 3.804% | 3.5%
6.45 | 4.104% | 3.704% | 3.4%
6.46 | 4.004% | 3.704% | 3.4%
6.47 | 4.004% | 3.604% | 3.3%
6.48 | 3.904% | 3.504% | 3.3%
6.49 | 3.804% | 3.504% | 3.2%
6.5 | 3.804% | 3.403% | 3.2%
6.51 | 3.704% | 3.403% | 3.1%
6.52 | 3.704% | 3.303% | 3.1%
6.53 | 3.604% | 3.303% | 3%
6.54 | 3.604% | 3.203% | 3%
6.55 | 3.504% | 3.203% | 2.9%
6.56 | 3.504% | 3.103% | 2.9%
6.57 | 3.403% | 3.103% | 2.9%
6.58 | 3.403% | 3.003% | 2.8%
6.59 | 3.303% | 3.003% | 2.8%
6.6 | 3.303% | 2.903% | 2.8%
6.61 | 3.203% | 2.903% | 2.7%
6.62 | 3.203% | 2.903% | 2.7%
6.63 | 3.203% | 2.803% | 2.7%
6.64 | 3.103% | 2.803% | 2.7%
6.65 | 3.103% | 2.803% | 2.6%
6.66 | 3.103% | 2.703% | 2.6%
6.67 | 3.003% | 2.703% | 2.6%
6.68 | 3.003% | 2.703% | 2.6%
6.69 | 3.003% | 2.603% | 2.6%
6.7 | 3.003% | 2.603% | 2.5%
6.71 | 2.903% | 2.603% | 2.5%
6.72 | 2.903% | 2.603% | 2.5%
6.73 | 2.903% | 2.603% | 2.5%
6.74 | 2.903% | 2.503% | 2.5%
6.75 | 2.803% | 2.503% | 2.4%
6.76 | 2.803% | 2.503% | 2.4%
6.77 | 2.803% | 2.503% | 2.4%
6.78 | 2.803% | 2.503% | 2.4%
6.79 | 2.803% | 2.402% | 2.4%
6.8 | 2.803% | 2.402% | 2.4%
6.81 | 2.703% | 2.402% | 2.4%
6.82 | 2.703% | 2.402% | 2.4%
6.83 | 2.703% | 2.402% | 2.3%
6.84 | 2.703% | 2.402% | 2.3%
6.85 | 2.703% | 2.402% | 2.3%
6.86 | 2.703% | 2.302% | 2.3%
6.87 | 2.703% | 2.302% | 2.3%
6.88 | 2.603% | 2.302% | 2.3%
6.89 | 2.603% | 2.302% | 2.3%
6.9 | 2.603% | 2.302% | 2.3%
6.91 | 2.603% | 2.302% | 2.3%
6.92 | 2.603% | 2.302% | 2.3%
6.93 | 2.603% | 2.302% | 2.3%
6.94 | 2.603% | 2.302% | 2.2%
6.95 | 2.603% | 2.302% | 2.2%
6.96 | 2.603% | 2.302% | 2.2%
6.97 | 2.603% | 2.202% | 2.2%
6.98 | 2.603% | 2.202% | 2.2%
6.99 | 2.503% | 2.202% | 2.2%
7 | 2.503% | 2.202% | 2.2%
7.01 | 2.503% | 2.202% | 2.2%
7.02 | 2.503% | 2.202% | 2.2%
7.03 | 2.503% | 2.202% | 2.2%
7.04 | 2.503% | 2.202% | 2.2%
7.05 | 2.503% | 2.202% | 2.2%
7.06 | 2.503% | 2.202% | 2.2%
7.07 | 2.503% | 2.202% | 2.2%
7.08 | 2.503% | 2.202% | 2.2%
7.09 | 2.503% | 2.202% | 2.2%
7.1 | 2.503% | 2.202% | 2.2%
7.11 | 2.503% | 2.202% | 2.2%
7.12 | 2.503% | 2.202% | 2.2%
7.13 | 2.503% | 2.202% | 2.2%
7.14 | 2.503% | 2.202% | 2.2%
7.15 | 2.503% | 2.202% | 2.2%
7.16 | 2.503% | 2.202% | 2.2%
7.17 | 2.503% | 2.202% | 2.2%
7.18 | 2.503% | 2.202% | 2.2%
7.19 | 2.503% | 2.202% | 2.2%
7.2 | 2.503% | 2.202% | 2.2%
7.21 | 2.503% | 2.202% | 2.2%
7.22 | 2.503% | 2.202% | 2.2%
7.23 | 2.503% | 2.202% | 2.2%
7.24 | 2.503% | 2.102% | 2.1%
7.25 | 2.402% | 2.102% | 2.1%
7.26 | 2.402% | 2.102% | 2.1%
7.27 | 2.402% | 2.102% | 2.1%
7.28 | 2.402% | 2.102% | 2.1%
7.29 | 2.402% | 2.102% | 2.1%
7.3 | 2.402% | 2.102% | 2.1%
7.31 | 2.402% | 2.102% | 2.1%
7.32 | 2.402% | 2.102% | 2.1%
7.33 | 2.402% | 2.102% | 2.1%
7.34 | 2.402% | 2.102% | 2.1%
7.35 | 2.402% | 2.102% | 2.1%
7.36 | 2.402% | 2.102% | 2.1%
7.37 | 2.402% | 2.102% | 2.1%
7.38 | 2.402% | 2.102% | 2.1%
7.39 | 2.402% | 2.102% | 2.1%
7.4 | 2.402% | 2.102% | 2.1%
7.41 | 2.402% | 2.102% | 2.1%
7.42 | 2.402% | 2.102% | 2.1%
7.43 | 2.402% | 2.102% | 2.1%
7.44 | 2.402% | 2.102% | 2.1%
7.45 | 2.402% | 2.102% | 2.1%
7.46 | 2.402% | 2.102% | 2.1%
7.47 | 2.402% | 2.102% | 2.1%
7.48 | 2.402% | 2.102% | 2.1%
7.49 | 2.402% | 2.102% | 2.1%
7.5 | 2.402% | 2.102% | 2.1%
7.51 | 2.402% | 2.102% | 2.1%
7.52 | 2.402% | 2.102% | 2.1%
7.53 | 2.402% | 2.102% | 2.1%
7.54 | 2.402% | 2.102% | 2.1%
7.55 | 2.402% | 2.102% | 2.1%
7.56 | 2.402% | 2.102% | 2.1%
7.57 | 2.402% | 2.102% | 2.1%
7.58 | 2.402% | 2.102% | 2.1%
7.59 | 2.402% | 2.102% | 2.1%
7.6 | 2.402% | 2.102% | 2.1%
7.61 | 2.402% | 2.102% | 2.1%
7.62 | 2.402% | 2.002% | 2.1%
7.63 | 2.402% | 2.002% | 2.1%
7.64 | 2.402% | 2.002% | 2.1%
7.65 | 2.402% | 2.002% | 2.1%
7.66 | 2.402% | 2.002% | 2.1%
7.67 | 2.402% | 2.002% | 2.1%
7.68 | 2.302% | 2.002% | 2.1%
7.69 | 2.302% | 2.002% | 2%
7.7 | 2.302% | 2.002% | 2%
7.71 | 2.302% | 2.002% | 2%
7.72 | 2.302% | 2.002% | 2%
7.73 | 2.302% | 2.002% | 2%
7.74 | 2.302% | 2.002% | 2%
7.75 | 2.302% | 2.002% | 2%
7.76 | 2.302% | 2.002% | 2%
7.77 | 2.302% | 2.002% | 2%
7.78 | 2.302% | 2.002% | 2%
7.79 | 2.302% | 2.002% | 2%
7.8 | 2.302% | 2.002% | 2%
7.81 | 2.302% | 2.002% | 2%
7.82 | 2.302% | 2.002% | 2%
7.83 | 2.302% | 2.002% | 2%
7.84 | 2.302% | 2.002% | 2%
7.85 | 2.302% | 2.002% | 2%
7.86 | 2.302% | 2.002% | 2%
7.87 | 2.302% | 2.002% | 2%
7.88 | 2.302% | 2.002% | 2%
7.89 | 2.302% | 2.002% | 2%
7.9 | 2.302% | 2.002% | 2%
7.91 | 2.302% | 2.002% | 2%
7.92 | 2.302% | 2.002% | 2%
7.93 | 2.302% | 2.002% | 2%
7.94 | 2.302% | 2.002% | 2%
7.95 | 2.302% | 2.002% | 2%
7.96 | 2.302% | 2.002% | 2%
7.97 | 2.302% | 2.002% | 2%
7.98 | 2.302% | 2.002% | 2%
7.99 | 2.302% | 2.002% | 2%
8 | 2.302% | 2.002% | 2%
8.01 | 2.302% | 2.002% | 2%
8.02 | 2.302% | 2.002% | 2%
8.03 | 2.302% | 2.002% | 2%
8.04 | 2.302% | 2.002% | 2%
8.05 | 2.302% | 2.002% | 2%
8.06 | 2.302% | 2.002% | 2%
8.07 | 2.302% | 2.002% | 2%
8.08 | 2.302% | 2.002% | 2%
8.09 | 2.302% | 2.002% | 2%
8.1 | 2.302% | 2.002% | 2%
8.11 | 2.302% | 2.002% | 2%
8.12 | 2.302% | 2.002% | 2%
8.13 | 2.302% | 2.002% | 2%
8.14 | 2.302% | 2.002% | 2%
8.15 | 2.302% | 2.002% | 2%
8.16 | 2.302% | 2.002% | 2%
8.17 | 2.302% | 1.902% | 2%
8.18 | 2.302% | 1.902% | 2%
8.19 | 2.302% | 1.902% | 2%
8.2 | 2.302% | 1.902% | 2%
8.21 | 2.302% | 1.902% | 2%
8.22 | 2.302% | 1.902% | 2%
8.23 | 2.302% | 1.902% | 2%
8.24 | 2.302% | 1.902% | 2%
8.25 | 2.302% | 1.902% | 1.9%
8.26 | 2.302% | 1.902% | 1.9%
8.27 | 2.302% | 1.902% | 1.9%
8.28 | 2.302% | 1.902% | 1.9%
8.29 | 2.302% | 1.902% | 1.9%
8.3 | 2.302% | 1.902% | 1.9%
8.31 | 2.302% | 1.902% | 1.9%
8.32 | 2.302% | 1.902% | 1.9%
8.33 | 2.302% | 1.902% | 1.9%
8.34 | 2.302% | 1.902% | 1.9%
8.35 | 2.202% | 1.902% | 1.9%
8.36 | 2.202% | 1.902% | 1.9%
8.37 | 2.202% | 1.902% | 1.9%
8.38 | 2.202% | 1.902% | 1.9%
8.39 | 2.202% | 1.902% | 1.9%
8.4 | 2.202% | 1.902% | 1.9%
8.41 | 2.202% | 1.902% | 1.9%
8.42 | 2.202% | 1.902% | 1.9%
8.43 | 2.202% | 1.902% | 1.9%
8.44 | 2.202% | 1.902% | 1.9%
8.45 | 2.202% | 1.902% | 1.9%
8.46 | 2.202% | 1.902% | 1.9%
8.47 | 2.202% | 1.902% | 1.9%
8.48 | 2.202% | 1.902% | 1.9%
8.49 | 2.202% | 1.902% | 1.9%
8.5 | 2.202% | 1.902% | 1.9%
8.51 | 2.202% | 1.902% | 1.9%
8.52 | 2.202% | 1.902% | 1.9%
8.53 | 2.202% | 1.902% | 1.9%
8.54 | 2.202% | 1.902% | 1.9%
8.55 | 2.202% | 1.902% | 1.9%
8.56 | 2.202% | 1.902% | 1.9%
8.57 | 2.202% | 1.902% | 1.9%
8.58 | 2.202% | 1.902% | 1.9%
8.59 | 2.202% | 1.902% | 1.9%
8.6 | 2.202% | 1.902% | 1.9%
8.61 | 2.202% | 1.902% | 1.9%
8.62 | 2.202% | 1.902% | 1.9%
8.63 | 2.202% | 1.902% | 1.9%
8.64 | 2.202% | 1.902% | 1.9%
8.65 | 2.202% | 1.902% | 1.9%
8.66 | 2.202% | 1.902% | 1.9%
8.67 | 2.202% | 1.902% | 1.9%
8.68 | 2.202% | 1.902% | 1.9%
8.69 | 2.202% | 1.902% | 1.9%
8.7 | 2.202% | 1.902% | 1.9%
8.71 | 2.202% | 1.902% | 1.9%
8.72 | 2.202% | 1.902% | 1.9%
8.73 | 2.202% | 1.902% | 1.9%
8.74 | 2.202% | 1.902% | 1.9%
8.75 | 2.202% | 1.902% | 1.9%
8.76 | 2.202% | 1.902% | 1.9%
8.77 | 2.202% | 1.902% | 1.9%
8.78 | 2.202% | 1.902% | 1.9%
8.79 | 2.202% | 1.902% | 1.9%
8.8 | 2.202% | 1.902% | 1.9%
8.81 | 2.202% | 1.902% | 1.9%
8.82 | 2.202% | 1.902% | 1.9%
8.83 | 2.202% | 1.902% | 1.9%
8.84 | 2.202% | 1.802% | 1.9%
8.85 | 2.202% | 1.802% | 1.9%
8.86 | 2.202% | 1.802% | 1.9%
8.87 | 2.202% | 1.802% | 1.9%
8.88 | 2.202% | 1.802% | 1.9%
8.89 | 2.202% | 1.802% | 1.9%
8.9 | 2.202% | 1.802% | 1.9%
8.91 | 2.202% | 1.802% | 1.9%
8.92 | 2.202% | 1.802% | 1.9%
8.93 | 2.202% | 1.802% | 1.9%
8.94 | 2.202% | 1.802% | 1.9%
8.95 | 2.102% | 1.802% | 1.9%
8.96 | 2.102% | 1.802% | 1.8%
8.97 | 2.102% | 1.802% | 1.8%
8.98 | 2.102% | 1.802% | 1.8%
8.99 | 2.102% | 1.802% | 1.8%
9 | 2.102% | 1.802% | 1.8%
9.01 | 2.102% | 1.802% | 1.8%
9.02 | 2.102% | 1.802% | 1.8%
9.03 | 2.102% | 1.802% | 1.8%
9.04 | 2.102% | 1.802% | 1.8%
9.05 | 2.102% | 1.802% | 1.8%
9.06 | 2.102% | 1.802% | 1.8%
9.07 | 2.102% | 1.802% | 1.8%
9.08 | 2.102% | 1.802% | 1.8%
9.09 | 2.102% | 1.802% | 1.8%
9.1 | 2.102% | 1.802% | 1.8%
9.11 | 2.102% | 1.802% | 1.8%
9.12 | 2.102% | 1.802% | 1.8%
9.13 | 2.102% | 1.802% | 1.8%
9.14 | 2.102% | 1.802% | 1.8%
9.15 | 2.102% | 1.802% | 1.8%
9.16 | 2.102% | 1.802% | 1.8%
9.17 | 2.102% | 1.802% | 1.8%
9.18 | 2.102% | 1.802% | 1.8%
9.19 | 2.102% | 1.802% | 1.8%
9.2 | 2.102% | 1.802% | 1.8%
9.21 | 2.102% | 1.802% | 1.8%
9.22 | 2.102% | 1.802% | 1.8%
9.23 | 2.102% | 1.802% | 1.8%
9.24 | 2.102% | 1.802% | 1.8%
9.25 | 2.102% | 1.802% | 1.8%
9.26 | 2.102% | 1.802% | 1.8%
9.27 | 2.102% | 1.802% | 1.8%
9.28 | 2.102% | 1.802% | 1.8%
9.29 | 2.102% | 1.802% | 1.8%
9.3 | 2.102% | 1.802% | 1.8%
9.31 | 2.102% | 1.802% | 1.8%
9.32 | 2.102% | 1.802% | 1.8%
9.33 | 2.102% | 1.802% | 1.8%
9.34 | 2.102% | 1.802% | 1.8%
9.35 | 2.102% | 1.802% | 1.8%
9.36 | 2.102% | 1.802% | 1.8%
9.37 | 2.102% | 1.802% | 1.8%
9.38 | 2.102% | 1.802% | 1.8%
9.39 | 2.102% | 1.802% | 1.8%
9.4 | 2.102% | 1.802% | 1.8%
9.41 | 2.102% | 1.802% | 1.8%
9.42 | 2.102% | 1.802% | 1.8%
9.43 | 2.102% | 1.802% | 1.8%
9.44 | 2.102% | 1.802% | 1.8%
9.45 | 2.102% | 1.802% | 1.8%
9.46 | 2.102% | 1.802% | 1.8%
9.47 | 2.102% | 1.802% | 1.8%
9.48 | 2.102% | 1.702% | 1.8%
9.49 | 2.102% | 1.702% | 1.8%
9.5 | 2.102% | 1.702% | 1.8%
9.51 | 2.102% | 1.702% | 1.8%
9.52 | 2.102% | 1.702% | 1.8%
9.53 | 2.102% | 1.702% | 1.8%
9.54 | 2.102% | 1.702% | 1.8%
9.55 | 2.102% | 1.702% | 1.8%
9.56 | 2.102% | 1.702% | 1.8%
9.57 | 2.102% | 1.702% | 1.8%
9.58 | 2.102% | 1.702% | 1.7%
9.59 | 2.102% | 1.702% | 1.7%
9.6 | 2.102% | 1.702% | 1.7%
9.61 | 2.102% | 1.702% | 1.7%
9.62 | 2.002% | 1.702% | 1.7%
9.63 | 2.002% | 1.702% | 1.7%
9.64 | 2.002% | 1.702% | 1.7%
9.65 | 2.002% | 1.702% | 1.7%
9.66 | 2.002% | 1.702% | 1.7%
9.67 | 2.002% | 1.702% | 1.7%
9.68 | 2.002% | 1.702% | 1.7%
9.69 | 2.002% | 1.702% | 1.7%
9.7 | 2.002% | 1.702% | 1.7%
9.71 | 2.002% | 1.702% | 1.7%
9.72 | 2.002% | 1.702% | 1.7%
9.73 | 2.002% | 1.702% | 1.7%
9.74 | 2.002% | 1.702% | 1.7%
9.75 | 2.002% | 1.702% | 1.7%
9.76 | 2.002% | 1.702% | 1.7%
9.77 | 2.002% | 1.702% | 1.7%
9.78 | 2.002% | 1.702% | 1.7%
9.79 | 2.002% | 1.702% | 1.7%
9.8 | 2.002% | 1.702% | 1.7%
9.81 | 2.002% | 1.702% | 1.7%
9.82 | 2.002% | 1.702% | 1.7%
9.83 | 2.002% | 1.702% | 1.7%
9.84 | 2.002% | 1.702% | 1.7%
9.85 | 2.002% | 1.702% | 1.7%
9.86 | 2.002% | 1.702% | 1.7%
9.87 | 2.002% | 1.702% | 1.7%
9.88 | 2.002% | 1.702% | 1.7%
9.89 | 2.002% | 1.702% | 1.7%
9.9 | 2.002% | 1.702% | 1.7%
9.91 | 2.002% | 1.702% | 1.7%
9.92 | 2.002% | 1.702% | 1.7%
9.93 | 2.002% | 1.702% | 1.7%
9.94 | 2.002% | 1.702% | 1.7%
9.95 | 2.002% | 1.702% | 1.7%
9.96 | 2.002% | 1.702% | 1.7%
9.97 | 2.002% | 1.702% | 1.7%
9.98 | 2.002% | 1.702% | 1.7%
9.99 | 2.002% | 1.702% | 1.7%
10 | 2.002% | 1.702% | 1.7%
10.01 | 2.002% | 1.702% | 1.7%
10.02 | 2.002% | 1.702% | 1.7%
10.03 | 2.002% | 1.702% | 1.7%
10.04 | 2.002% | 1.702% | 1.7%
10.05 | 2.002% | 1.702% | 1.7%
10.06 | 2.002% | 1.702% | 1.7%
10.07 | 2.002% | 1.702% | 1.7%
10.08 | 2.002% | 1.702% | 1.7%
10.09 | 2.002% | 1.702% | 1.7%
10.1 | 2.002% | 1.702% | 1.7%
10.11 | 2.002% | 1.702% | 1.7%
10.12 | 2.002% | 1.702% | 1.7%
10.13 | 2.002% | 1.702% | 1.7%
10.14 | 2.002% | 1.702% | 1.7%
10.15 | 2.002% | 1.702% | 1.7%
10.16 | 2.002% | 1.702% | 1.7%
10.17 | 2.002% | 1.602% | 1.7%
10.18 | 2.002% | 1.602% | 1.7%
10.19 | 2.002% | 1.602% | 1.7%
10.2 | 2.002% | 1.602% | 1.7%
10.21 | 2.002% | 1.602% | 1.7%
10.22 | 2.002% | 1.602% | 1.7%
10.23 | 2.002% | 1.602% | 1.7%
10.24 | 2.002% | 1.602% | 1.7%
10.25 | 2.002% | 1.602% | 1.7%
10.26 | 2.002% | 1.602% | 1.7%
10.27 | 2.002% | 1.602% | 1.7%
10.28 | 2.002% | 1.602% | 1.7%
10.29 | 2.002% | 1.602% | 1.7%
10.3 | 2.002% | 1.602% | 1.7%
10.31 | 2.002% | 1.602% | 1.7%
10.32 | 2.002% | 1.602% | 1.7%
10.33 | 1.902% | 1.602% | 1.7%
10.34 | 1.902% | 1.602% | 1.7%
10.35 | 1.902% | 1.602% | 1.7%
10.36 | 1.902% | 1.602% | 1.7%
10.37 | 1.902% | 1.602% | 1.7%
10.38 | 1.902% | 1.602% | 1.6%
10.39 | 1.902% | 1.602% | 1.6%
10.4 | 1.902% | 1.602% | 1.6%
10.41 | 1.902% | 1.602% | 1.6%
10.42 | 1.902% | 1.602% | 1.6%
10.43 | 1.902% | 1.602% | 1.6%
10.44 | 1.902% | 1.602% | 1.6%
10.45 | 1.902% | 1.602% | 1.6%
10.46 | 1.902% | 1.602% | 1.6%
10.47 | 1.902% | 1.602% | 1.6%
10.48 | 1.902% | 1.602% | 1.6%
10.49 | 1.902% | 1.602% | 1.6%
10.5 | 1.902% | 1.602% | 1.6%
10.51 | 1.902% | 1.602% | 1.6%
10.52 | 1.902% | 1.602% | 1.6%
10.53 | 1.902% | 1.602% | 1.6%
10.54 | 1.902% | 1.602% | 1.6%
10.55 | 1.902% | 1.602% | 1.6%
10.56 | 1.902% | 1.602% | 1.6%
10.57 | 1.902% | 1.602% | 1.6%
10.58 | 1.902% | 1.602% | 1.6%
10.59 | 1.902% | 1.602% | 1.6%
10.6 | 1.902% | 1.602% | 1.6%
10.61 | 1.902% | 1.602% | 1.6%
10.62 | 1.902% | 1.602% | 1.6%
10.63 | 1.902% | 1.602% | 1.6%
10.64 | 1.902% | 1.602% | 1.6%
10.65 | 1.902% | 1.602% | 1.6%
10.66 | 1.902% | 1.602% | 1.6%
10.67 | 1.902% | 1.602% | 1.6%
10.68 | 1.902% | 1.602% | 1.6%
10.69 | 1.902% | 1.602% | 1.6%
10.7 | 1.902% | 1.602% | 1.6%
10.71 | 1.902% | 1.602% | 1.6%
10.72 | 1.902% | 1.602% | 1.6%
10.73 | 1.902% | 1.602% | 1.6%
10.74 | 1.902% | 1.602% | 1.6%
10.75 | 1.902% | 1.602% | 1.6%
10.76 | 1.902% | 1.602% | 1.6%
10.77 | 1.902% | 1.602% | 1.6%
10.78 | 1.902% | 1.602% | 1.6%
10.79 | 1.902% | 1.602% | 1.6%
10.8 | 1.902% | 1.602% | 1.6%
10.81 | 1.902% | 1.602% | 1.6%
10.82 | 1.902% | 1.602% | 1.6%
10.83 | 1.902% | 1.602% | 1.6%
10.84 | 1.902% | 1.602% | 1.6%
10.85 | 1.902% | 1.602% | 1.6%
10.86 | 1.902% | 1.602% | 1.6%
10.87 | 1.902% | 1.602% | 1.6%
10.88 | 1.902% | 1.602% | 1.6%
10.89 | 1.902% | 1.602% | 1.6%
10.9 | 1.902% | 1.602% | 1.6%
10.91 | 1.902% | 1.602% | 1.6%
10.92 | 1.902% | 1.602% | 1.6%
10.93 | 1.902% | 1.602% | 1.6%
10.94 | 1.902% | 1.602% | 1.6%
10.95 | 1.902% | 1.602% | 1.6%
10.96 | 1.902% | 1.602% | 1.6%
10.97 | 1.902% | 1.602% | 1.6%
10.98 | 1.902% | 1.602% | 1.6%
10.99 | 1.902% | 1.602% | 1.6%
11 | 1.902% | 1.602% | 1.6%
11.01 | 1.902% | 1.602% | 1.6%
11.02 | 1.902% | 1.602% | 1.6%
11.03 | 1.902% | 1.602% | 1.6%
11.04 | 1.902% | 1.602% | 1.6%
11.05 | 1.902% | 1.602% | 1.6%
11.06 | 1.902% | 1.602% | 1.6%
11.07 | 1.902% | 1.502% | 1.6%
11.08 | 1.902% | 1.502% | 1.6%
11.09 | 1.902% | 1.502% | 1.6%
11.1 | 1.902% | 1.502% | 1.6%
11.11 | 1.902% | 1.502% | 1.6%
11.12 | 1.902% | 1.502% | 1.6%
11.13 | 1.802% | 1.502% | 1.6%
11.14 | 1.802% | 1.502% | 1.6%
11.15 | 1.802% | 1.502% | 1.6%
11.16 | 1.802% | 1.502% | 1.6%
11.17 | 1.802% | 1.502% | 1.6%
11.18 | 1.802% | 1.502% | 1.6%
11.19 | 1.802% | 1.502% | 1.5%
11.2 | 1.802% | 1.502% | 1.5%
11.21 | 1.802% | 1.502% | 1.5%
11.22 | 1.802% | 1.502% | 1.5%
11.23 | 1.802% | 1.502% | 1.5%
11.24 | 1.802% | 1.502% | 1.5%
11.25 | 1.802% | 1.502% | 1.5%
11.26 | 1.802% | 1.502% | 1.5%
11.27 | 1.802% | 1.502% | 1.5%
11.28 | 1.802% | 1.502% | 1.5%
11.29 | 1.802% | 1.502% | 1.5%
11.3 | 1.802% | 1.502% | 1.5%
11.31 | 1.802% | 1.502% | 1.5%
11.32 | 1.802% | 1.502% | 1.5%
11.33 | 1.802% | 1.502% | 1.5%
11.34 | 1.802% | 1.502% | 1.5%
11.35 | 1.802% | 1.502% | 1.5%
11.36 | 1.802% | 1.502% | 1.5%
11.37 | 1.802% | 1.502% | 1.5%
11.38 | 1.802% | 1.502% | 1.5%
11.39 | 1.802% | 1.502% | 1.5%
11.4 | 1.802% | 1.502% | 1.5%
11.41 | 1.802% | 1.502% | 1.5%
11.42 | 1.802% | 1.502% | 1.5%
11.43 | 1.802% | 1.502% | 1.5%
11.44 | 1.802% | 1.502% | 1.5%
11.45 | 1.802% | 1.502% | 1.5%
11.46 | 1.802% | 1.502% | 1.5%
11.47 | 1.802% | 1.502% | 1.5%
11.48 | 1.802% | 1.502% | 1.5%
11.49 | 1.802% | 1.502% | 1.5%
11.5 | 1.802% | 1.502% | 1.5%
11.51 | 1.802% | 1.502% | 1.5%
11.52 | 1.802% | 1.502% | 1.5%
11.53 | 1.802% | 1.502% | 1.5%
11.54 | 1.802% | 1.502% | 1.5%
11.55 | 1.802% | 1.502% | 1.5%
11.56 | 1.802% | 1.502% | 1.5%
11.57 | 1.802% | 1.502% | 1.5%
11.58 | 1.802% | 1.502% | 1.5%
11.59 | 1.802% | 1.502% | 1.5%
11.6 | 1.802% | 1.502% | 1.5%
11.61 | 1.802% | 1.502% | 1.5%
11.62 | 1.802% | 1.502% | 1.5%
11.63 | 1.802% | 1.502% | 1.5%
11.64 | 1.802% | 1.502% | 1.5%
11.65 | 1.802% | 1.502% | 1.5%
11.66 | 1.802% | 1.502% | 1.5%
11.67 | 1.802% | 1.502% | 1.5%
11.68 | 1.802% | 1.502% | 1.5%
11.69 | 1.802% | 1.502% | 1.5%
11.7 | 1.802% | 1.502% | 1.5%
11.71 | 1.802% | 1.502% | 1.5%
11.72 | 1.802% | 1.502% | 1.5%
11.73 | 1.802% | 1.502% | 1.5%
11.74 | 1.802% | 1.502% | 1.5%
11.75 | 1.802% | 1.502% | 1.5%
11.76 | 1.802% | 1.502% | 1.5%
11.77 | 1.802% | 1.502% | 1.5%
11.78 | 1.802% | 1.502% | 1.5%
11.79 | 1.802% | 1.502% | 1.5%
11.8 | 1.802% | 1.502% | 1.5%
11.81 | 1.802% | 1.502% | 1.5%
11.82 | 1.802% | 1.502% | 1.5%
|  | 1.502% | 1.5%
|  | 1.502% | 1.5%
11.86 | 5.506% | 1.502% | 1.5%
11.87 | 5.506% | 1.502%
11.88 | 5.506% | 1.502%
11.89 | 5.506% | 1.502% | 0.048
11.9 | 5.506% | 1.502% | 0.048
11.91 | 5.506% |  | 0.048
11.92 | 5.506% |  | 0.048
11.93 | 5.506% | 0.052052 | 0.048
11.94 | 5.506% | 0.052052 | 0.048
11.95 | 5.506% | 0.052052 | 0.048
11.96 | 5.506% | 0.052052 | 0.048
11.97 | 5.506% | 0.051051 | 0.048
11.98 | 5.506% | 0.051051 | 0.048
11.99 | 5.506% | 0.051051 | 0.048
12 | 5.506% | 0.051051 | 0.048
12.01 | 5.506% | 0.051051 | 0.048
12.02 | 5.506% | 0.051051 | 0.048
12.03 | 5.506% | 0.051051 | 0.048
12.04 | 5.506% | 0.051051 | 0.048
12.05 | 5.506% | 0.051051 | 0.048
12.06 | 5.506% | 0.051051 | 0.048
12.07 | 5.506% | 0.051051 | 0.048
12.08 | 5.405% | 0.051051 | 0.048
12.09 | 5.405% | 0.051051 | 0.048
12.1 | 5.405% | 0.051051 | 0.048
12.11 | 5.405% | 0.051051 | 0.048
12.12 | 5.405% | 0.051051 | 0.047
12.13 | 5.405% | 0.051051 | 0.047
12.14 | 5.405% | 0.051051 | 0.047
12.15 | 5.405% | 0.051051 | 0.047
12.16 | 5.405% | 0.051051 | 0.047
12.17 | 5.405% | 0.051051 | 0.047
12.18 | 5.405% | 0.051051 | 0.047
12.19 | 5.405% | 0.051051 | 0.047
12.2 | 5.405% | 0.051051 | 0.047
12.21 | 5.405% | 0.051051 | 0.047
12.22 | 5.405% | 0.051051 | 0.047
12.23 | 5.405% | 0.051051 | 0.047
12.24 | 5.405% | 0.05005 | 0.047
12.25 | 5.405% | 0.05005 | 0.047
12.26 | 5.405% | 0.05005 | 0.047
12.27 | 5.405% | 0.05005 | 0.047
12.28 | 5.405% | 0.05005 | 0.047
12.29 | 5.405% | 0.05005 | 0.047
12.3 | 5.405% | 0.05005 | 0.047
12.31 | 5.405% | 0.05005 | 0.047
12.32 | 5.405% | 0.05005 | 0.047
12.33 | 5.405% | 0.05005 | 0.047
12.34 | 5.405% | 0.05005 | 0.047
12.35 | 5.405% | 0.05005 | 0.047
12.36 | 5.405% | 0.05005 | 0.047
12.37 | 5.405% | 0.05005 | 0.047
12.38 | 5.405% | 0.05005 | 0.047
12.39 | 5.405% | 0.05005 | 0.047
12.4 | 5.305% | 0.05005 | 0.047
12.41 | 5.305% | 0.05005 | 0.046
12.42 | 5.305% | 0.05005 | 0.046
12.43 | 5.305% | 0.05005 | 0.046
12.44 | 5.305% | 0.05005 | 0.046
12.45 | 5.305% | 0.05005 | 0.046
12.46 | 5.305% | 0.05005 | 0.046
12.47 | 5.305% | 0.05005 | 0.046
12.48 | 5.305% | 0.05005 | 0.046
12.49 | 5.305% | 0.05005 | 0.046
12.5 | 5.305% | 0.05005 | 0.046
12.51 | 5.305% | 0.05005 | 0.046
12.52 | 5.305% | 0.05005 | 0.046
12.53 | 5.305% | 0.05005 | 0.046
12.54 | 5.305% | 0.05005 | 0.046
12.55 | 5.305% | 0.05005 | 0.046
12.56 | 5.305% | 0.05005 | 0.046
12.57 | 5.305% | 0.05005 | 0.046
12.58 | 5.305% | 0.05005 | 0.046
12.59 | 5.305% | 0.05005 | 0.046
12.6 | 5.305% | 0.05005 | 0.046
12.61 | 5.305% | 0.049049 | 0.046
12.62 | 5.305% | 0.049049 | 0.046
12.63 | 5.305% | 0.049049 | 0.046
12.64 | 5.305% | 0.049049 | 0.046
12.65 | 5.305% | 0.049049 | 0.046
12.66 | 5.305% | 0.049049 | 0.046
12.67 | 5.305% | 0.049049 | 0.046
12.68 | 5.305% | 0.049049 | 0.046
12.69 | 5.305% | 0.049049 | 0.046
12.7 | 5.305% | 0.049049 | 0.046
12.71 | 5.305% | 0.049049 | 0.046
12.72 | 5.305% | 0.049049 | 0.046
12.73 | 5.305% | 0.049049 | 0.046
12.74 | 5.305% | 0.049049 | 0.046
12.75 | 5.305% | 0.049049 | 0.046
12.76 | 5.305% | 0.049049 | 0.046
12.77 | 5.305% | 0.049049 | 0.046
12.78 | 5.305% | 0.049049 | 0.046
12.79 | 5.305% | 0.049049 | 0.046
12.8 | 5.305% | 0.049049 | 0.045
12.81 | 5.305% | 0.049049 | 0.045
12.82 | 5.205% | 0.049049 | 0.045
12.83 | 5.205% | 0.049049 | 0.045
12.84 | 5.205% | 0.049049 | 0.045
12.85 | 5.205% | 0.049049 | 0.045
12.86 | 5.205% | 0.048048 | 0.045
12.87 | 5.205% | 0.048048 | 0.045
12.88 | 5.205% | 0.048048 | 0.045
12.89 | 5.205% | 0.048048 | 0.045
12.9 | 5.205% | 0.048048 | 0.045
12.91 | 5.205% | 0.048048 | 0.045
12.92 | 5.205% | 0.048048 | 0.045
12.93 | 5.205% | 0.048048 | 0.045
12.94 | 5.205% | 0.048048 | 0.045
12.95 | 5.205% | 0.048048 | 0.045
12.96 | 5.205% | 0.048048 | 0.045
12.97 | 5.205% | 0.048048 | 0.045
12.98 | 5.205% | 0.048048 | 0.045
12.99 | 5.205% | 0.048048 | 0.045
13 | 5.205% | 0.048048 | 0.045
13.01 | 5.205% | 0.048048 | 0.045
13.02 | 5.205% | 0.048048 | 0.045
13.03 | 5.205% | 0.048048 | 0.045
13.04 | 5.205% | 0.048048 | 0.045
13.05 | 5.205% | 0.048048 | 0.045
13.06 | 5.205% | 0.048048 | 0.045
13.07 | 5.205% | 0.048048 | 0.045
13.08 | 5.205% | 0.048048 | 0.045
13.09 | 5.205% | 0.048048 | 0.045
13.1 | 5.205% | 0.048048 | 0.045
13.11 | 5.205% | 0.047047 | 0.045
13.12 | 5.105% | 0.047047 | 0.044
13.13 | 5.105% | 0.047047 | 0.044
13.14 | 5.105% | 0.047047 | 0.044
13.15 | 5.105% | 0.047047 | 0.044
13.16 | 5.105% | 0.047047 | 0.044
13.17 | 5.105% | 0.047047 | 0.044
13.18 | 5.105% | 0.047047 | 0.044
13.19 | 5.105% | 0.047047 | 0.044
13.2 | 5.105% | 0.047047 | 0.044
13.21 | 5.105% | 0.047047 | 0.044
13.22 | 5.105% | 0.047047 | 0.044
13.23 | 5.105% | 0.047047 | 0.044
13.24 | 5.105% | 0.047047 | 0.044
13.25 | 5.105% | 0.047047 | 0.044
13.26 | 5.105% | 0.047047 | 0.044
13.27 | 5.105% | 0.047047 | 0.044
13.28 | 5.105% | 0.047047 | 0.044
13.29 | 5.105% | 0.047047 | 0.044
13.3 | 5.105% | 0.047047 | 0.044
13.31 | 5.105% | 0.047047 | 0.044
13.32 | 5.105% | 0.047047 | 0.044
13.33 | 5.105% | 0.047047 | 0.044
13.34 | 5.105% | 0.047047 | 0.044
13.35 | 5.105% | 0.047047 | 0.044
13.36 | 5.105% | 0.047047 | 0.044
13.37 | 5.105% | 0.047047 | 0.044
13.38 | 5.105% | 0.047047 | 0.044
13.39 | 5.105% | 0.047047 | 0.044
13.4 | 5.105% | 0.047047 | 0.044
13.41 | 5.105% | 0.047047 | 0.044
13.42 | 5.105% | 0.047047 | 0.043
13.43 | 5.105% | 0.047047 | 0.043
13.44 | 5.105% | 0.047047 | 0.043
13.45 | 5.105% | 0.047047 | 0.043
13.46 | 5.105% | 0.047047 | 0.043
13.47 | 5.105% | 0.047047 | 0.043
13.48 | 5.105% | 0.047047 | 0.043
13.49 | 5.105% | 0.047047 | 0.043
13.5 | 5.105% | 0.047047 | 0.043
13.51 | 5.105% | 0.047047 | 0.043
13.52 | 5.005% | 0.047047 | 0.043
13.53 | 5.005% | 0.047047 | 0.043
13.54 | 5.005% | 0.047047 | 0.043
13.55 | 5.005% | 0.047047 | 0.043
13.56 | 5.005% | 0.047047 | 0.043
13.57 | 5.005% | 0.047047 | 0.043
13.58 | 5.005% | 0.047047 | 0.043
13.59 | 5.005% | 0.047047 | 0.043
13.6 | 5.005% | 0.047047 | 0.043
13.61 | 5.005% | 0.047047 | 0.043
13.62 | 5.005% | 0.047047 | 0.043
13.63 | 5.005% | 0.047047 | 0.043
13.64 | 5.005% | 0.047047 | 0.043
13.65 | 5.005% | 0.047047 | 0.043
13.66 | 5.005% | 0.047047 | 0.043
13.67 | 5.005% | 0.046046 | 0.043
13.68 | 5.005% | 0.046046 | 0.043
13.69 | 5.005% | 0.046046 | 0.043
13.7 | 5.005% | 0.046046 | 0.043
13.71 | 5.005% | 0.046046 | 0.043
13.72 | 5.005% | 0.046046 | 0.043
13.73 | 5.005% | 0.046046 | 0.043
13.74 | 5.005% | 0.046046 | 0.043
13.75 | 5.005% | 0.046046 | 0.043
13.76 | 5.005% | 0.046046 | 0.043
13.77 | 5.005% | 0.046046 | 0.043
13.78 | 5.005% | 0.046046 | 0.043
13.79 | 5.005% | 0.046046 | 0.043
13.8 | 5.005% | 0.046046 | 0.043
13.81 | 5.005% | 0.046046 | 0.043
13.82 | 5.005% | 0.046046 | 0.043
13.83 | 5.005% | 0.046046 | 0.043
13.84 | 5.005% | 0.046046 | 0.043
13.85 | 5.005% | 0.046046 | 0.043
13.86 | 5.005% | 0.046046 | 0.043
13.87 | 5.005% | 0.046046 | 0.043
13.88 | 5.005% | 0.046046 | 0.043
13.89 | 5.005% | 0.046046 | 0.043
13.9 | 5.005% | 0.046046 | 0.043
13.91 | 4.905% | 0.046046 | 0.043
13.92 | 4.905% | 0.046046 | 0.043
13.93 | 4.905% | 0.046046 | 0.043
13.94 | 4.905% | 0.046046 | 0.043
13.95 | 4.905% | 0.046046 | 0.043
13.96 | 4.905% | 0.046046 | 0.042
13.97 | 4.905% | 0.046046 | 0.042
13.98 | 4.905% | 0.046046 | 0.042
13.99 | 4.905% | 0.046046 | 0.042
14 | 4.905% | 0.046046 | 0.042
14.01 | 4.905% | 0.046046 | 0.042
14.02 | 4.905% | 0.046046 | 0.042
14.03 | 4.905% | 0.046046 | 0.042
14.04 | 4.905% | 0.045045 | 0.042
14.05 | 4.905% | 0.045045 | 0.042
14.06 | 4.905% | 0.045045 | 0.042
14.07 | 4.905% | 0.045045 | 0.042
14.08 | 4.905% | 0.045045 | 0.042
14.09 | 4.905% | 0.045045 | 0.042
14.1 | 4.905% | 0.045045 | 0.042
14.11 | 4.905% | 0.045045 | 0.042
14.12 | 4.905% | 0.045045 | 0.042
14.13 | 4.905% | 0.045045 | 0.042
14.14 | 4.905% | 0.045045 | 0.042
14.15 | 4.905% | 0.045045 | 0.042
14.16 | 4.905% | 0.045045 | 0.042
14.17 | 4.905% | 0.045045 | 0.042
14.18 | 4.905% | 0.045045 | 0.042
14.19 | 4.905% | 0.045045 | 0.042
14.2 | 4.905% | 0.045045 | 0.042
14.21 | 4.905% | 0.045045 | 0.042
14.22 | 4.905% | 0.045045 | 0.042
14.23 | 4.805% | 0.045045 | 0.042
14.24 | 4.805% | 0.045045 | 0.042
14.25 | 4.805% | 0.045045 | 0.042
14.26 | 4.805% | 0.045045 | 0.042
14.27 | 4.805% | 0.045045 | 0.042
14.28 | 4.805% | 0.045045 | 0.042
14.29 | 4.805% | 0.045045 | 0.041
14.3 | 4.805% | 0.044044 | 0.041
14.31 | 4.805% | 0.044044 | 0.041
14.32 | 4.805% | 0.044044 | 0.041
14.33 | 4.805% | 0.044044 | 0.041
14.34 | 4.805% | 0.044044 | 0.041
14.35 | 4.805% | 0.044044 | 0.041
14.36 | 4.805% | 0.044044 | 0.041
14.37 | 4.805% | 0.044044 | 0.041
14.38 | 4.805% | 0.044044 | 0.041
14.39 | 4.805% | 0.044044 | 0.041
14.4 | 4.805% | 0.044044 | 0.041
14.41 | 4.805% | 0.044044 | 0.041
14.42 | 4.805% | 0.044044 | 0.041
14.43 | 4.805% | 0.044044 | 0.041
14.44 | 4.805% | 0.044044 | 0.041
14.45 | 4.805% | 0.044044 | 0.041
14.46 | 4.805% | 0.044044 | 0.041
14.47 | 4.805% | 0.044044 | 0.041
14.48 | 4.805% | 0.044044 | 0.041
14.49 | 4.805% | 0.044044 | 0.041
14.5 | 4.805% | 0.044044 | 0.041
14.51 | 4.805% | 0.044044 | 0.041
14.52 | 4.805% | 0.044044 | 0.041
14.53 | 4.805% | 0.044044 | 0.041
14.54 | 4.805% | 0.044044 | 0.041
14.55 | 4.805% | 0.044044 | 0.041
14.56 | 4.805% | 0.044044 | 0.041
14.57 | 4.805% | 0.044044 | 0.041
14.58 | 4.805% | 0.044044 | 0.041
14.59 | 4.805% | 0.044044 | 0.041
14.6 | 4.805% | 0.044044 | 0.041
14.61 | 4.805% | 0.044044 | 0.041
14.62 | 4.805% | 0.044044 | 0.041
14.63 | 4.705% | 0.044044 | 0.041
14.64 | 4.705% | 0.044044 | 0.04
14.65 | 4.705% | 0.044044 | 0.04
14.66 | 4.705% | 0.044044 | 0.04
14.67 | 4.705% | 0.043043 | 0.04
14.68 | 4.705% | 0.043043 | 0.04
14.69 | 4.705% | 0.043043 | 0.04
14.7 | 4.705% | 0.043043 | 0.04
14.71 | 4.705% | 0.043043 | 0.04
14.72 | 4.705% | 0.043043 | 0.04
14.73 | 4.705% | 0.043043 | 0.04
14.74 | 4.705% | 0.043043 | 0.04
14.75 | 4.705% | 0.043043 | 0.04
14.76 | 4.705% | 0.043043 | 0.04
14.77 | 4.705% | 0.043043 | 0.04
14.78 | 4.705% | 0.043043 | 0.04
14.79 | 4.705% | 0.043043 | 0.04
14.8 | 4.705% | 0.043043 | 0.04
14.81 | 4.705% | 0.043043 | 0.04
14.82 | 4.705% | 0.043043 | 0.04
14.83 | 4.705% | 0.043043 | 0.04
14.84 | 4.705% | 0.043043 | 0.04
14.85 | 4.705% | 0.043043 | 0.04
14.86 | 4.705% | 0.043043 | 0.04
14.87 | 4.705% | 0.043043 | 0.04
14.88 | 4.705% | 0.043043 | 0.04
14.89 | 4.705% | 0.043043 | 0.04
14.9 | 4.705% | 0.043043 | 0.04
14.91 | 4.705% | 0.043043 | 0.04
14.92 | 4.705% | 0.042042 | 0.04
14.93 | 4.705% | 0.042042 | 0.04
14.94 | 4.705% | 0.042042 | 0.04
14.95 | 4.705% | 0.042042 | 0.04
14.96 | 4.705% | 0.042042 | 0.04
14.97 | 4.705% | 0.042042 | 0.04
14.98 | 4.705% | 0.042042 | 0.04
14.99 | 4.605% | 0.042042 | 0.04
15 | 4.605% | 0.042042 | 0.04
15.01 | 4.605% | 0.042042 | 0.04
15.02 | 4.605% | 0.042042 | 0.04
15.03 | 4.605% | 0.042042 | 0.04
15.04 | 4.605% | 0.042042 | 0.039
15.05 | 4.605% | 0.042042 | 0.039
15.06 | 4.605% | 0.042042 | 0.039
15.07 | 4.605% | 0.042042 | 0.039
15.08 | 4.605% | 0.042042 | 0.039
15.09 | 4.605% | 0.042042 | 0.039
15.1 | 4.605% | 0.042042 | 0.039
15.11 | 4.605% | 0.042042 | 0.039
15.12 | 4.605% | 0.042042 | 0.039
15.13 | 4.605% | 0.042042 | 0.039
15.14 | 4.605% | 0.042042 | 0.039
15.15 | 4.605% | 0.042042 | 0.039
15.16 | 4.605% | 0.042042 | 0.039
15.17 | 4.605% | 0.042042 | 0.039
15.18 | 4.605% | 0.042042 | 0.039
15.19 | 4.605% | 0.042042 | 0.039
15.2 | 4.605% | 0.042042 | 0.039
15.21 | 4.605% | 0.042042 | 0.039
15.22 | 4.605% | 0.042042 | 0.039
15.23 | 4.605% | 0.042042 | 0.039
15.24 | 4.605% | 0.042042 | 0.039
15.25 | 4.605% | 0.042042 | 0.039
15.26 | 4.605% | 0.042042 | 0.039
15.27 | 4.605% | 0.042042 | 0.039
15.28 | 4.605% | 0.042042 | 0.039
15.29 | 4.605% | 0.042042 | 0.039
15.3 | 4.605% | 0.042042 | 0.039
15.31 | 4.605% | 0.042042 | 0.039
15.32 | 4.605% | 0.042042 | 0.039
15.33 | 4.605% | 0.042042 | 0.039
15.34 | 4.605% | 0.042042 | 0.039
15.35 | 4.605% | 0.042042 | 0.039
15.36 | 4.605% | 0.042042 | 0.039
15.37 | 4.605% | 0.042042 | 0.039
15.38 | 4.605% | 0.042042 | 0.039
15.39 | 4.605% | 0.042042 | 0.039
15.4 | 4.605% | 0.041041 | 0.039
15.41 | 4.605% | 0.041041 | 0.039
15.42 | 4.605% | 0.041041 | 0.039
15.43 | 4.605% | 0.041041 | 0.039
15.44 | 4.605% | 0.041041 | 0.039
15.45 | 4.605% | 0.041041 | 0.039
15.46 | 4.605% | 0.041041 | 0.039
15.47 | 4.505% | 0.041041 | 0.039
15.48 | 4.505% | 0.041041 | 0.039
15.49 | 4.505% | 0.041041 | 0.039
15.5 | 4.505% | 0.041041 | 0.039
15.51 | 4.505% | 0.041041 | 0.039
15.52 | 4.505% | 0.041041 | 0.039
15.53 | 4.505% | 0.041041 | 0.039
15.54 | 4.505% | 0.041041 | 0.038
15.55 | 4.505% | 0.041041 | 0.038
15.56 | 4.505% | 0.041041 | 0.038
15.57 | 4.505% | 0.041041 | 0.038
15.58 | 4.505% | 0.041041 | 0.038
15.59 | 4.505% | 0.041041 | 0.038
15.6 | 4.505% | 0.041041 | 0.038
15.61 | 4.505% | 0.041041 | 0.038
15.62 | 4.505% | 0.041041 | 0.038
15.63 | 4.505% | 0.041041 | 0.038
15.64 | 4.505% | 0.041041 | 0.038
15.65 | 4.505% | 0.041041 | 0.038
15.66 | 4.505% | 0.041041 | 0.038
15.67 | 4.505% | 0.041041 | 0.038
15.68 | 4.505% | 0.041041 | 0.038
15.69 | 4.505% | 0.041041 | 0.038
15.7 | 4.505% | 0.041041 | 0.038
15.71 | 4.505% | 0.041041 | 0.038
15.72 | 4.505% | 0.041041 | 0.038
15.73 | 4.505% | 0.041041 | 0.038
15.74 | 4.505% | 0.041041 | 0.038
15.75 | 4.505% | 0.041041 | 0.038
15.76 | 4.505% | 0.041041 | 0.038
15.77 | 4.505% | 0.041041 | 0.038
15.78 | 4.505% | 0.041041 | 0.038
15.79 | 4.505% | 0.041041 | 0.038
15.8 | 4.505% | 0.041041 | 0.038
15.81 | 4.505% | 0.041041 | 0.038
15.82 | 4.505% | 0.041041 | 0.038
15.83 | 4.505% | 0.041041 | 0.038
15.84 | 4.505% | 0.041041 | 0.038
15.85 | 4.505% | 0.041041 | 0.038
15.86 | 4.505% | 0.041041 | 0.038
15.87 | 4.505% | 0.041041 | 0.038
|  | 0.041041 | 0.038
|  | 0.041041 | 0.038
|  | 0.04004 | 0.038
|  | 0.04004 | 0.037
|  | 0.04004 | 0.037
|  | 0.041041 | 0.037
0.037
0.037
0.037
0.037
0.037
0.037
0.037

## 16 MeV PDDs
Energy: | 16 MeV
SSD: | 100 cm
| PDD |  |  |  | Energy | SSD | Field size inline | R100 | R50 | R95 | R90 | R80 | Rp | Ds | PDD(3.22) | PDD(3.0)
Depth | 10x10+ | 6x6 & 7x7 | 4x4 | 3x3 | 16 MeV | 100.00 cm | 3x3 | 1.68 cm | 5.55 cm | 3.11 cm | 3.57 cm | 4.17 cm | 7.75 cm | 93.1% | 94.09% | 96%
0 |  | 91.69% | 92.89% |  | 16 MeV | 100.00 cm | 4x4 | 2.02 cm | 6.08 cm | 3.64 cm | 4.13 cm | 4.74 cm | 8.00 cm | 92.8% | 97.9% | 98.8%
0.01 |  | 91.79% | 92.89% |  | 16 MeV | 100.00 cm | 6x6 & 7x7 | 2.89 cm | 6.50 cm | 4.35 cm | 4.84 cm | 5.42 cm | 7.96 cm | 91.9% | 99.7% | 100%
0.02 |  | 91.79% | 92.89% |  | 16 MeV | 100.00 cm | 10x10+ | 3.22 cm | 6.63 cm | 4.64 cm | 5.09 cm | 5.63 cm | 8.00 cm |  | 100% | 100%
0.03 |  | 91.89% | 92.89%
0.04 |  | 91.99% | 92.89%
0.05 | 91.49% | 91.99% | 92.89% | 93.19%
0.06 | 91.59% | 92.09% | 92.89% | 93.29%
0.07 | 91.59% | 92.19% | 92.89% | 93.39%
0.08 | 91.69% | 92.19% | 92.99% | 93.39%
0.09 | 91.79% | 92.29% | 93.09% | 93.49%
0.1 | 91.79% | 92.29% | 93.09% | 93.59%
0.11 | 91.89% | 92.39% | 93.19% | 93.59%
0.12 | 91.89% | 92.49% | 93.19% | 93.69%
0.13 | 91.99% | 92.59% | 93.29% | 93.79%
0.14 | 92.09% | 92.69% | 93.39% | 93.79%
0.15 | 92.09% | 92.69% | 93.39% | 93.89%
0.16 | 92.19% | 92.79% | 93.49% | 93.99%
0.17 | 92.29% | 92.89% | 93.49% | 93.99%
0.18 | 92.29% | 92.99% | 93.59% | 94.09%
0.19 | 92.39% | 93.09% | 93.69% | 94.19%
0.2 | 92.49% | 93.19% | 93.69% | 94.29%
0.21 | 92.49% | 93.29% | 93.79% | 94.39%
0.22 | 92.59% | 93.39% | 93.89% | 94.49%
0.23 | 92.69% | 93.49% | 93.99% | 94.59%
0.24 | 92.69% | 93.59% | 93.99% | 94.69%
0.25 | 92.79% | 93.69% | 94.09% | 94.79%
0.26 | 92.89% | 93.79% | 94.19% | 94.89%
0.27 | 92.99% | 93.89% | 94.29% | 94.99%
0.28 | 93.09% | 93.99% | 94.39% | 95.2%
0.29 | 93.19% | 94.09% | 94.49% | 95.3%
0.3 | 93.29% | 94.19% | 94.49% | 95.4%
0.31 | 93.39% | 94.29% | 94.59% | 95.5%
0.32 | 93.49% | 94.39% | 94.69% | 95.6%
0.33 | 93.59% | 94.49% | 94.79% | 95.7%
0.34 | 93.69% | 94.59% | 94.89% | 95.8%
0.35 | 93.79% | 94.69% | 94.99% | 95.9%
0.36 | 93.89% | 94.79% | 94.99% | 96%
0.37 | 93.99% | 94.79% | 95.1% | 96%
0.38 | 93.99% | 94.89% | 95.2% | 96.1%
0.39 | 94.09% | 94.99% | 95.3% | 96.2%
0.4 | 94.19% | 95.1% | 95.4% | 96.3%
0.41 | 94.29% | 95.1% | 95.4% | 96.4%
0.42 | 94.39% | 95.2% | 95.5% | 96.5%
0.43 | 94.49% | 95.3% | 95.6% | 96.6%
0.44 | 94.49% | 95.4% | 95.7% | 96.7%
0.45 | 94.59% | 95.5% | 95.8% | 96.8%
0.46 | 94.69% | 95.5% | 95.8% | 96.8%
0.47 | 94.79% | 95.6% | 95.9% | 96.9%
0.48 | 94.89% | 95.7% | 96% | 97%
0.49 | 94.89% | 95.7% | 96.1% | 97.1%
0.5 | 94.99% | 95.8% | 96.1% | 97.2%
0.51 | 95.1% | 95.9% | 96.2% | 97.2%
0.52 | 95.2% | 95.9% | 96.3% | 97.3%
0.53 | 95.2% | 96% | 96.4% | 97.4%
0.54 | 95.3% | 96% | 96.4% | 97.5%
0.55 | 95.4% | 96.1% | 96.5% | 97.5%
0.56 | 95.4% | 96.2% | 96.6% | 97.5%
0.57 | 95.5% | 96.2% | 96.6% | 97.6%
0.58 | 95.6% | 96.3% | 96.7% | 97.6%
0.59 | 95.6% | 96.3% | 96.8% | 97.7%
0.6 | 95.7% | 96.4% | 96.8% | 97.7%
0.61 | 95.7% | 96.4% | 96.9% | 97.8%
0.62 | 95.8% | 96.5% | 97% | 97.8%
0.63 | 95.8% | 96.5% | 97% | 97.8%
0.64 | 95.9% | 96.5% | 97.1% | 97.9%
0.65 | 96% | 96.5% | 97.2% | 97.9%
0.66 | 96% | 96.6% | 97.2% | 98%
0.67 | 96.1% | 96.6% | 97.3% | 98%
0.68 | 96.1% | 96.6% | 97.3% | 98.1%
0.69 | 96.2% | 96.7% | 97.4% | 98.1%
0.7 | 96.2% | 96.7% | 97.5% | 98.2%
0.71 | 96.3% | 96.7% | 97.5% | 98.2%
0.72 | 96.3% | 96.7% | 97.6% | 98.3%
0.73 | 96.4% | 96.8% | 97.6% | 98.3%
0.74 | 96.4% | 96.8% | 97.7% | 98.3%
0.75 | 96.5% | 96.9% | 97.7% | 98.4%
0.76 | 96.5% | 96.9% | 97.8% | 98.4%
0.77 | 96.6% | 96.9% | 97.8% | 98.4%
0.78 | 96.6% | 97% | 97.8% | 98.5%
0.79 | 96.6% | 97% | 97.9% | 98.5%
0.8 | 96.7% | 97.1% | 97.9% | 98.5%
0.81 | 96.7% | 97.1% | 98% | 98.6%
0.82 | 96.8% | 97.2% | 98% | 98.6%
0.83 | 96.8% | 97.2% | 98% | 98.6%
0.84 | 96.9% | 97.2% | 98.1% | 98.7%
0.85 | 96.9% | 97.3% | 98.1% | 98.7%
0.86 | 97% | 97.3% | 98.2% | 98.7%
0.87 | 97% | 97.4% | 98.2% | 98.8%
0.88 | 97% | 97.4% | 98.2% | 98.8%
0.89 | 97.1% | 97.5% | 98.3% | 98.9%
0.9 | 97.1% | 97.5% | 98.3% | 98.9%
0.91 | 97.1% | 97.6% | 98.3% | 98.9%
0.92 | 97.2% | 97.6% | 98.4% | 98.9%
0.93 | 97.2% | 97.6% | 98.4% | 98.9%
0.94 | 97.2% | 97.7% | 98.4% | 99%
0.95 | 97.2% | 97.7% | 98.5% | 99%
0.96 | 97.3% | 97.8% | 98.5% | 99%
0.97 | 97.3% | 97.8% | 98.5% | 99%
0.98 | 97.3% | 97.9% | 98.6% | 99%
0.99 | 97.3% | 97.9% | 98.6% | 99%
1 | 97.4% | 97.9% | 98.6% | 99%
1.01 | 97.4% | 98% | 98.7% | 99.1%
1.02 | 97.4% | 98% | 98.7% | 99.1%
1.03 | 97.4% | 98.1% | 98.7% | 99.1%
1.04 | 97.5% | 98.1% | 98.8% | 99.1%
1.05 | 97.5% | 98.2% | 98.8% | 99.2%
1.06 | 97.5% | 98.2% | 98.8% | 99.2%
1.07 | 97.6% | 98.3% | 98.8% | 99.3%
1.08 | 97.6% | 98.3% | 98.9% | 99.3%
1.09 | 97.7% | 98.3% | 98.9% | 99.3%
1.1 | 97.7% | 98.3% | 98.9% | 99.4%
1.11 | 97.7% | 98.3% | 98.9% | 99.4%
1.12 | 97.8% | 98.4% | 99% | 99.4%
1.13 | 97.8% | 98.4% | 99% | 99.5%
1.14 | 97.8% | 98.4% | 99% | 99.5%
1.15 | 97.9% | 98.4% | 99% | 99.5%
1.16 | 97.9% | 98.4% | 99.1% | 99.5%
1.17 | 97.9% | 98.4% | 99.1% | 99.6%
1.18 | 97.9% | 98.4% | 99.1% | 99.6%
1.19 | 98% | 98.5% | 99.1% | 99.6%
1.2 | 98% | 98.5% | 99.1% | 99.6%
1.21 | 98% | 98.5% | 99.2% | 99.6%
1.22 | 98% | 98.5% | 99.2% | 99.7%
1.23 | 98.1% | 98.5% | 99.2% | 99.7%
1.24 | 98.1% | 98.5% | 99.2% | 99.7%
1.25 | 98.1% | 98.5% | 99.2% | 99.7%
1.26 | 98.1% | 98.6% | 99.3% | 99.7%
1.27 | 98.1% | 98.6% | 99.3% | 99.7%
1.28 | 98.2% | 98.6% | 99.3% | 99.7%
1.29 | 98.2% | 98.6% | 99.3% | 99.7%
1.3 | 98.2% | 98.6% | 99.3% | 99.7%
1.31 | 98.2% | 98.7% | 99.4% | 99.7%
1.32 | 98.3% | 98.7% | 99.4% | 99.7%
1.33 | 98.3% | 98.7% | 99.4% | 99.7%
1.34 | 98.3% | 98.7% | 99.4% | 99.7%
1.35 | 98.3% | 98.8% | 99.4% | 99.8%
1.36 | 98.3% | 98.8% | 99.4% | 99.8%
1.37 | 98.4% | 98.8% | 99.5% | 99.8%
1.38 | 98.4% | 98.8% | 99.5% | 99.8%
1.39 | 98.4% | 98.8% | 99.5% | 99.8%
1.4 | 98.4% | 98.9% | 99.5% | 99.8%
1.41 | 98.4% | 98.9% | 99.5% | 99.8%
1.42 | 98.4% | 98.9% | 99.5% | 99.8%
1.43 | 98.4% | 98.9% | 99.6% | 99.9%
1.44 | 98.5% | 98.9% | 99.6% | 99.9%
1.45 | 98.5% | 99% | 99.6% | 99.9%
1.46 | 98.5% | 99% | 99.6% | 99.9%
1.47 | 98.5% | 99% | 99.6% | 99.9%
1.48 | 98.5% | 99% | 99.6% | 99.9%
1.49 | 98.5% | 99% | 99.6% | 99.9%
1.5 | 98.6% | 99% | 99.7% | 100%
1.51 | 98.6% | 99% | 99.7% | 100%
1.52 | 98.6% | 99% | 99.7% | 100%
1.53 | 98.7% | 99% | 99.7% | 100%
1.54 | 98.7% | 99% | 99.7% | 100%
1.55 | 98.7% | 99% | 99.7% | 100%
1.56 | 98.8% | 99.1% | 99.7% | 100%
1.57 | 98.8% | 99.1% | 99.7% | 100%
1.58 | 98.8% | 99.1% | 99.7% | 100%
1.59 | 98.8% | 99.1% | 99.7% | 100%
1.6 | 98.9% | 99.1% | 99.7% | 99.9%
1.61 | 98.9% | 99.1% | 99.8% | 99.9%
1.62 | 98.9% | 99.1% | 99.8% | 99.9%
1.63 | 98.9% | 99.1% | 99.8% | 100%
1.64 | 98.9% | 99.2% | 99.8% | 100%
1.65 | 98.9% | 99.2% | 99.8% | 100%
1.66 | 98.9% | 99.2% | 99.8% | 100%
1.67 | 98.9% | 99.2% | 99.8% | 100%
1.68 | 98.9% | 99.2% | 99.8% | 100%
1.69 | 98.9% | 99.2% | 99.8% | 100%
1.7 | 99% | 99.2% | 99.8% | 100%
1.71 | 99% | 99.3% | 99.9% | 100%
1.72 | 99% | 99.3% | 99.9% | 100%
1.73 | 99% | 99.3% | 99.9% | 100%
1.74 | 99% | 99.4% | 99.9% | 100%
1.75 | 99% | 99.4% | 99.9% | 100%
1.76 | 99% | 99.4% | 99.9% | 100%
1.77 | 99% | 99.4% | 99.9% | 100%
1.78 | 99% | 99.5% | 99.9% | 99.9%
1.79 | 99.1% | 99.5% | 99.9% | 99.9%
1.8 | 99.1% | 99.5% | 99.9% | 99.9%
1.81 | 99.1% | 99.5% | 99.9% | 99.9%
1.82 | 99.1% | 99.5% | 99.9% | 99.9%
1.83 | 99.1% | 99.5% | 100% | 99.9%
1.84 | 99.2% | 99.5% | 100% | 99.9%
1.85 | 99.2% | 99.6% | 100% | 99.9%
1.86 | 99.2% | 99.6% | 100% | 99.9%
1.87 | 99.2% | 99.6% | 100% | 99.9%
1.88 | 99.2% | 99.6% | 100% | 99.9%
1.89 | 99.2% | 99.6% | 100% | 100%
1.9 | 99.2% | 99.6% | 100% | 100%
1.91 | 99.2% | 99.6% | 100% | 100%
1.92 | 99.2% | 99.6% | 100% | 100%
1.93 | 99.2% | 99.6% | 100% | 99.9%
1.94 | 99.2% | 99.6% | 100% | 99.9%
1.95 | 99.3% | 99.6% | 100% | 99.9%
1.96 | 99.3% | 99.6% | 100% | 99.9%
1.97 | 99.3% | 99.7% | 100% | 99.9%
1.98 | 99.3% | 99.7% | 100% | 99.9%
1.99 | 99.3% | 99.7% | 100% | 99.8%
2 | 99.3% | 99.7% | 100% | 99.8%
2.01 | 99.3% | 99.7% | 100% | 99.8%
2.02 | 99.4% | 99.7% | 100% | 99.7%
2.03 | 99.4% | 99.7% | 100% | 99.7%
2.04 | 99.4% | 99.7% | 100% | 99.7%
2.05 | 99.4% | 99.7% | 100% | 99.6%
2.06 | 99.4% | 99.7% | 100% | 99.6%
2.07 | 99.4% | 99.7% | 100% | 99.6%
2.08 | 99.5% | 99.7% | 100% | 99.6%
2.09 | 99.5% | 99.7% | 100% | 99.6%
2.1 | 99.5% | 99.7% | 100% | 99.6%
2.11 | 99.5% | 99.7% | 100% | 99.6%
2.12 | 99.5% | 99.7% | 100% | 99.6%
2.13 | 99.5% | 99.7% | 100% | 99.6%
2.14 | 99.5% | 99.7% | 100% | 99.6%
2.15 | 99.5% | 99.7% | 100% | 99.6%
2.16 | 99.5% | 99.7% | 100% | 99.6%
2.17 | 99.5% | 99.7% | 100% | 99.6%
2.18 | 99.5% | 99.8% | 100% | 99.6%
2.19 | 99.6% | 99.8% | 100% | 99.5%
2.2 | 99.6% | 99.8% | 100% | 99.5%
2.21 | 99.6% | 99.8% | 100% | 99.5%
2.22 | 99.6% | 99.8% | 99.9% | 99.5%
2.23 | 99.6% | 99.8% | 99.9% | 99.5%
2.24 | 99.6% | 99.9% | 99.9% | 99.5%
2.25 | 99.6% | 99.9% | 99.9% | 99.5%
2.26 | 99.6% | 99.9% | 99.9% | 99.5%
2.27 | 99.6% | 99.9% | 99.9% | 99.5%
2.28 | 99.6% | 99.9% | 99.9% | 99.4%
2.29 | 99.6% | 99.9% | 99.9% | 99.4%
2.3 | 99.6% | 99.9% | 99.9% | 99.4%
2.31 | 99.6% | 99.9% | 99.9% | 99.3%
2.32 | 99.6% | 99.9% | 99.9% | 99.3%
2.33 | 99.6% | 99.9% | 99.8% | 99.3%
2.34 | 99.6% | 99.9% | 99.8% | 99.2%
2.35 | 99.6% | 99.9% | 99.8% | 99.2%
2.36 | 99.6% | 99.9% | 99.8% | 99.1%
2.37 | 99.6% | 99.9% | 99.8% | 99.1%
2.38 | 99.6% | 99.9% | 99.8% | 99%
2.39 | 99.6% | 99.9% | 99.8% | 99%
2.4 | 99.7% | 99.9% | 99.8% | 99%
2.41 | 99.7% | 99.9% | 99.8% | 98.9%
2.42 | 99.7% | 99.9% | 99.8% | 98.9%
2.43 | 99.7% | 99.9% | 99.8% | 98.8%
2.44 | 99.7% | 99.9% | 99.8% | 98.8%
2.45 | 99.7% | 99.9% | 99.8% | 98.8%
2.46 | 99.7% | 99.9% | 99.8% | 98.7%
2.47 | 99.7% | 100% | 99.7% | 98.7%
2.48 | 99.7% | 100% | 99.7% | 98.7%
2.49 | 99.7% | 100% | 99.7% | 98.7%
2.5 | 99.8% | 100% | 99.7% | 98.6%
2.51 | 99.8% | 100% | 99.7% | 98.6%
2.52 | 99.8% | 100% | 99.7% | 98.6%
2.53 | 99.8% | 100% | 99.7% | 98.6%
2.54 | 99.8% | 100% | 99.7% | 98.5%
2.55 | 99.9% | 100% | 99.7% | 98.5%
2.56 | 99.9% | 100% | 99.7% | 98.5%
2.57 | 99.9% | 100% | 99.7% | 98.5%
2.58 | 99.9% | 100% | 99.7% | 98.4%
2.59 | 99.9% | 100% | 99.6% | 98.4%
2.6 | 99.9% | 100% | 99.6% | 98.4%
2.61 | 99.9% | 100% | 99.6% | 98.3%
2.62 | 99.9% | 100% | 99.6% | 98.3%
2.63 | 99.9% | 100% | 99.6% | 98.3%
2.64 | 100% | 100% | 99.6% | 98.2%
2.65 | 100% | 100% | 99.6% | 98.2%
2.66 | 100% | 100% | 99.6% | 98.2%
2.67 | 99.9% | 100% | 99.6% | 98.1%
2.68 | 99.9% | 100% | 99.6% | 98.1%
2.69 | 99.9% | 100% | 99.6% | 98%
2.7 | 99.9% | 100% | 99.5% | 98%
2.71 | 99.9% | 100% | 99.5% | 97.9%
2.72 | 99.9% | 100% | 99.5% | 97.8%
2.73 | 99.9% | 100% | 99.5% | 97.8%
2.74 | 99.9% | 100% | 99.5% | 97.7%
2.75 | 99.9% | 100% | 99.5% | 97.6%
2.76 | 99.9% | 100% | 99.4% | 97.6%
2.77 | 99.9% | 100% | 99.4% | 97.5%
2.78 | 99.9% | 100% | 99.4% | 97.4%
2.79 | 99.9% | 100% | 99.4% | 97.4%
2.8 | 99.9% | 100% | 99.4% | 97.3%
2.81 | 99.9% | 100% | 99.3% | 97.3%
2.82 | 99.9% | 100% | 99.3% | 97.2%
2.83 | 99.9% | 100% | 99.3% | 97.1%
2.84 | 100% | 100% | 99.3% | 97.1%
2.85 | 100% | 100% | 99.3% | 97%
2.86 | 100% | 100% | 99.2% | 97%
2.87 | 100% | 100% | 99.2% | 96.9%
2.88 | 100% | 100% | 99.2% | 96.9%
2.89 | 100% | 100% | 99.2% | 96.8%
2.9 | 100% | 100% | 99.1% | 96.7%
2.91 | 100% | 100% | 99.1% | 96.7%
2.92 | 100% | 100% | 99.1% | 96.6%
2.93 | 100% | 100% | 99.1% | 96.5%
2.94 | 100% | 100% | 99% | 96.5%
2.95 | 100% | 100% | 99% | 96.4%
2.96 | 100% | 100% | 99% | 96.3%
2.97 | 100% | 100% | 98.9% | 96.2%
2.98 | 100% | 100% | 98.9% | 96.1%
2.99 | 100% | 100% | 98.9% | 96.1%
3 | 100% | 100% | 98.8% | 96%
3.01 | 100% | 100% | 98.8% | 95.9%
3.02 | 100% | 100% | 98.8% | 95.8%
3.03 | 100% | 100% | 98.7% | 95.7%
3.04 | 100% | 100% | 98.7% | 95.6%
3.05 | 100% | 100% | 98.7% | 95.5%
3.06 | 100% | 100% | 98.6% | 95.4%
3.07 | 100% | 99.9% | 98.6% | 95.4%
3.08 | 100% | 99.9% | 98.5% | 95.3%
3.09 | 100% | 99.9% | 98.5% | 95.2%
3.1 | 100% | 99.9% | 98.5% | 95.1%
3.11 | 100% | 99.9% | 98.4% | 94.99%
3.12 | 100% | 99.9% | 98.4% | 94.99%
3.13 | 100% | 99.9% | 98.3% | 94.89%
3.14 | 100% | 99.9% | 98.3% | 94.79%
3.15 | 100% | 99.8% | 98.3% | 94.69%
3.16 | 100% | 99.8% | 98.2% | 94.59%
3.17 | 100% | 99.8% | 98.2% | 94.49%
3.18 | 100% | 99.8% | 98.1% | 94.49%
3.19 | 100% | 99.7% | 98.1% | 94.39%
3.2 | 100% | 99.7% | 98% | 94.29%
3.21 | 100% | 99.7% | 98% | 94.19%
3.22 | 100% | 99.7% | 97.9% | 94.09%
3.23 | 100% | 99.6% | 97.8% | 93.99%
3.24 | 100% | 99.6% | 97.8% | 93.89%
3.25 | 100% | 99.6% | 97.7% | 93.79%
3.26 | 100% | 99.6% | 97.7% | 93.69%
3.27 | 100% | 99.6% | 97.6% | 93.59%
3.28 | 100% | 99.6% | 97.6% | 93.49%
3.29 | 100% | 99.6% | 97.5% | 93.39%
3.3 | 99.9% | 99.5% | 97.5% | 93.29%
3.31 | 99.9% | 99.5% | 97.4% | 93.19%
3.32 | 99.9% | 99.5% | 97.3% | 93.09%
3.33 | 99.9% | 99.5% | 97.3% | 92.99%
3.34 | 99.9% | 99.5% | 97.2% | 92.89%
3.35 | 99.9% | 99.5% | 97.2% | 92.79%
3.36 | 99.9% | 99.5% | 97.1% | 92.59%
3.37 | 99.9% | 99.5% | 97% | 92.49%
3.38 | 99.8% | 99.5% | 97% | 92.39%
3.39 | 99.8% | 99.4% | 96.9% | 92.29%
3.4 | 99.8% | 99.4% | 96.8% | 92.19%
3.41 | 99.8% | 99.4% | 96.8% | 91.99%
3.42 | 99.8% | 99.4% | 96.7% | 91.89%
3.43 | 99.7% | 99.3% | 96.6% | 91.79%
3.44 | 99.7% | 99.3% | 96.6% | 91.69%
3.45 | 99.7% | 99.3% | 96.5% | 91.59%
3.46 | 99.7% | 99.3% | 96.4% | 91.39%
3.47 | 99.7% | 99.2% | 96.3% | 91.29%
3.48 | 99.7% | 99.2% | 96.2% | 91.19%
3.49 | 99.7% | 99.2% | 96.2% | 91.09%
3.5 | 99.6% | 99.1% | 96.1% | 90.89%
3.51 | 99.6% | 99.1% | 96% | 90.79%
3.52 | 99.6% | 99.1% | 95.9% | 90.69%
3.53 | 99.6% | 99.1% | 95.9% | 90.49%
3.54 | 99.6% | 99.1% | 95.8% | 90.39%
3.55 | 99.6% | 99% | 95.7% | 90.29%
3.56 | 99.6% | 99% | 95.6% | 90.09%
3.57 | 99.6% | 99% | 95.6% | 89.99%
3.58 | 99.6% | 99% | 95.5% | 89.89%
3.59 | 99.6% | 98.9% | 95.4% | 89.69%
3.6 | 99.6% | 98.9% | 95.3% | 89.59%
3.61 | 99.6% | 98.9% | 95.3% | 89.49%
3.62 | 99.6% | 98.9% | 95.2% | 89.29%
3.63 | 99.6% | 98.8% | 95.1% | 89.19%
3.64 | 99.6% | 98.8% | 94.99% | 88.99%
3.65 | 99.5% | 98.8% | 94.89% | 88.89%
3.66 | 99.5% | 98.8% | 94.79% | 88.79%
3.67 | 99.5% | 98.8% | 94.79% | 88.59%
3.68 | 99.5% | 98.7% | 94.69% | 88.49%
3.69 | 99.5% | 98.7% | 94.59% | 88.29%
3.7 | 99.4% | 98.7% | 94.49% | 88.19%
3.71 | 99.4% | 98.7% | 94.39% | 87.99%
3.72 | 99.4% | 98.6% | 94.29% | 87.79%
3.73 | 99.4% | 98.6% | 94.19% | 87.69%
3.74 | 99.4% | 98.6% | 94.09% | 87.49%
3.75 | 99.4% | 98.5% | 93.99% | 87.39%
3.76 | 99.3% | 98.5% | 93.99% | 87.19%
3.77 | 99.3% | 98.5% | 93.89% | 86.99%
3.78 | 99.3% | 98.4% | 93.79% | 86.89%
3.79 | 99.3% | 98.4% | 93.69% | 86.69%
3.8 | 99.3% | 98.3% | 93.59% | 86.59%
3.81 | 99.2% | 98.3% | 93.49% | 86.39%
3.82 | 99.2% | 98.3% | 93.39% | 86.19%
3.83 | 99.2% | 98.2% | 93.29% | 86.09%
3.84 | 99.2% | 98.2% | 93.19% | 85.89%
3.85 | 99.1% | 98.1% | 93.09% | 85.79%
3.86 | 99.1% | 98.1% | 92.99% | 85.59%
3.87 | 99.1% | 98% | 92.89% | 85.39%
3.88 | 99.1% | 98% | 92.79% | 85.29%
3.89 | 99% | 97.9% | 92.69% | 85.09%
3.9 | 99% | 97.8% | 92.59% | 84.88%
3.91 | 99% | 97.8% | 92.49% | 84.68%
3.92 | 99% | 97.7% | 92.39% | 84.58%
3.93 | 98.9% | 97.7% | 92.29% | 84.38%
3.94 | 98.9% | 97.6% | 92.19% | 84.18%
3.95 | 98.9% | 97.5% | 92.09% | 83.98%
3.96 | 98.9% | 97.5% | 91.99% | 83.78%
3.97 | 98.8% | 97.4% | 91.89% | 83.58%
3.98 | 98.8% | 97.4% | 91.79% | 83.48%
3.99 | 98.8% | 97.4% | 91.69% | 83.28%
4 | 98.7% | 97.3% | 91.59% | 83.08%
4.01 | 98.7% | 97.3% | 91.49% | 82.88%
4.02 | 98.7% | 97.3% | 91.29% | 82.68%
4.03 | 98.7% | 97.3% | 91.19% | 82.58%
4.04 | 98.6% | 97.2% | 91.09% | 82.38%
4.05 | 98.6% | 97.2% | 90.99% | 82.18%
4.06 | 98.5% | 97.2% | 90.89% | 81.98%
4.07 | 98.5% | 97.2% | 90.69% | 81.88%
4.08 | 98.4% | 97.1% | 90.59% | 81.68%
4.09 | 98.4% | 97.1% | 90.49% | 81.48%
4.1 | 98.3% | 97.1% | 90.39% | 81.28%
4.11 | 98.3% | 97% | 90.29% | 81.08%
4.12 | 98.2% | 96.9% | 90.09% | 80.98%
4.13 | 98.2% | 96.8% | 89.99% | 80.78%
4.14 | 98.2% | 96.8% | 89.89% | 80.58%
4.15 | 98.1% | 96.7% | 89.69% | 80.38%
4.16 | 98.1% | 96.6% | 89.59% | 80.18%
4.17 | 98.1% | 96.5% | 89.49% | 79.98%
4.18 | 98% | 96.5% | 89.29% | 79.78%
4.19 | 98% | 96.4% | 89.19% | 79.58%
4.2 | 98% | 96.3% | 88.99% | 79.48%
4.21 | 97.9% | 96.2% | 88.89% | 79.28%
4.22 | 97.9% | 96.1% | 88.69% | 79.08%
4.23 | 97.8% | 96.1% | 88.59% | 78.78%
4.24 | 97.8% | 96% | 88.39% | 78.58%
4.25 | 97.8% | 95.9% | 88.29% | 78.38%
4.26 | 97.7% | 95.8% | 88.19% | 78.18%
4.27 | 97.7% | 95.7% | 87.99% | 77.98%
4.28 | 97.6% | 95.7% | 87.89% | 77.78%
4.29 | 97.6% | 95.6% | 87.69% | 77.58%
4.3 | 97.5% | 95.5% | 87.49% | 77.38%
4.31 | 97.4% | 95.4% | 87.39% | 77.08%
4.32 | 97.4% | 95.3% | 87.19% | 76.88%
4.33 | 97.3% | 95.2% | 87.09% | 76.68%
4.34 | 97.2% | 95.2% | 86.89% | 76.48%
4.35 | 97.2% | 95.1% | 86.79% | 76.28%
4.36 | 97.1% | 94.99% | 86.59% | 76.08%
4.37 | 97.1% | 94.89% | 86.49% | 75.78%
4.38 | 97% | 94.89% | 86.29% | 75.58%
4.39 | 97% | 94.79% | 86.09% | 75.38%
4.4 | 96.9% | 94.69% | 85.99% | 75.18%
4.41 | 96.8% | 94.69% | 85.79% | 74.97%
4.42 | 96.8% | 94.59% | 85.69% | 74.67%
4.43 | 96.7% | 94.49% | 85.49% | 74.47%
4.44 | 96.6% | 94.49% | 85.29% | 74.27%
4.45 | 96.6% | 94.39% | 85.09% | 74.07%
4.46 | 96.5% | 94.29% | 84.98% | 73.87%
4.47 | 96.4% | 94.19% | 84.78% | 73.67%
4.48 | 96.4% | 94.19% | 84.58% | 73.47%
4.49 | 96.3% | 94.09% | 84.48% | 73.27%
4.5 | 96.2% | 93.99% | 84.28% | 73.07%
4.51 | 96.1% | 93.79% | 84.08% | 72.87%
4.52 | 96.1% | 93.69% | 83.88% | 72.67%
4.53 | 96% | 93.59% | 83.78% | 72.47%
4.54 | 95.9% | 93.49% | 83.58% | 72.27%
4.55 | 95.8% | 93.39% | 83.38% | 72.07%
4.56 | 95.7% | 93.29% | 83.18% | 71.87%
4.57 | 95.7% | 93.19% | 83.08% | 71.67%
4.58 | 95.6% | 93.09% | 82.88% | 71.37%
4.59 | 95.5% | 92.99% | 82.68% | 71.17%
4.6 | 95.4% | 92.89% | 82.58% | 70.97%
4.61 | 95.3% | 92.79% | 82.38% | 70.77%
4.62 | 95.2% | 92.69% | 82.18% | 70.47%
4.63 | 95.1% | 92.49% | 81.98% | 70.27%
4.64 | 94.99% | 92.39% | 81.78% | 70.07%
4.65 | 94.99% | 92.29% | 81.68% | 69.87%
4.66 | 94.89% | 92.19% | 81.48% | 69.67%
4.67 | 94.79% | 91.99% | 81.28% | 69.37%
4.68 | 94.69% | 91.89% | 81.08% | 69.17%
4.69 | 94.59% | 91.79% | 80.88% | 68.97%
4.7 | 94.49% | 91.69% | 80.78% | 68.77%
4.71 | 94.39% | 91.49% | 80.58% | 68.57%
4.72 | 94.29% | 91.39% | 80.38% | 68.37%
4.73 | 94.19% | 91.29% | 80.18% | 68.17%
4.74 | 94.09% | 91.19% | 79.98% | 67.87%
4.75 | 93.99% | 91.09% | 79.88% | 67.67%
4.76 | 93.89% | 90.99% | 79.68% | 67.47%
4.77 | 93.89% | 90.89% | 79.48% | 67.27%
4.78 | 93.79% | 90.69% | 79.28% | 67.07%
4.79 | 93.69% | 90.59% | 79.08% | 66.87%
4.8 | 93.59% | 90.49% | 78.88% | 66.67%
4.81 | 93.49% | 90.39% | 78.78% | 66.37%
4.82 | 93.39% | 90.29% | 78.58% | 66.17%
4.83 | 93.19% | 90.19% | 78.38% | 65.97%
4.84 | 93.09% | 89.99% | 78.18% | 65.77%
4.85 | 92.99% | 89.89% | 77.98% | 65.57%
4.86 | 92.89% | 89.69% | 77.78% | 65.27%
4.87 | 92.79% | 89.59% | 77.58% | 65.07%
4.88 | 92.69% | 89.49% | 77.38% | 64.86%
4.89 | 92.49% | 89.29% | 77.18% | 64.66%
4.9 | 92.39% | 89.19% | 76.98% | 64.36%
4.91 | 92.29% | 88.99% | 76.78% | 64.16%
4.92 | 92.19% | 88.89% | 76.58% | 63.96%
4.93 | 91.99% | 88.79% | 76.38% | 63.66%
4.94 | 91.89% | 88.59% | 76.18% | 63.46%
4.95 | 91.79% | 88.49% | 75.98% | 63.26%
4.96 | 91.69% | 88.29% | 75.78% | 63.06%
4.97 | 91.59% | 88.19% | 75.58% | 62.76%
4.98 | 91.39% | 87.99% | 75.38% | 62.56%
4.99 | 91.29% | 87.89% | 75.18% | 62.36%
5 | 91.19% | 87.69% | 74.97% | 62.16%
5.01 | 91.09% | 87.49% | 74.77% | 61.86%
5.02 | 90.99% | 87.39% | 74.57% | 61.66%
5.03 | 90.79% | 87.19% | 74.37% | 61.46%
5.04 | 90.69% | 87.09% | 74.07% | 61.26%
5.05 | 90.59% | 86.89% | 73.87% | 60.96%
5.06 | 90.39% | 86.79% | 73.67% | 60.76%
5.07 | 90.29% | 86.59% | 73.47% | 60.56%
5.08 | 90.19% | 86.39% | 73.27% | 60.36%
5.09 | 89.99% | 86.19% | 73.07% | 60.16%
5.1 | 89.89% | 85.99% | 72.77% | 59.96%
5.11 | 89.69% | 85.89% | 72.57% | 59.66%
5.12 | 89.59% | 85.69% | 72.37% | 59.46%
5.13 | 89.39% | 85.49% | 72.17% | 59.26%
5.14 | 89.19% | 85.29% | 71.97% | 59.06%
5.15 | 89.09% | 85.09% | 71.77% | 58.76%
5.16 | 88.89% | 84.98% | 71.47% | 58.56%
5.17 | 88.79% | 84.78% | 71.27% | 58.36%
5.18 | 88.59% | 84.58% | 71.07% | 58.16%
5.19 | 88.39% | 84.38% | 70.87% | 57.86%
5.2 | 88.29% | 84.18% | 70.67% | 57.66%
5.21 | 88.09% | 84.08% | 70.47% | 57.46%
5.22 | 87.89% | 83.88% | 70.17% | 57.16%
5.23 | 87.79% | 83.68% | 69.97% | 56.96%
5.24 | 87.59% | 83.48% | 69.77% | 56.76%
5.25 | 87.49% | 83.38% | 69.57% | 56.56%
5.26 | 87.29% | 83.18% | 69.27% | 56.36%
5.27 | 87.09% | 82.98% | 69.07% | 56.16%
5.28 | 86.99% | 82.78% | 68.87% | 55.86%
5.29 | 86.79% | 82.68% | 68.67% | 55.66%
5.3 | 86.59% | 82.48% | 68.37% | 55.46%
5.31 | 86.49% | 82.28% | 68.17% | 55.26%
5.32 | 86.29% | 82.08% | 67.97% | 55.06%
5.33 | 86.09% | 81.88% | 67.77% | 54.85%
5.34 | 85.89% | 81.68% | 67.47% | 54.55%
5.35 | 85.79% | 81.48% | 67.27% | 54.35%
5.36 | 85.59% | 81.28% | 67.07% | 54.15%
5.37 | 85.39% | 80.98% | 66.77% | 53.95%
5.38 | 85.19% | 80.78% | 66.57% | 53.65%
5.39 | 84.98% | 80.58% | 66.37% | 53.45%
5.4 | 84.78% | 80.38% | 66.17% | 53.25%
5.41 | 84.58% | 80.18% | 65.87% | 52.95%
5.42 | 84.38% | 79.98% | 65.67% | 52.75%
5.43 | 84.18% | 79.78% | 65.47% | 52.55%
5.44 | 83.98% | 79.58% | 65.27% | 52.35%
5.45 | 83.78% | 79.38% | 64.96% | 52.15%
5.46 | 83.58% | 79.08% | 64.76% | 51.85%
5.47 | 83.38% | 78.88% | 64.56% | 51.65%
5.48 | 83.18% | 78.68% | 64.36% | 51.45%
5.49 | 82.98% | 78.48% | 64.06% | 51.25%
5.5 | 82.78% | 78.28% | 63.86% | 51.05%
5.51 | 82.58% | 78.08% | 63.66% | 50.85%
5.52 | 82.38% | 77.78% | 63.46% | 50.65%
5.53 | 82.18% | 77.58% | 63.16% | 50.45%
5.54 | 81.98% | 77.38% | 62.96% | 50.15%
5.55 | 81.68% | 77.18% | 62.76% | 49.95%
5.56 | 81.48% | 76.88% | 62.56% | 49.75%
5.57 | 81.28% | 76.68% | 62.26% | 49.55%
5.58 | 81.08% | 76.38% | 62.06% | 49.35%
5.59 | 80.88% | 76.18% | 61.86% | 49.15%
5.6 | 80.58% | 75.98% | 61.66% | 48.85%
5.61 | 80.38% | 75.68% | 61.36% | 48.65%
5.62 | 80.18% | 75.48% | 61.16% | 48.45%
5.63 | 79.98% | 75.18% | 60.96% | 48.25%
5.64 | 79.68% | 74.97% | 60.66% | 47.95%
5.65 | 79.48% | 74.77% | 60.46% | 47.75%
5.66 | 79.28% | 74.47% | 60.26% | 47.55%
5.67 | 78.98% | 74.27% | 59.96% | 47.35%
5.68 | 78.78% | 73.97% | 59.76% | 47.05%
5.69 | 78.48% | 73.77% | 59.56% | 46.85%
5.7 | 78.28% | 73.47% | 59.26% | 46.65%
5.71 | 78.08% | 73.27% | 59.06% | 46.45%
5.72 | 77.78% | 72.97% | 58.76% | 46.15%
5.73 | 77.58% | 72.67% | 58.56% | 45.95%
5.74 | 77.28% | 72.47% | 58.36% | 45.75%
5.75 | 77.08% | 72.17% | 58.06% | 45.55%
5.76 | 76.78% | 71.97% | 57.86% | 45.25%
5.77 | 76.58% | 71.67% | 57.56% | 45.05%
5.78 | 76.28% | 71.47% | 57.36% | 44.84%
5.79 | 75.98% | 71.17% | 57.06% | 44.64%
5.8 | 75.78% | 70.97% | 56.86% | 44.44%
5.81 | 75.48% | 70.67% | 56.56% | 44.24%
5.82 | 75.28% | 70.37% | 56.36% | 44.04%
5.83 | 74.97% | 70.17% | 56.16% | 43.84%
5.84 | 74.67% | 69.87% | 55.86% | 43.64%
5.85 | 74.47% | 69.67% | 55.66% | 43.44%
5.86 | 74.17% | 69.37% | 55.36% | 43.24%
5.87 | 73.97% | 69.17% | 55.16% | 43.04%
5.88 | 73.67% | 68.87% | 54.85% | 42.84%
5.89 | 73.37% | 68.67% | 54.65% | 42.64%
5.9 | 73.17% | 68.37% | 54.35% | 42.44%
5.91 | 72.87% | 68.07% | 54.15% | 42.24%
5.92 | 72.57% | 67.87% | 53.95% | 42.04%
5.93 | 72.27% | 67.57% | 53.65% | 41.84%
5.94 | 71.97% | 67.27% | 53.45% | 41.64%
5.95 | 71.77% | 66.97% | 53.15% | 41.44%
5.96 | 71.47% | 66.67% | 52.95% | 41.24%
5.97 | 71.17% | 66.47% | 52.65% | 41.04%
5.98 | 70.87% | 66.17% | 52.45% | 40.84%
5.99 | 70.57% | 65.87% | 52.15% | 40.64%
6 | 70.27% | 65.57% | 51.95% | 40.44%
6.01 | 70.07% | 65.27% | 51.65% | 40.24%
6.02 | 69.77% | 64.96% | 51.45% | 39.94%
6.03 | 69.47% | 64.66% | 51.15% | 39.74%
6.04 | 69.17% | 64.36% | 50.85% | 39.54%
6.05 | 68.87% | 64.06% | 50.65% | 39.34%
6.06 | 68.57% | 63.76% | 50.35% | 39.14%
6.07 | 68.27% | 63.46% | 50.15% | 38.94%
6.08 | 67.97% | 63.16% | 49.85% | 38.74%
6.09 | 67.67% | 62.86% | 49.65% | 38.44%
6.1 | 67.37% | 62.56% | 49.35% | 38.24%
6.11 | 67.07% | 62.26% | 49.15% | 38.04%
6.12 | 66.77% | 61.96% | 48.85% | 37.84%
6.13 | 66.47% | 61.66% | 48.65% | 37.64%
6.14 | 66.17% | 61.36% | 48.35% | 37.44%
6.15 | 65.87% | 61.06% | 48.15% | 37.24%
6.16 | 65.57% | 60.76% | 47.85% | 37.04%
6.17 | 65.27% | 60.36% | 47.65% | 36.84%
6.18 | 64.96% | 60.06% | 47.35% | 36.64%
6.19 | 64.66% | 59.76% | 47.15% | 36.44%
6.2 | 64.26% | 59.46% | 46.85% | 36.24%
6.21 | 63.96% | 59.16% | 46.65% | 36.04%
6.22 | 63.66% | 58.76% | 46.35% | 35.84%
6.23 | 63.36% | 58.46% | 46.15% | 35.64%
6.24 | 63.06% | 58.16% | 45.85% | 35.44%
6.25 | 62.76% | 57.86% | 45.65% | 35.24%
6.26 | 62.46% | 57.56% | 45.35% | 35.04%
6.27 | 62.16% | 57.16% | 45.15% | 34.93%
6.28 | 61.86% | 56.86% | 44.84% | 34.73%
6.29 | 61.56% | 56.56% | 44.64% | 34.53%
6.3 | 61.26% | 56.26% | 44.34% | 34.33%
6.31 | 60.96% | 55.96% | 44.14% | 34.13%
6.32 | 60.56% | 55.66% | 43.84% | 33.93%
6.33 | 60.26% | 55.36% | 43.64% | 33.73%
6.34 | 59.96% | 55.06% | 43.44% | 33.53%
6.35 | 59.66% | 54.75% | 43.14% | 33.33%
6.36 | 59.26% | 54.35% | 42.94% | 33.13%
6.37 | 58.96% | 54.05% | 42.64% | 32.93%
6.38 | 58.66% | 53.75% | 42.44% | 32.73%
6.39 | 58.26% | 53.45% | 42.14% | 32.53%
6.4 | 57.96% | 53.15% | 41.94% | 32.33%
6.41 | 57.56% | 52.85% | 41.64% | 32.13%
6.42 | 57.26% | 52.45% | 41.44% | 31.93%
6.43 | 56.86% | 52.15% | 41.14% | 31.73%
6.44 | 56.56% | 51.85% | 40.94% | 31.53%
6.45 | 56.16% | 51.55% | 40.74% | 31.33%
6.46 | 55.86% | 51.15% | 40.44% | 31.13%
6.47 | 55.46% | 50.85% | 40.24% | 31.03%
6.48 | 55.16% | 50.55% | 39.94% | 30.83%
6.49 | 54.85% | 50.25% | 39.74% | 30.63%
6.5 | 54.45% | 49.95% | 39.44% | 30.43%
6.51 | 54.15% | 49.55% | 39.24% | 30.23%
6.52 | 53.75% | 49.25% | 39.04% | 30.03%
6.53 | 53.45% | 48.95% | 38.74% | 29.83%
6.54 | 53.15% | 48.65% | 38.54% | 29.63%
6.55 | 52.75% | 48.25% | 38.24% | 29.53%
6.56 | 52.45% | 47.95% | 38.04% | 29.33%
6.57 | 52.15% | 47.65% | 37.74% | 29.13%
6.58 | 51.75% | 47.35% | 37.54% | 28.93%
6.59 | 51.45% | 46.95% | 37.34% | 28.73%
6.6 | 51.15% | 46.65% | 37.04% | 28.53%
6.61 | 50.75% | 46.35% | 36.84% | 28.33%
6.62 | 50.45% | 46.05% | 36.54% | 28.13%
6.63 | 50.15% | 45.75% | 36.34% | 27.93%
6.64 | 49.75% | 45.35% | 36.04% | 27.73%
6.65 | 49.45% | 45.05% | 35.84% | 27.53%
6.66 | 49.15% | 44.64% | 35.54% | 27.33%
6.67 | 48.75% | 44.34% | 35.34% | 27.13%
6.68 | 48.45% | 44.04% | 35.14% | 27.03%
6.69 | 48.05% | 43.64% | 34.83% | 26.83%
6.7 | 47.75% | 43.34% | 34.63% | 26.63%
6.71 | 47.35% | 43.04% | 34.33% | 26.43%
6.72 | 47.05% | 42.64% | 34.13% | 26.23%
6.73 | 46.65% | 42.34% | 33.83% | 26.03%
6.74 | 46.35% | 41.94% | 33.63% | 25.93%
6.75 | 45.95% | 41.64% | 33.43% | 25.73%
6.76 | 45.65% | 41.34% | 33.13% | 25.53%
6.77 | 45.25% | 41.04% | 32.93% | 25.33%
6.78 | 44.94% | 40.64% | 32.63% | 25.23%
6.79 | 44.54% | 40.34% | 32.43% | 25.03%
6.8 | 44.24% | 40.04% | 32.23% | 24.82%
6.81 | 43.84% | 39.74% | 31.93% | 24.62%
6.82 | 43.54% | 39.34% | 31.73% | 24.52%
6.83 | 43.14% | 39.04% | 31.43% | 24.32%
6.84 | 42.84% | 38.74% | 31.23% | 24.12%
6.85 | 42.44% | 38.44% | 31.03% | 24.02%
6.86 | 42.14% | 38.04% | 30.73% | 23.82%
6.87 | 41.74% | 37.74% | 30.53% | 23.62%
6.88 | 41.44% | 37.44% | 30.33% | 23.42%
6.89 | 41.14% | 37.04% | 30.03% | 23.32%
6.9 | 40.74% | 36.74% | 29.83% | 23.12%
6.91 | 40.44% | 36.44% | 29.53% | 22.92%
6.92 | 40.04% | 36.14% | 29.33% | 22.72%
6.93 | 39.74% | 35.84% | 29.13% | 22.52%
6.94 | 39.44% | 35.44% | 28.83% | 22.32%
6.95 | 39.04% | 35.14% | 28.63% | 22.22%
6.96 | 38.74% | 34.83% | 28.43% | 22.02%
6.97 | 38.44% | 34.53% | 28.13% | 21.82%
6.98 | 38.04% | 34.23% | 27.93% | 21.62%
6.99 | 37.74% | 33.93% | 27.73% | 21.42%
7 | 37.34% | 33.53% | 27.43% | 21.32%
7.01 | 37.04% | 33.23% | 27.23% | 21.12%
7.02 | 36.74% | 32.93% | 27.03% | 20.92%
7.03 | 36.34% | 32.63% | 26.73% | 20.72%
7.04 | 36.04% | 32.33% | 26.53% | 20.62%
7.05 | 35.64% | 32.03% | 26.33% | 20.42%
7.06 | 35.34% | 31.73% | 26.13% | 20.22%
7.07 | 34.93% | 31.43% | 25.83% | 20.12%
7.08 | 34.63% | 31.03% | 25.63% | 19.92%
7.09 | 34.33% | 30.73% | 25.43% | 19.82%
7.1 | 34.03% | 30.43% | 25.23% | 19.62%
7.11 | 33.63% | 30.13% | 24.92% | 19.42%
7.12 | 33.33% | 29.83% | 24.72% | 19.22%
7.13 | 33.03% | 29.53% | 24.52% | 19.12%
7.14 | 32.73% | 29.23% | 24.32% | 18.92%
7.15 | 32.43% | 28.93% | 24.12% | 18.72%
7.16 | 32.03% | 28.63% | 23.82% | 18.62%
7.17 | 31.73% | 28.33% | 23.62% | 18.42%
7.18 | 31.43% | 28.03% | 23.42% | 18.22%
7.19 | 31.13% | 27.73% | 23.22% | 18.12%
7.2 | 30.83% | 27.43% | 23.02% | 17.92%
7.21 | 30.53% | 27.13% | 22.82% | 17.72%
7.22 | 30.13% | 26.83% | 22.62% | 17.62%
7.23 | 29.83% | 26.63% | 22.32% | 17.42%
7.24 | 29.53% | 26.33% | 22.12% | 17.32%
7.25 | 29.23% | 26.03% | 21.92% | 17.12%
7.26 | 28.93% | 25.73% | 21.72% | 17.02%
7.27 | 28.63% | 25.43% | 21.52% | 16.82%
7.28 | 28.23% | 25.13% | 21.32% | 16.72%
7.29 | 27.93% | 24.82% | 21.12% | 16.52%
7.3 | 27.63% | 24.52% | 20.92% | 16.42%
7.31 | 27.33% | 24.22% | 20.72% | 16.22%
7.32 | 27.03% | 23.92% | 20.52% | 16.12%
7.33 | 26.73% | 23.72% | 20.32% | 15.92%
7.34 | 26.43% | 23.42% | 20.02% | 15.82%
7.35 | 26.13% | 23.12% | 19.82% | 15.62%
7.36 | 25.83% | 22.82% | 19.62% | 15.52%
7.37 | 25.53% | 22.62% | 19.42% | 15.32%
7.38 | 25.23% | 22.32% | 19.22% | 15.22%
7.39 | 24.92% | 22.12% | 19.02% | 15.02%
7.4 | 24.62% | 21.82% | 18.82% | 14.91%
7.41 | 24.32% | 21.52% | 18.62% | 14.71%
7.42 | 24.02% | 21.32% | 18.42% | 14.61%
7.43 | 23.72% | 21.02% | 18.22% | 14.41%
7.44 | 23.42% | 20.82% | 18.02% | 14.31%
7.45 | 23.22% | 20.52% | 17.82% | 14.11%
7.46 | 22.92% | 20.22% | 17.62% | 14.01%
7.47 | 22.62% | 20.02% | 17.42% | 13.81%
7.48 | 22.32% | 19.72% | 17.22% | 13.71%
7.49 | 22.02% | 19.52% | 17.02% | 13.51%
7.5 | 21.82% | 19.22% | 16.82% | 13.41%
7.51 | 21.52% | 18.92% | 16.62% | 13.31%
7.52 | 21.22% | 18.72% | 16.42% | 13.11%
7.53 | 21.02% | 18.42% | 16.22% | 13.01%
7.54 | 20.72% | 18.22% | 16.02% | 12.81%
7.55 | 20.42% | 17.92% | 15.92% | 12.71%
7.56 | 20.22% | 17.62% | 15.72% | 12.51%
7.57 | 19.92% | 17.42% | 15.52% | 12.41%
7.58 | 19.72% | 17.12% | 15.32% | 12.21%
7.59 | 19.42% | 16.92% | 15.12% | 12.11%
7.6 | 19.22% | 16.72% | 14.91% | 12.01%
7.61 | 18.92% | 16.52% | 14.71% | 11.81%
7.62 | 18.72% | 16.22% | 14.61% | 11.71%
7.63 | 18.42% | 16.02% | 14.41% | 11.61%
7.64 | 18.22% | 15.82% | 14.21% | 11.41%
7.65 | 17.92% | 15.62% | 14.01% | 11.31%
7.66 | 17.72% | 15.32% | 13.81% | 11.21%
7.67 | 17.52% | 15.12% | 13.71% | 11.11%
7.68 | 17.22% | 14.91% | 13.51% | 10.91%
7.69 | 17.02% | 14.71% | 13.31% | 10.81%
7.7 | 16.72% | 14.41% | 13.11% | 10.71%
7.71 | 16.52% | 14.21% | 13.01% | 10.61%
7.72 | 16.32% | 14.01% | 12.81% | 10.51%
7.73 | 16.02% | 13.81% | 12.61% | 10.31%
7.74 | 15.82% | 13.71% | 12.51% | 10.21%
7.75 | 15.62% | 13.51% | 12.31% | 10.11%
7.76 | 15.32% | 13.31% | 12.11% | 10.01%
7.77 | 15.12% | 13.11% | 12.01% | 9.91%
7.78 | 14.91% | 12.91% | 11.81% | 9.71%
7.79 | 14.71% | 12.71% | 11.71% | 9.61%
7.8 | 14.51% | 12.51% | 11.51% | 9.51%
7.81 | 14.31% | 12.31% | 11.31% | 9.409%
7.82 | 14.11% | 12.21% | 11.21% | 9.309%
7.83 | 13.91% | 12.01% | 11.01% | 9.209%
7.84 | 13.71% | 11.81% | 10.91% | 9.109%
7.85 | 13.51% | 11.71% | 10.71% | 9.009%
7.86 | 13.31% | 11.51% | 10.61% | 8.809%
7.87 | 13.11% | 11.31% | 10.51% | 8.709%
7.88 | 12.91% | 11.21% | 10.31% | 8.609%
7.89 | 12.71% | 11.01% | 10.21% | 8.509%
7.9 | 12.51% | 10.91% | 10.01% | 8.408%
7.91 | 12.31% | 10.71% | 9.91% | 8.308%
7.92 | 12.11% | 10.51% | 9.81% | 8.208%
7.93 | 12.01% | 10.41% | 9.61% | 8.108%
7.94 | 11.81% | 10.21% | 9.51% | 8.008%
7.95 | 11.61% | 10.11% | 9.409% | 7.908%
7.96 | 11.51% | 9.91% | 9.309% | 7.808%
7.97 | 11.31% | 9.81% | 9.109% | 7.708%
7.98 | 11.11% | 9.61% | 9.009% | 7.708%
7.99 | 11.01% | 9.51% | 8.909% | 7.608%
8 | 10.81% | 9.309% | 8.809% | 7.508%
8.01 | 10.61% | 9.209% | 8.609% | 7.407%
8.02 | 10.51% | 9.009% | 8.509% | 7.307%
8.03 | 10.31% | 8.909% | 8.408% | 7.207%
8.04 | 10.21% | 8.709% | 8.308% | 7.107%
8.05 | 10.01% | 8.609% | 8.208% | 7.107%
8.06 | 9.91% | 8.509% | 8.108% | 7.007%
8.07 | 9.71% | 8.308% | 8.008% | 6.907%
8.08 | 9.61% | 8.208% | 7.908% | 6.807%
8.09 | 9.51% | 8.108% | 7.808% | 6.707%
8.1 | 9.309% | 8.008% | 7.708% | 6.707%
8.11 | 9.209% | 7.908% | 7.608% | 6.607%
8.12 | 9.109% | 7.808% | 7.508% | 6.507%
8.13 | 8.909% | 7.708% | 7.407% | 6.507%
8.14 | 8.809% | 7.608% | 7.307% | 6.406%
8.15 | 8.709% | 7.508% | 7.207% | 6.306%
8.16 | 8.509% | 7.307% | 7.107% | 6.206%
8.17 | 8.408% | 7.207% | 7.007% | 6.206%
8.18 | 8.308% | 7.107% | 6.907% | 6.106%
8.19 | 8.208% | 7.007% | 6.807% | 6.006%
8.2 | 8.108% | 6.907% | 6.707% | 6.006%
8.21 | 8.008% | 6.907% | 6.707% | 5.906%
8.22 | 7.808% | 6.807% | 6.607% | 5.806%
8.23 | 7.708% | 6.707% | 6.507% | 5.806%
8.24 | 7.608% | 6.607% | 6.406% | 5.706%
8.25 | 7.508% | 6.507% | 6.306% | 5.606%
8.26 | 7.407% | 6.406% | 6.306% | 5.606%
8.27 | 7.307% | 6.306% | 6.206% | 5.506%
8.28 | 7.207% | 6.206% | 6.106% | 5.506%
8.29 | 7.107% | 6.206% | 6.006% | 5.405%
8.3 | 7.007% | 6.106% | 6.006% | 5.305%
8.31 | 6.907% | 6.006% | 5.906% | 5.305%
8.32 | 6.907% | 5.906% | 5.806% | 5.205%
8.33 | 6.807% | 5.906% | 5.706% | 5.205%
8.34 | 6.707% | 5.806% | 5.706% | 5.105%
8.35 | 6.607% | 5.706% | 5.606% | 5.105%
8.36 | 6.507% | 5.706% | 5.506% | 5.005%
8.37 | 6.507% | 5.606% | 5.506% | 5.005%
8.38 | 6.406% | 5.606% | 5.405% | 4.905%
8.39 | 6.306% | 5.506% | 5.405% | 4.905%
8.4 | 6.206% | 5.405% | 5.305% | 4.805%
8.41 | 6.206% | 5.405% | 5.205% | 4.805%
8.42 | 6.106% | 5.305% | 5.205% | 4.805%
8.43 | 6.006% | 5.305% | 5.105% | 4.705%
8.44 | 6.006% | 5.205% | 5.105% | 4.705%
8.45 | 5.906% | 5.205% | 5.005% | 4.605%
8.46 | 5.806% | 5.105% | 5.005% | 4.605%
8.47 | 5.806% | 5.105% | 4.905% | 4.605%
8.48 | 5.706% | 5.005% | 4.905% | 4.505%
8.49 | 5.706% | 5.005% | 4.805% | 4.505%
8.5 | 5.606% | 4.905% | 4.805% | 4.505%
8.51 | 5.506% | 4.905% | 4.705% | 4.404%
8.52 | 5.506% | 4.805% | 4.705% | 4.404%
8.53 | 5.405% | 4.805% | 4.605% | 4.404%
8.54 | 5.405% | 4.705% | 4.605% | 4.304%
8.55 | 5.305% | 4.705% | 4.605% | 4.304%
8.56 | 5.305% | 4.605% | 4.505% | 4.304%
8.57 | 5.205% | 4.605% | 4.505% | 4.204%
8.58 | 5.205% | 4.605% | 4.404% | 4.204%
8.59 | 5.205% | 4.505% | 4.404% | 4.204%
8.6 | 5.105% | 4.505% | 4.404% | 4.204%
8.61 | 5.105% | 4.505% | 4.304% | 4.104%
8.62 | 5.005% | 4.404% | 4.304% | 4.104%
8.63 | 5.005% | 4.404% | 4.304% | 4.104%
8.64 | 5.005% | 4.404% | 4.204% | 4.104%
8.65 | 4.905% | 4.304% | 4.204% | 4.004%
8.66 | 4.905% | 4.304% | 4.204% | 4.004%
8.67 | 4.905% | 4.304% | 4.104% | 4.004%
8.68 | 4.805% | 4.204% | 4.104% | 4.004%
8.69 | 4.805% | 4.204% | 4.104% | 4.004%
8.7 | 4.805% | 4.204% | 4.104% | 3.904%
8.71 | 4.705% | 4.204% | 4.004% | 3.904%
8.72 | 4.705% | 4.204% | 4.004% | 3.904%
8.73 | 4.705% | 4.104% | 4.004% | 3.904%
8.74 | 4.605% | 4.104% | 4.004% | 3.904%
8.75 | 4.605% | 4.104% | 3.904% | 3.904%
8.76 | 4.605% | 4.104% | 3.904% | 3.804%
8.77 | 4.605% | 4.104% | 3.904% | 3.804%
8.78 | 4.505% | 4.004% | 3.904% | 3.804%
8.79 | 4.505% | 4.004% | 3.904% | 3.804%
8.8 | 4.505% | 4.004% | 3.804% | 3.804%
8.81 | 4.505% | 4.004% | 3.804% | 3.804%
8.82 | 4.505% | 4.004% | 3.804% | 3.704%
8.83 | 4.404% | 4.004% | 3.804% | 3.704%
8.84 | 4.404% | 3.904% | 3.804% | 3.704%
8.85 | 4.404% | 3.904% | 3.804% | 3.704%
8.86 | 4.404% | 3.904% | 3.704% | 3.704%
8.87 | 4.404% | 3.904% | 3.704% | 3.704%
8.88 | 4.304% | 3.904% | 3.704% | 3.704%
8.89 | 4.304% | 3.904% | 3.704% | 3.604%
8.9 | 4.304% | 3.804% | 3.704% | 3.604%
8.91 | 4.304% | 3.804% | 3.704% | 3.604%
8.92 | 4.304% | 3.804% | 3.704% | 3.604%
8.93 | 4.304% | 3.804% | 3.604% | 3.604%
8.94 | 4.304% | 3.804% | 3.604% | 3.604%
8.95 | 4.204% | 3.804% | 3.604% | 3.604%
8.96 | 4.204% | 3.804% | 3.604% | 3.604%
8.97 | 4.204% | 3.804% | 3.604% | 3.604%
8.98 | 4.204% | 3.804% | 3.604% | 3.504%
8.99 | 4.204% | 3.804% | 3.604% | 3.504%
9 | 4.204% | 3.804% | 3.604% | 3.504%
9.01 | 4.204% | 3.804% | 3.604% | 3.504%
9.02 | 4.204% | 3.704% | 3.604% | 3.504%
9.03 | 4.204% | 3.704% | 3.604% | 3.504%
9.04 | 4.204% | 3.704% | 3.504% | 3.504%
9.05 | 4.204% | 3.704% | 3.504% | 3.504%
9.06 | 4.104% | 3.704% | 3.504% | 3.504%
9.07 | 4.104% | 3.704% | 3.504% | 3.504%
9.08 | 4.104% | 3.704% | 3.504% | 3.504%
9.09 | 4.104% | 3.704% | 3.504% | 3.504%
9.1 | 4.104% | 3.704% | 3.504% | 3.504%
9.11 | 4.104% | 3.704% | 3.504% | 3.504%
9.12 | 4.104% | 3.704% | 3.504% | 3.504%
9.13 | 4.104% | 3.704% | 3.504% | 3.504%
9.14 | 4.104% | 3.704% | 3.504% | 3.504%
9.15 | 4.104% | 3.704% | 3.504% | 3.403%
9.16 | 4.104% | 3.704% | 3.504% | 3.403%
9.17 | 4.104% | 3.704% | 3.504% | 3.403%
9.18 | 4.104% | 3.704% | 3.504% | 3.403%
9.19 | 4.104% | 3.704% | 3.403% | 3.403%
9.2 | 4.104% | 3.704% | 3.403% | 3.403%
9.21 | 4.004% | 3.704% | 3.403% | 3.403%
9.22 | 4.004% | 3.604% | 3.403% | 3.403%
9.23 | 4.004% | 3.604% | 3.403% | 3.403%
9.24 | 4.004% | 3.604% | 3.403% | 3.403%
9.25 | 4.004% | 3.604% | 3.403% | 3.403%
9.26 | 4.004% | 3.604% | 3.403% | 3.403%
9.27 | 4.004% | 3.604% | 3.403% | 3.403%
9.28 | 4.004% | 3.604% | 3.403% | 3.403%
9.29 | 4.004% | 3.604% | 3.403% | 3.403%
9.3 | 4.004% | 3.604% | 3.403% | 3.403%
9.31 | 4.004% | 3.604% | 3.403% | 3.403%
9.32 | 4.004% | 3.604% | 3.403% | 3.403%
9.33 | 4.004% | 3.604% | 3.403% | 3.403%
9.34 | 4.004% | 3.604% | 3.403% | 3.403%
9.35 | 4.004% | 3.604% | 3.403% | 3.403%
9.36 | 4.004% | 3.604% | 3.403% | 3.403%
9.37 | 4.004% | 3.604% | 3.403% | 3.303%
9.38 | 4.004% | 3.604% | 3.403% | 3.303%
9.39 | 4.004% | 3.604% | 3.403% | 3.303%
9.4 | 4.004% | 3.604% | 3.403% | 3.303%
9.41 | 4.004% | 3.604% | 3.403% | 3.303%
9.42 | 4.004% | 3.604% | 3.403% | 3.303%
9.43 | 3.904% | 3.604% | 3.403% | 3.303%
9.44 | 3.904% | 3.604% | 3.403% | 3.303%
9.45 | 3.904% | 3.604% | 3.403% | 3.303%
9.46 | 3.904% | 3.604% | 3.403% | 3.303%
9.47 | 3.904% | 3.604% | 3.303% | 3.303%
9.48 | 3.904% | 3.504% | 3.303% | 3.303%
9.49 | 3.904% | 3.504% | 3.303% | 3.303%
9.5 | 3.904% | 3.504% | 3.303% | 3.303%
9.51 | 3.904% | 3.504% | 3.303% | 3.303%
9.52 | 3.904% | 3.504% | 3.303% | 3.303%
9.53 | 3.904% | 3.504% | 3.303% | 3.303%
9.54 | 3.904% | 3.504% | 3.303% | 3.303%
9.55 | 3.904% | 3.504% | 3.303% | 3.303%
9.56 | 3.904% | 3.504% | 3.303% | 3.303%
9.57 | 3.904% | 3.504% | 3.303% | 3.303%
9.58 | 3.904% | 3.504% | 3.303% | 3.303%
9.59 | 3.904% | 3.504% | 3.303% | 3.303%
9.6 | 3.904% | 3.504% | 3.303% | 3.303%
9.61 | 3.904% | 3.504% | 3.303% | 3.303%
9.62 | 3.904% | 3.504% | 3.303% | 3.303%
9.63 | 3.904% | 3.504% | 3.303% | 3.303%
9.64 | 3.904% | 3.504% | 3.303% | 3.303%
9.65 | 3.904% | 3.504% | 3.303% | 3.303%
9.66 | 3.904% | 3.504% | 3.303% | 3.303%
9.67 | 3.904% | 3.504% | 3.303% | 3.303%
9.68 | 3.904% | 3.504% | 3.303% | 3.303%
9.69 | 3.904% | 3.504% | 3.303% | 3.303%
9.7 | 3.904% | 3.504% | 3.303% | 3.303%
9.71 | 3.904% | 3.504% | 3.303% | 3.303%
9.72 | 3.904% | 3.504% | 3.303% | 3.303%
9.73 | 3.904% | 3.504% | 3.303% | 3.303%
9.74 | 3.904% | 3.504% | 3.303% | 3.303%
9.75 | 3.904% | 3.504% | 3.303% | 3.203%
9.76 | 3.904% | 3.504% | 3.303% | 3.303%
9.77 | 3.904% | 3.504% | 3.203% | 3.303%
9.78 | 3.904% | 3.504% | 3.203% | 3.303%
9.79 | 3.904% | 3.504% | 3.203% | 3.303%
9.8 | 3.904% | 3.504% | 3.203% | 3.303%
9.81 | 3.904% | 3.504% | 3.203% | 3.203%
9.82 | 3.904% | 3.403% | 3.203% | 3.203%
9.83 | 3.904% | 3.403% | 3.203% | 3.203%
9.84 | 3.904% | 3.403% | 3.203% | 3.203%
9.85 | 3.904% | 3.403% | 3.203% | 3.203%
9.86 | 3.904% | 3.403% | 3.203% | 3.203%
9.87 | 3.904% | 3.403% | 3.203% | 3.203%
9.88 | 3.804% | 3.403% | 3.203% | 3.203%
9.89 | 3.804% | 3.403% | 3.203% | 3.203%
9.9 | 3.804% | 3.403% | 3.203% | 3.203%
9.91 | 3.804% | 3.403% | 3.203% | 3.203%
9.92 | 3.804% | 3.403% | 3.203% | 3.203%
9.93 | 3.804% | 3.403% | 3.203% | 3.203%
9.94 | 3.804% | 3.403% | 3.203% | 3.203%
9.95 | 3.804% | 3.403% | 3.203% | 3.203%
9.96 | 3.804% | 3.403% | 3.203% | 3.203%
9.97 | 3.804% | 3.403% | 3.203% | 3.203%
9.98 | 3.804% | 3.403% | 3.203% | 3.203%
9.99 | 3.804% | 3.403% | 3.203% | 3.203%
10 | 3.804% | 3.403% | 3.203% | 3.203%
10.01 | 3.804% | 3.403% | 3.203% | 3.203%
10.02 | 3.804% | 3.403% | 3.203% | 3.203%
10.03 | 3.804% | 3.403% | 3.203% | 3.203%
10.04 | 3.804% | 3.403% | 3.203% | 3.203%
10.05 | 3.804% | 3.403% | 3.203% | 3.203%
10.06 | 3.804% | 3.403% | 3.203% | 3.203%
10.07 | 3.804% | 3.403% | 3.203% | 3.203%
10.08 | 3.804% | 3.403% | 3.203% | 3.203%
10.09 | 3.804% | 3.403% | 3.203% | 3.203%
10.1 | 3.804% | 3.403% | 3.203% | 3.203%
10.11 | 3.804% | 3.403% | 3.203% | 3.203%
10.12 | 3.804% | 3.403% | 3.203% | 3.203%
10.13 | 3.804% | 3.403% | 3.203% | 3.203%
10.14 | 3.804% | 3.403% | 3.203% | 3.203%
10.15 | 3.804% | 3.403% | 3.203% | 3.203%
10.16 | 3.804% | 3.403% | 3.203% | 3.203%
10.17 | 3.804% | 3.403% | 3.203% | 3.203%
10.18 | 3.804% | 3.403% | 3.203% | 3.203%
10.19 | 3.804% | 3.403% | 3.203% | 3.203%
10.2 | 3.804% | 3.403% | 3.103% | 3.203%
10.21 | 3.804% | 3.403% | 3.103% | 3.203%
10.22 | 3.804% | 3.403% | 3.103% | 3.203%
10.23 | 3.804% | 3.403% | 3.103% | 3.203%
10.24 | 3.804% | 3.303% | 3.103% | 3.103%
10.25 | 3.804% | 3.303% | 3.103% | 3.103%
10.26 | 3.804% | 3.303% | 3.103% | 3.103%
10.27 | 3.804% | 3.303% | 3.103% | 3.103%
10.28 | 3.804% | 3.303% | 3.103% | 3.103%
10.29 | 3.804% | 3.303% | 3.103% | 3.103%
10.3 | 3.704% | 3.303% | 3.103% | 3.103%
10.31 | 3.704% | 3.303% | 3.103% | 3.103%
10.32 | 3.704% | 3.303% | 3.103% | 3.103%
10.33 | 3.704% | 3.303% | 3.103% | 3.103%
10.34 | 3.704% | 3.303% | 3.103% | 3.103%
10.35 | 3.704% | 3.303% | 3.103% | 3.103%
10.36 | 3.704% | 3.303% | 3.103% | 3.103%
10.37 | 3.704% | 3.303% | 3.103% | 3.103%
10.38 | 3.704% | 3.303% | 3.103% | 3.103%
10.39 | 3.704% | 3.303% | 3.103% | 3.103%
10.4 | 3.704% | 3.303% | 3.103% | 3.103%
10.41 | 3.704% | 3.303% | 3.103% | 3.103%
10.42 | 3.704% | 3.303% | 3.103% | 3.103%
10.43 | 3.704% | 3.303% | 3.103% | 3.103%
10.44 | 3.704% | 3.303% | 3.103% | 3.103%
10.45 | 3.704% | 3.303% | 3.103% | 3.103%
10.46 | 3.704% | 3.303% | 3.103% | 3.103%
10.47 | 3.704% | 3.303% | 3.103% | 3.103%
10.48 | 3.704% | 3.303% | 3.103% | 3.103%
10.49 | 3.704% | 3.303% | 3.103% | 3.103%
10.5 | 3.704% | 3.303% | 3.103% | 3.103%
10.51 | 3.704% | 3.303% | 3.103% | 3.103%
10.52 | 3.704% | 3.303% | 3.103% | 3.103%
10.53 | 3.704% | 3.303% | 3.103% | 3.103%
10.54 | 3.704% | 3.303% | 3.103% | 3.103%
10.55 | 3.704% | 3.303% | 3.103% | 3.103%
10.56 | 3.704% | 3.303% | 3.103% | 3.103%
10.57 | 3.704% | 3.303% | 3.103% | 3.103%
10.58 | 3.704% | 3.303% | 3.103% | 3.103%
10.59 | 3.704% | 3.203% | 3.103% | 3.103%
10.6 | 3.704% | 3.203% | 3.103% | 3.103%
10.61 | 3.704% | 3.203% | 3.103% | 3.003%
10.62 | 3.704% | 3.203% | 3.103% | 3.003%
10.63 | 3.704% | 3.203% | 3.103% | 3.003%
10.64 | 3.704% | 3.203% | 3.103% | 3.003%
10.65 | 3.704% | 3.203% | 3.103% | 3.003%
10.66 | 3.704% | 3.203% | 3.103% | 3.003%
10.67 | 3.704% | 3.203% | 3.103% | 3.003%
10.68 | 3.704% | 3.203% | 3.103% | 3.003%
10.69 | 3.704% | 3.203% | 3.103% | 3.003%
10.7 | 3.704% | 3.203% | 3.103% | 3.003%
10.71 | 3.704% | 3.203% | 3.103% | 3.003%
10.72 | 3.704% | 3.203% | 3.103% | 3.003%
10.73 | 3.604% | 3.203% | 3.103% | 3.003%
10.74 | 3.604% | 3.203% | 3.103% | 3.003%
10.75 | 3.604% | 3.203% | 3.103% | 3.003%
10.76 | 3.604% | 3.203% | 3.003% | 3.003%
10.77 | 3.604% | 3.203% | 3.003% | 3.003%
10.78 | 3.604% | 3.203% | 3.003% | 3.003%
10.79 | 3.604% | 3.203% | 3.003% | 3.003%
10.8 | 3.604% | 3.203% | 3.003% | 3.003%
10.81 | 3.604% | 3.203% | 3.003% | 3.003%
10.82 | 3.604% | 3.203% | 3.003% | 3.003%
10.83 | 3.604% | 3.203% | 3.003% | 3.003%
10.84 | 3.604% | 3.203% | 3.003% | 3.003%
10.85 | 3.604% | 3.203% | 3.003% | 3.003%
10.86 | 3.604% | 3.203% | 3.003% | 3.003%
10.87 | 3.604% | 3.203% | 3.003% | 3.003%
10.88 | 3.604% | 3.203% | 3.003% | 3.003%
10.89 | 3.604% | 3.203% | 3.003% | 3.003%
10.9 | 3.604% | 3.203% | 3.003% | 3.003%
10.91 | 3.604% | 3.203% | 3.003% | 3.003%
10.92 | 3.604% | 3.203% | 3.003% | 3.003%
10.93 | 3.604% | 3.203% | 3.003% | 3.003%
10.94 | 3.604% | 3.203% | 3.003% | 3.003%
10.95 | 3.604% | 3.203% | 3.003% | 3.003%
10.96 | 3.604% | 3.203% | 3.003% | 3.003%
10.97 | 3.604% | 3.203% | 3.003% | 3.003%
10.98 | 3.604% | 3.203% | 3.003% | 3.003%
10.99 | 3.604% | 3.203% | 3.003% | 3.003%
11 | 3.604% | 3.203% | 3.003% | 3.003%
11.01 | 3.604% | 3.203% | 3.003% | 3.003%
11.02 | 3.604% | 3.103% | 3.003% | 3.003%
11.03 | 3.604% | 3.103% | 3.003% | 3.003%
11.04 | 3.604% | 3.103% | 3.003% | 3.003%
11.05 | 3.604% | 3.103% | 3.003% | 3.003%
11.06 | 3.604% | 3.103% | 3.003% | 3.003%
11.07 | 3.604% | 3.103% | 3.003% | 3.003%
11.08 | 3.604% | 3.103% | 3.003% | 3.003%
11.09 | 3.604% | 3.103% | 3.003% | 2.903%
11.1 | 3.604% | 3.103% | 3.003% | 2.903%
11.11 | 3.604% | 3.103% | 3.003% | 2.903%
11.12 | 3.604% | 3.103% | 3.003% | 2.903%
11.13 | 3.604% | 3.103% | 3.003% | 2.903%
11.14 | 3.604% | 3.103% | 3.003% | 2.903%
11.15 | 3.604% | 3.103% | 3.003% | 2.903%
11.16 | 3.604% | 3.103% | 3.003% | 2.903%
11.17 | 3.604% | 3.103% | 3.003% | 2.903%
11.18 | 3.604% | 3.103% | 3.003% | 2.903%
11.19 | 3.504% | 3.103% | 2.903% | 2.903%
11.2 | 3.504% | 3.103% | 2.903% | 2.903%
11.21 | 3.504% | 3.103% | 2.903% | 2.903%
11.22 | 3.504% | 3.103% | 2.903% | 2.903%
11.23 | 3.504% | 3.103% | 2.903% | 2.903%
11.24 | 3.504% | 3.103% | 2.903% | 2.903%
11.25 | 3.504% | 3.103% | 2.903% | 2.903%
11.26 | 3.504% | 3.103% | 2.903% | 2.903%
11.27 | 3.504% | 3.103% | 2.903% | 2.903%
11.28 | 3.504% | 3.103% | 2.903% | 2.903%
11.29 | 3.504% | 3.103% | 2.903% | 2.903%
11.3 | 3.504% | 3.103% | 2.903% | 2.903%
11.31 | 3.504% | 3.103% | 2.903% | 2.903%
11.32 | 3.504% | 3.103% | 2.903% | 2.903%
11.33 | 3.504% | 3.103% | 2.903% | 2.903%
11.34 | 3.504% | 3.103% | 2.903% | 2.903%
11.35 | 3.504% | 3.103% | 2.903% | 2.903%
11.36 | 3.504% | 3.103% | 2.903% | 2.903%
11.37 | 3.504% | 3.103% | 2.903% | 2.903%
11.38 | 3.504% | 3.103% | 2.903% | 2.903%
11.39 | 3.504% | 3.103% | 2.903% | 2.903%
11.4 | 3.504% | 3.103% | 2.903% | 2.903%
11.41 | 3.504% | 3.103% | 2.903% | 2.903%
11.42 | 3.504% | 3.103% | 2.903% | 2.903%
11.43 | 3.504% | 3.103% | 2.903% | 2.903%
11.44 | 3.504% | 3.103% | 2.903% | 2.903%
11.45 | 3.504% | 3.103% | 2.903% | 2.903%
11.46 | 3.504% | 3.103% | 2.903% | 2.903%
11.47 | 3.504% | 3.103% | 2.903% | 2.903%
11.48 | 3.504% | 3.103% | 2.903% | 2.903%
11.49 | 3.504% | 3.103% | 2.903% | 2.903%
11.5 | 3.504% | 3.103% | 2.903% | 2.903%
11.51 | 3.504% | 3.103% | 2.903% | 2.903%
11.52 | 3.504% | 3.103% | 2.903% | 2.903%
11.53 | 3.504% | 3.103% | 2.903% | 2.903%
11.54 | 3.504% | 3.103% | 2.903% | 2.903%
11.55 | 3.504% | 3.103% | 2.903% | 2.903%
11.56 | 3.504% | 3.103% | 2.903% | 2.903%
11.57 | 3.504% | 3.103% | 2.903% | 2.903%
11.58 | 3.504% | 3.103% | 2.903% | 2.903%
11.59 | 3.504% | 3.103% | 2.903% | 2.903%
11.6 | 3.504% | 3.103% | 2.903% | 2.903%
11.61 | 3.504% | 3.103% | 2.903% | 2.903%
11.62 | 3.504% | 3.103% | 2.903% | 2.903%
11.63 | 3.504% | 3.103% | 2.903% | 2.903%
11.64 | 3.504% | 3.003% | 2.903% | 2.903%
11.65 | 3.504% | 3.003% | 2.903% | 2.803%
11.66 | 3.504% | 3.003% | 2.903% | 2.803%
11.67 | 3.504% | 3.003% | 2.803% | 2.803%
11.68 | 3.504% | 3.003% | 2.803% | 2.803%
11.69 | 3.403% | 3.003% | 2.803% | 2.803%
11.7 | 3.403% | 3.003% | 2.803% | 2.803%
11.71 | 3.403% | 3.003% | 2.803% | 2.803%
11.72 | 3.403% | 3.003% | 2.803% | 2.803%
11.73 | 3.403% | 3.003% | 2.803% | 2.803%
11.74 | 3.403% | 3.003% | 2.803% | 2.803%
11.75 | 3.403% | 3.003% | 2.803% | 2.803%
11.76 | 3.403% | 3.003% | 2.803% | 2.803%
11.77 | 3.403% | 3.003% | 2.803% | 2.803%
11.78 | 3.403% | 3.003% | 2.803% | 2.803%
11.79 | 3.403% | 3.003% | 2.803% | 2.803%
11.8 | 3.403% | 3.003% | 2.803% | 2.803%
11.81 | 3.403% | 3.003% | 2.803% | 2.803%
11.82 | 3.403% | 3.003% | 2.803% | 2.803%
11.83 | 3.403% | 3.003% | 2.803% | 2.803%
11.84 | 3.403% | 3.003% | 2.803% | 2.803%
11.85 | 3.403% | 3.003% | 2.803% | 2.803%
11.86 | 3.403% | 3.003% | 2.803% | 2.803%
11.87 | 3.403% | 3.003% | 2.803% | 2.803%
11.88 | 3.403% | 3.003% | 2.803% | 2.803%
11.89 | 3.403% | 3.003% | 2.803% | 2.803%
11.9 | 3.403% | 3.003% | 2.803% | 2.803%
11.91 | 3.403% | 3.003% | 2.803% | 2.803%
11.92 | 3.403% | 3.003% | 2.803% | 2.803%
11.93 | 3.403% | 3.003% | 2.803% | 2.803%
11.94 | 3.403% | 3.003% | 2.803% | 2.803%
11.95 | 3.403% | 3.003% | 2.803% | 2.803%
11.96 | 3.403% | 3.003% | 2.803% | 2.803%
11.97 | 3.403% | 3.003% | 2.803% | 2.803%
11.98 | 3.403% | 3.003% | 2.803% | 2.803%
11.99 | 3.403% | 3.003% | 2.803% | 2.803%
12 | 3.403% | 3.003% | 2.803% | 2.803%
12.01 | 3.403% | 3.003% | 2.803% | 2.803%
12.02 | 3.403% | 3.003% | 2.803% | 2.803%
12.03 | 3.403% | 3.003% | 2.703% | 2.803%
12.04 | 3.403% | 3.003% | 2.703% | 2.803%
12.05 | 3.403% | 3.003% | 2.703% | 2.803%
12.06 | 3.403% | 3.003% | 2.703% | 2.803%
12.07 | 3.403% | 3.003% | 2.703% | 2.803%
12.08 | 3.403% | 2.903% | 2.703% | 2.803%
12.09 | 3.403% | 2.903% | 2.703% | 2.803%
12.1 | 3.403% | 2.903% | 2.703% | 2.803%
12.11 | 3.403% | 2.903% | 2.703% | 2.803%
12.12 | 3.403% | 2.903% | 2.703% | 2.803%
12.13 | 3.403% | 2.903% | 2.703% | 2.803%
12.14 | 3.403% | 2.903% | 2.703% | 2.803%
12.15 | 3.403% | 2.903% | 2.703% | 2.803%
12.16 | 3.403% | 2.903% | 2.703% | 2.803%
12.17 | 3.403% | 2.903% | 2.703% | 2.803%
12.18 | 3.403% | 2.903% | 2.703% | 2.703%
12.19 | 3.403% | 2.903% | 2.703% | 2.703%
12.2 | 3.303% | 2.903% | 2.703% | 2.703%
12.21 | 3.303% | 2.903% | 2.703% | 2.703%
12.22 | 3.303% | 2.903% | 2.703% | 2.703%
12.23 | 3.303% | 2.903% | 2.703% | 2.703%
12.24 | 3.303% | 2.903% | 2.703% | 2.703%
12.25 | 3.303% | 2.903% | 2.703% | 2.703%
12.26 | 3.303% | 2.903% | 2.703% | 2.703%
12.27 | 3.303% | 2.903% | 2.703% | 2.703%
12.28 | 3.303% | 2.903% | 2.703% | 2.703%
12.29 | 3.303% | 2.903% | 2.703% | 2.703%
12.3 | 3.303% | 2.903% | 2.703% | 2.703%
12.31 | 3.303% | 2.903% | 2.703% | 2.703%
12.32 | 3.303% | 2.903% | 2.703% | 2.703%
12.33 | 3.303% | 2.903% | 2.703% | 2.703%
12.34 | 3.303% | 2.903% | 2.703% | 2.703%
12.35 | 3.303% | 2.903% | 2.703% | 2.703%
12.36 | 3.303% | 2.903% | 2.703% | 2.703%
12.37 | 3.303% | 2.903% | 2.703% | 2.703%
12.38 | 3.303% | 2.903% | 2.703% | 2.703%
12.39 | 3.303% | 2.903% | 2.703% | 2.703%
12.4 | 3.303% | 2.903% | 2.703% | 2.703%
12.41 | 3.303% | 2.903% | 2.703% | 2.703%
12.42 | 3.303% | 2.903% | 2.703% | 2.703%
12.43 | 3.303% | 2.903% | 2.703% | 2.703%
12.44 | 3.303% | 2.903% | 2.703% | 2.703%
12.45 | 3.303% | 2.903% | 2.703% | 2.703%
12.46 | 3.303% | 2.903% | 2.703% | 2.703%
12.47 | 3.303% | 2.903% | 2.703% | 2.703%
12.48 | 3.303% | 2.903% | 2.703% | 2.703%
12.49 | 3.303% | 2.903% | 2.703% | 2.703%
12.5 | 3.303% | 2.903% | 2.703% | 2.703%
12.51 | 3.303% | 2.903% | 2.703% | 2.703%
12.52 | 3.303% | 2.903% | 2.703% | 2.703%
12.53 | 3.303% | 2.903% | 2.703% | 2.703%
12.54 | 3.303% | 2.903% | 2.703% | 2.703%
12.55 | 3.303% | 2.903% | 2.703% | 2.703%
12.56 | 3.303% | 2.903% | 2.703% | 2.703%
12.57 | 3.303% | 2.903% | 2.703% | 2.703%
12.58 | 3.303% | 2.903% | 2.703% | 2.703%
12.59 | 3.303% | 2.903% | 2.703% | 2.703%
12.6 | 3.303% | 2.903% | 2.703% | 2.703%
12.61 | 3.303% | 2.903% | 2.703% | 2.703%
12.62 | 3.303% | 2.903% | 2.703% | 2.703%
12.63 | 3.303% | 2.903% | 2.703% | 2.703%
12.64 | 3.303% | 2.903% | 2.703% | 2.703%
12.65 | 3.303% | 2.903% | 2.703% | 2.703%
12.66 | 3.303% | 2.803% | 2.703% | 2.703%
12.67 | 3.303% | 2.803% | 2.703% | 2.703%
12.68 | 3.303% | 2.803% | 2.703% | 2.703%
12.69 | 3.303% | 2.803% | 2.703% | 2.703%
12.7 | 3.303% | 2.803% | 2.703% | 2.703%
12.71 | 3.303% | 2.803% | 2.703% | 2.703%
12.72 | 3.303% | 2.803% | 2.703% | 2.703%
12.73 | 3.203% | 2.803% | 2.703% | 2.703%
12.74 | 3.203% | 2.803% | 2.703% | 2.703%
12.75 | 3.203% | 2.803% | 2.703% | 2.703%
12.76 | 3.203% | 2.803% | 2.703% | 2.703%
12.77 | 3.203% | 2.803% | 2.703% | 2.703%
12.78 | 3.203% | 2.803% | 2.703% | 2.703%
12.79 | 3.203% | 2.803% | 2.603% | 2.703%
12.8 | 3.203% | 2.803% | 2.603% | 2.703%
12.81 | 3.203% | 2.803% | 2.603% | 2.703%
12.82 | 3.203% | 2.803% | 2.603% | 2.703%
12.83 | 3.203% | 2.803% | 2.603% | 2.703%
12.84 | 3.203% | 2.803% | 2.603% | 2.703%
12.85 | 3.203% | 2.803% | 2.603% | 2.703%
12.86 | 3.203% | 2.803% | 2.603% | 2.703%
12.87 | 3.203% | 2.803% | 2.603% | 2.603%
12.88 | 3.203% | 2.803% | 2.603% | 2.603%
12.89 | 3.203% | 2.803% | 2.603% | 2.603%
12.9 | 3.203% | 2.803% | 2.603% | 2.603%
12.91 | 3.203% | 2.803% | 2.603% | 2.603%
12.92 | 3.203% | 2.803% | 2.603% | 2.603%
12.93 | 3.203% | 2.803% | 2.603% | 2.603%
12.94 | 3.203% | 2.803% | 2.603% | 2.603%
12.95 | 3.203% | 2.803% | 2.603% | 2.603%
12.96 | 3.203% | 2.803% | 2.603% | 2.603%
12.97 | 3.203% | 2.803% | 2.603% | 2.603%
12.98 | 3.203% | 2.803% | 2.603% | 2.603%
12.99 | 3.203% | 2.803% | 2.603% | 2.603%
13 | 3.203% | 2.803% | 2.603% | 2.603%
13.01 | 3.203% | 2.803% | 2.603% | 2.603%
13.02 | 3.203% | 2.803% | 2.603% | 2.603%
13.03 | 3.203% | 2.803% | 2.603% | 2.603%
13.04 | 3.203% | 2.803% | 2.603% | 2.603%
13.05 | 3.203% | 2.803% | 2.603% | 2.603%
13.06 | 3.203% | 2.803% | 2.603% | 2.603%
13.07 | 3.203% | 2.803% | 2.603% | 2.603%
13.08 | 3.203% | 2.803% | 2.603% | 2.603%
13.09 | 3.203% | 2.803% | 2.603% | 2.603%
13.1 | 3.203% | 2.703% | 2.603% | 2.603%
13.11 | 3.203% | 2.703% | 2.603% | 2.603%
13.12 | 3.203% | 2.703% | 2.603% | 2.603%
13.13 | 3.203% | 2.703% | 2.603% | 2.603%
13.14 | 3.203% | 2.703% | 2.603% | 2.603%
13.15 | 3.203% | 2.703% | 2.603% | 2.603%
13.16 | 3.203% | 2.703% | 2.603% | 2.603%
13.17 | 3.203% | 2.703% | 2.603% | 2.603%
13.18 | 3.203% | 2.703% | 2.603% | 2.603%
13.19 | 3.203% | 2.703% | 2.603% | 2.603%
13.2 | 3.203% | 2.703% | 2.603% | 2.603%
13.21 | 3.203% | 2.703% | 2.603% | 2.603%
13.22 | 3.203% | 2.703% | 2.603% | 2.603%
13.23 | 3.203% | 2.703% | 2.603% | 2.603%
13.24 | 3.203% | 2.703% | 2.503% | 2.603%
13.25 | 3.203% | 2.703% | 2.503% | 2.603%
13.26 | 3.203% | 2.703% | 2.503% | 2.603%
13.27 | 3.203% | 2.703% | 2.503% | 2.603%
13.28 | 3.203% | 2.703% | 2.503% | 2.603%
13.29 | 3.103% | 2.703% | 2.503% | 2.603%
13.3 | 3.103% | 2.703% | 2.503% | 2.603%
13.31 | 3.103% | 2.703% | 2.503% | 2.603%
13.32 | 3.103% | 2.703% | 2.503% | 2.603%
13.33 | 3.103% | 2.703% | 2.503% | 2.603%
13.34 | 3.103% | 2.703% | 2.503% | 2.603%
13.35 | 3.103% | 2.703% | 2.503% | 2.603%
13.36 | 3.103% | 2.703% | 2.503% | 2.603%
13.37 | 3.103% | 2.703% | 2.503% | 2.603%
13.38 | 3.103% | 2.703% | 2.503% | 2.603%
13.39 | 3.103% | 2.703% | 2.503% | 2.603%
13.4 | 3.103% | 2.703% | 2.503% | 2.603%
13.41 | 3.103% | 2.703% | 2.503% | 2.603%
13.42 | 3.103% | 2.703% | 2.503% | 2.603%
13.43 | 3.103% | 2.703% | 2.503% | 2.603%
13.44 | 3.103% | 2.703% | 2.503% | 2.503%
13.45 | 3.103% | 2.703% | 2.503% | 2.503%
13.46 | 3.103% | 2.703% | 2.503% | 2.503%
13.47 | 3.103% | 2.703% | 2.503% | 2.503%
13.48 | 3.103% | 2.703% | 2.503% | 2.503%
13.49 | 3.103% | 2.703% | 2.503% | 2.503%
13.5 | 3.103% | 2.703% | 2.503% | 2.503%
13.51 | 3.103% | 2.703% | 2.503% | 2.503%
13.52 | 3.103% | 2.703% | 2.503% | 2.503%
13.53 | 3.103% | 2.703% | 2.503% | 2.503%
13.54 | 3.103% | 2.703% | 2.503% | 2.503%
13.55 | 3.103% | 2.703% | 2.503% | 2.503%
13.56 | 3.103% | 2.703% | 2.503% | 2.503%
13.57 | 3.103% | 2.703% | 2.503% | 2.503%
13.58 | 3.103% | 2.703% | 2.503% | 2.503%
13.59 | 3.103% | 2.703% | 2.503% | 2.503%
13.6 | 3.103% | 2.703% | 2.503% | 2.503%
13.61 | 3.103% | 2.703% | 2.503% | 2.503%
13.62 | 3.103% | 2.703% | 2.503% | 2.503%
13.63 | 3.103% | 2.703% | 2.503% | 2.503%
13.64 | 3.103% | 2.703% | 2.503% | 2.503%
13.65 | 3.103% | 2.703% | 2.503% | 2.503%
13.66 | 3.103% | 2.703% | 2.503% | 2.503%
13.67 | 3.103% | 2.703% | 2.503% | 2.503%
13.68 | 3.103% | 2.703% | 2.503% | 2.503%
13.69 | 3.103% | 2.703% | 2.503% | 2.503%
13.7 | 3.103% | 2.703% | 2.503% | 2.503%
13.71 | 3.103% | 2.703% | 2.503% | 2.503%
13.72 | 3.103% | 2.703% | 2.503% | 2.503%
13.73 | 3.103% | 2.603% | 2.503% | 2.503%
13.74 | 3.103% |  | 2.503% | 2.503%
13.75 | 3.103% |  | 2.503% | 2.503%
13.76 | 3.103% |  | 2.503% | 2.503%
13.77 | 3.103% |  | 2.503% | 2.503%
13.78 | 3.103% |  | 2.503% | 2.503%
13.79 | 3.103% |  | 2.503% | 2.503%
13.8 | 3.103% |  | 2.503% | 2.503%
13.81 | 3.103% |  | 2.503% | 2.503%
13.82 | 3.103% |  | 2.503% | 2.503%
13.83 | 3.103% |  | 2.503%
13.84 | 3.103% |  | 2.503%
13.85 | 3.103% |  | 2.503%
13.86 | 3.103% |  | 2.503%
13.87 | 3.103% |  | 2.503%
2.503%
2.503%
2.503%
2.503%

## 20 MeV PDDs
Energy: | 20 MeV
SSD: | 100 cm
PDD
Depth | 10x10+ | 7x7 | 6x6 | 4x4 | 3x3 | Energy | SSD | Field size crossline | R100 | R50 | R95 | R90 | R80 | Rp | Ds | PDD(2.37) | PDD(2.2)
0 |  | 93% | 93.2% | 93.3% |  | 20 MeV | 100.00 cm | 3x3 | 1.61 cm | 6.42 cm | 3.36 cm | 3.95 cm | 4.71 cm | 9.23 cm | 94.2% | 99% | 99.3%
0.01 |  | 93% | 93.3% | 93.3% |  | 20 MeV | 100.00 cm | 4x4 | 1.78 cm | 7.13 cm | 3.85 cm | 4.56 cm | 5.36 cm | 9.80 cm | 93.7% | 99.3% | 99.6%
0.02 |  | 93% | 93.4% | 93.4% |  | 20 MeV | 100.00 cm | 6x6 | 1.85 cm | 7.85 cm | 4.49 cm | 5.28 cm | 6.19 cm | 9.98 cm | 93.8% | 100% | 99.9%
0.03 |  | 93% | 93.4% | 93.4% |  | 20 MeV | 100.00 cm | 7x7 | 1.86 cm | 7.98 cm | 4.58 cm | 5.45 cm | 6.36 cm | 9.97 cm | 93% | 99.6% | 99.9%
0.04 | 92.19% | 93% | 93.5% | 93.5% | 94% | 20 MeV | 100.00 cm | 10x10+ | 2.37 cm | 8.20 cm | 5.10 cm | 5.88 cm | 6.75 cm | 9.99 cm | 92.3% | 99.9% | 99.8%
0.05 | 92.19% | 93% | 93.5% | 93.6% | 94.1%
0.06 | 92.29% | 93% | 93.6% | 93.6% | 94.2%
0.07 | 92.39% | 93% | 93.7% | 93.7% | 94.2%
0.08 | 92.39% | 93% | 93.7% | 93.7% | 94.3%
0.09 | 92.49% | 93% | 93.8% | 93.8% | 94.4%
0.1 | 92.49% | 93.1% | 93.8% | 93.9% | 94.4%
0.11 | 92.59% | 93.2% | 93.9% | 93.9% | 94.5%
0.12 | 92.59% | 93.3% | 94% | 94% | 94.6%
0.13 | 92.69% | 93.4% | 94% | 94% | 94.6%
0.14 | 92.69% | 93.6% | 94.1% | 94.1% | 94.7%
0.15 | 92.79% | 93.7% | 94.1% | 94.1% | 94.8%
0.16 | 92.79% | 93.8% | 94.2% | 94.2% | 94.9%
0.17 | 92.89% | 93.9% | 94.3% | 94.3% | 95%
0.18 | 92.99% | 94% | 94.3% | 94.3% | 95.1%
0.19 | 93.09% | 94.1% | 94.4% | 94.4% | 95.2%
0.2 | 93.19% | 94.3% | 94.5% | 94.4% | 95.3%
0.21 | 93.29% | 94.4% | 94.6% | 94.5% | 95.4%
0.22 | 93.39% | 94.5% | 94.7% | 94.6% | 95.5%
0.23 | 93.49% | 94.7% | 94.8% | 94.6% | 95.6%
0.24 | 93.59% | 94.8% | 94.9% | 94.7% | 95.7%
0.25 | 93.69% | 94.9% | 95.1% | 94.8% | 95.8%
0.26 | 93.79% | 95.1% | 95.2% | 94.9% | 95.9%
0.27 | 93.89% | 95.2% | 95.3% | 95% | 96%
0.28 | 93.99% | 95.3% | 95.4% | 95.1% | 96.1%
0.29 | 94.09% | 95.5% | 95.5% | 95.3% | 96.2%
0.3 | 94.19% | 95.6% | 95.7% | 95.4% | 96.2%
0.31 | 94.29% | 95.7% | 95.8% | 95.5% | 96.3%
0.32 | 94.39% | 95.9% | 95.9% | 95.6% | 96.4%
0.33 | 94.49% | 96% | 96% | 95.7% | 96.5%
0.34 | 94.59% | 96.1% | 96.1% | 95.8% | 96.6%
0.35 | 94.69% | 96.2% | 96.1% | 95.9% | 96.7%
0.36 | 94.79% | 96.2% | 96.2% | 96% | 96.8%
0.37 | 94.89% | 96.3% | 96.3% | 96.1% | 96.9%
0.38 | 94.99% | 96.4% | 96.4% | 96.2% | 97%
0.39 | 95.1% | 96.5% | 96.5% | 96.3% | 97.1%
0.4 | 95.2% | 96.6% | 96.6% | 96.4% | 97.1%
0.41 | 95.3% | 96.7% | 96.6% | 96.5% | 97.2%
0.42 | 95.4% | 96.8% | 96.7% | 96.6% | 97.3%
0.43 | 95.5% | 96.9% | 96.8% | 96.7% | 97.4%
0.44 | 95.5% | 96.9% | 96.9% | 96.8% | 97.5%
0.45 | 95.6% | 97% | 97% | 96.9% | 97.5%
0.46 | 95.7% | 97.1% | 97% | 97% | 97.6%
0.47 | 95.8% | 97.1% | 97.1% | 97% | 97.7%
0.48 | 95.9% | 97.2% | 97.1% | 97.1% | 97.7%
0.49 | 96% | 97.2% | 97.2% | 97.2% | 97.8%
0.5 | 96.1% | 97.3% | 97.2% | 97.3% | 97.8%
0.51 | 96.1% | 97.3% | 97.3% | 97.3% | 97.9%
0.52 | 96.2% | 97.4% | 97.4% | 97.4% | 97.9%
0.53 | 96.3% | 97.4% | 97.4% | 97.4% | 97.9%
0.54 | 96.3% | 97.4% | 97.5% | 97.5% | 98%
0.55 | 96.4% | 97.5% | 97.5% | 97.5% | 98%
0.56 | 96.5% | 97.5% | 97.6% | 97.6% | 98%
0.57 | 96.5% | 97.6% | 97.6% | 97.6% | 98.1%
0.58 | 96.6% | 97.6% | 97.7% | 97.7% | 98.1%
0.59 | 96.6% | 97.7% | 97.7% | 97.8% | 98.1%
0.6 | 96.7% | 97.7% | 97.7% | 97.8% | 98.2%
0.61 | 96.7% | 97.8% | 97.8% | 97.9% | 98.2%
0.62 | 96.7% | 97.8% | 97.8% | 97.9% | 98.3%
0.63 | 96.8% | 97.8% | 97.8% | 98% | 98.3%
0.64 | 96.8% | 97.9% | 97.9% | 98% | 98.3%
0.65 | 96.9% | 97.9% | 97.9% | 98.1% | 98.4%
0.66 | 96.9% | 98% | 98% | 98.1% | 98.4%
0.67 | 97% | 98% | 98% | 98.2% | 98.5%
0.68 | 97% | 98.1% | 98% | 98.2% | 98.5%
0.69 | 97.1% | 98.1% | 98% | 98.3% | 98.5%
0.7 | 97.1% | 98.2% | 98.1% | 98.3% | 98.6%
0.71 | 97.2% | 98.2% | 98.1% | 98.4% | 98.6%
0.72 | 97.2% | 98.3% | 98.1% | 98.4% | 98.6%
0.73 | 97.3% | 98.3% | 98.1% | 98.4% | 98.7%
0.74 | 97.3% | 98.4% | 98.1% | 98.5% | 98.7%
0.75 | 97.4% | 98.4% | 98.1% | 98.5% | 98.8%
0.76 | 97.4% | 98.5% | 98.1% | 98.5% | 98.8%
0.77 | 97.5% | 98.5% | 98.1% | 98.6% | 98.8%
0.78 | 97.5% | 98.6% | 98.1% | 98.6% | 98.9%
0.79 | 97.6% | 98.6% | 98.1% | 98.6% | 98.9%
0.8 | 97.6% | 98.7% | 98.1% | 98.7% | 98.9%
0.81 | 97.6% | 98.7% | 98.2% | 98.7% | 99%
0.82 | 97.7% | 98.7% | 98.3% | 98.7% | 99%
0.83 | 97.7% | 98.8% | 98.3% | 98.8% | 99%
0.84 | 97.7% | 98.8% | 98.4% | 98.8% | 99%
0.85 | 97.8% | 98.8% | 98.5% | 98.8% | 99.1%
0.86 | 97.8% | 98.8% | 98.5% | 98.9% | 99.1%
0.87 | 97.8% | 98.9% | 98.6% | 98.9% | 99.1%
0.88 | 97.9% | 98.9% | 98.7% | 98.9% | 99.1%
0.89 | 97.9% | 98.9% | 98.7% | 98.9% | 99.1%
0.9 | 98% | 98.9% | 98.8% | 99% | 99.1%
0.91 | 98% | 98.9% | 98.9% | 99% | 99.2%
0.92 | 98% | 99% | 98.9% | 99% | 99.2%
0.93 | 98.1% | 99% | 99% | 99.1% | 99.2%
0.94 | 98.1% | 99% | 99% | 99.1% | 99.2%
0.95 | 98.1% | 99% | 99% | 99.1% | 99.3%
0.96 | 98.2% | 99% | 99% | 99.2% | 99.3%
0.97 | 98.2% | 99.1% | 99.1% | 99.2% | 99.3%
0.98 | 98.2% | 99.1% | 99.1% | 99.2% | 99.3%
0.99 | 98.3% | 99.1% | 99.1% | 99.2% | 99.4%
1 | 98.3% | 99.1% | 99.1% | 99.3% | 99.4%
1.01 | 98.3% | 99.1% | 99.2% | 99.3% | 99.4%
1.02 | 98.4% | 99.2% | 99.2% | 99.3% | 99.4%
1.03 | 98.4% | 99.2% | 99.2% | 99.3% | 99.5%
1.04 | 98.4% | 99.2% | 99.2% | 99.3% | 99.5%
1.05 | 98.5% | 99.2% | 99.2% | 99.4% | 99.5%
1.06 | 98.5% | 99.2% | 99.2% | 99.4% | 99.6%
1.07 | 98.5% | 99.2% | 99.2% | 99.4% | 99.6%
1.08 | 98.5% | 99.2% | 99.2% | 99.4% | 99.6%
1.09 | 98.6% | 99.2% | 99.2% | 99.5% | 99.6%
1.1 | 98.6% | 99.2% | 99.2% | 99.5% | 99.7%
1.11 | 98.6% | 99.2% | 99.2% | 99.5% | 99.7%
1.12 | 98.6% | 99.2% | 99.2% | 99.5% | 99.7%
1.13 | 98.7% | 99.2% | 99.2% | 99.5% | 99.8%
1.14 | 98.7% | 99.3% | 99.2% | 99.5% | 99.8%
1.15 | 98.7% | 99.3% | 99.2% | 99.5% | 99.8%
1.16 | 98.7% | 99.3% | 99.2% | 99.6% | 99.8%
1.17 | 98.8% | 99.3% | 99.3% | 99.6% | 99.8%
1.18 | 98.8% | 99.3% | 99.3% | 99.6% | 99.9%
1.19 | 98.8% | 99.3% | 99.3% | 99.6% | 99.9%
1.2 | 98.8% | 99.3% | 99.4% | 99.6% | 99.9%
1.21 | 98.8% | 99.3% | 99.4% | 99.6% | 99.9%
1.22 | 98.8% | 99.4% | 99.4% | 99.7% | 99.9%
1.23 | 98.9% | 99.4% | 99.5% | 99.7% | 99.9%
1.24 | 98.9% | 99.4% | 99.5% | 99.7% | 99.9%
1.25 | 98.9% | 99.4% | 99.6% | 99.7% | 99.9%
1.26 | 98.9% | 99.4% | 99.6% | 99.7% | 99.9%
1.27 | 98.9% | 99.5% | 99.6% | 99.8% | 99.9%
1.28 | 99% | 99.5% | 99.7% | 99.8% | 99.9%
1.29 | 99% | 99.5% | 99.7% | 99.8% | 99.9%
1.3 | 99% | 99.5% | 99.7% | 99.8% | 99.9%
1.31 | 99% | 99.5% | 99.7% | 99.8% | 99.9%
1.32 | 99.1% | 99.5% | 99.6% | 99.9% | 99.9%
1.33 | 99.1% | 99.5% | 99.6% | 99.9% | 99.9%
1.34 | 99.1% | 99.5% | 99.6% | 99.9% | 99.9%
1.35 | 99.1% | 99.5% | 99.6% | 99.9% | 99.9%
1.36 | 99.1% | 99.5% | 99.6% | 99.9% | 99.9%
1.37 | 99.2% | 99.5% | 99.6% | 99.9% | 99.9%
1.38 | 99.2% | 99.6% | 99.6% | 99.9% | 99.9%
1.39 | 99.2% | 99.6% | 99.6% | 99.9% | 99.9%
1.4 | 99.2% | 99.6% | 99.5% | 99.9% | 99.9%
1.41 | 99.2% | 99.6% | 99.6% | 99.9% | 99.9%
1.42 | 99.2% | 99.6% | 99.6% | 99.9% | 99.9%
1.43 | 99.3% | 99.6% | 99.6% | 99.9% | 99.9%
1.44 | 99.3% | 99.6% | 99.7% | 99.9% | 99.9%
1.45 | 99.3% | 99.6% | 99.7% | 99.9% | 99.9%
1.46 | 99.3% | 99.6% | 99.7% | 99.9% | 99.9%
1.47 | 99.3% | 99.6% | 99.8% | 99.9% | 99.9%
1.48 | 99.3% | 99.6% | 99.8% | 99.9% | 99.9%
1.49 | 99.3% | 99.6% | 99.8% | 99.9% | 99.9%
1.5 | 99.3% | 99.7% | 99.9% | 99.9% | 99.9%
1.51 | 99.4% | 99.7% | 99.9% | 99.9% | 99.9%
1.52 | 99.4% | 99.7% | 99.9% | 99.9% | 99.9%
1.53 | 99.4% | 99.7% | 99.9% | 100% | 100%
1.54 | 99.4% | 99.7% | 99.9% | 100% | 100%
1.55 | 99.4% | 99.7% | 99.9% | 100% | 100%
1.56 | 99.4% | 99.7% | 99.9% | 100% | 100%
1.57 | 99.5% | 99.7% | 99.8% | 100% | 100%
1.58 | 99.5% | 99.7% | 99.8% | 100% | 100%
1.59 | 99.5% | 99.7% | 99.8% | 99.9% | 100%
1.6 | 99.5% | 99.7% | 99.8% | 99.9% | 100%
1.61 | 99.5% | 99.7% | 99.8% | 99.9% | 100%
1.62 | 99.5% | 99.7% | 99.8% | 99.9% | 100%
1.63 | 99.5% | 99.7% | 99.7% | 99.9% | 100%
1.64 | 99.5% | 99.7% | 99.7% | 99.9% | 100%
1.65 | 99.5% | 99.7% | 99.7% | 99.9% | 100%
1.66 | 99.5% | 99.8% | 99.7% | 99.9% | 100%
1.67 | 99.5% | 99.8% | 99.7% | 99.8% | 100%
1.68 | 99.6% | 99.8% | 99.7% | 99.8% | 100%
1.69 | 99.6% | 99.8% | 99.8% | 99.8% | 99.9%
1.7 | 99.6% | 99.8% | 99.8% | 99.9% | 99.9%
1.71 | 99.6% | 99.8% | 99.8% | 99.9% | 99.9%
1.72 | 99.6% | 99.8% | 99.8% | 99.9% | 99.9%
1.73 | 99.6% | 99.8% | 99.8% | 99.9% | 99.9%
1.74 | 99.6% | 99.8% | 99.8% | 99.9% | 99.9%
1.75 | 99.6% | 99.8% | 99.8% | 99.9% | 99.9%
1.76 | 99.6% | 99.8% | 99.8% | 99.9% | 99.8%
1.77 | 99.6% | 99.9% | 99.8% | 100% | 99.8%
1.78 | 99.7% | 99.9% | 99.8% | 100% | 99.8%
1.79 | 99.7% | 99.9% | 99.8% | 100% | 99.8%
1.8 | 99.7% | 99.9% | 99.9% | 100% | 99.8%
1.81 | 99.7% | 99.9% | 99.9% | 100% | 99.8%
1.82 | 99.7% | 99.9% | 99.9% | 100% | 99.8%
1.83 | 99.7% | 99.9% | 99.9% | 99.9% | 99.8%
1.84 | 99.7% | 99.9% | 99.9% | 99.9% | 99.7%
1.85 | 99.7% | 100% | 99.9% | 99.9% | 99.7%
1.86 | 99.7% | 100% | 100% | 99.9% | 99.7%
1.87 | 99.7% | 100% | 100% | 99.8% | 99.7%
1.88 | 99.7% | 100% | 100% | 99.8% | 99.7%
1.89 | 99.7% | 100% | 100% | 99.8% | 99.7%
1.9 | 99.7% | 100% | 100% | 99.8% | 99.7%
1.91 | 99.7% | 99.9% | 100% | 99.8% | 99.7%
1.92 | 99.7% | 99.9% | 99.9% | 99.7% | 99.7%
1.93 | 99.7% | 99.9% | 99.9% | 99.7% | 99.8%
1.94 | 99.7% | 99.9% | 99.9% | 99.7% | 99.8%
1.95 | 99.8% | 99.9% | 99.9% | 99.7% | 99.8%
1.96 | 99.8% | 99.8% | 99.8% | 99.8% | 99.8%
1.97 | 99.8% | 99.8% | 99.8% | 99.8% | 99.8%
1.98 | 99.8% | 99.8% | 99.8% | 99.8% | 99.8%
1.99 | 99.8% | 99.8% | 99.7% | 99.8% | 99.8%
2 | 99.8% | 99.7% | 99.7% | 99.8% | 99.8%
2.01 | 99.8% | 99.7% | 99.7% | 99.8% | 99.8%
2.02 | 99.8% | 99.7% | 99.7% | 99.8% | 99.7%
2.03 | 99.8% | 99.7% | 99.7% | 99.8% | 99.7%
2.04 | 99.8% | 99.7% | 99.7% | 99.8% | 99.7%
2.05 | 99.8% | 99.7% | 99.7% | 99.8% | 99.7%
2.06 | 99.8% | 99.7% | 99.7% | 99.8% | 99.7%
2.07 | 99.8% | 99.7% | 99.8% | 99.8% | 99.6%
2.08 | 99.8% | 99.7% | 99.8% | 99.8% | 99.6%
2.09 | 99.8% | 99.7% | 99.8% | 99.8% | 99.6%
2.1 | 99.8% | 99.7% | 99.8% | 99.8% | 99.6%
2.11 | 99.8% | 99.7% | 99.8% | 99.8% | 99.5%
2.12 | 99.8% | 99.7% | 99.8% | 99.7% | 99.5%
2.13 | 99.8% | 99.8% | 99.8% | 99.7% | 99.5%
2.14 | 99.8% | 99.8% | 99.8% | 99.7% | 99.5%
2.15 | 99.8% | 99.8% | 99.8% | 99.7% | 99.4%
2.16 | 99.8% | 99.8% | 99.8% | 99.7% | 99.4%
2.17 | 99.8% | 99.8% | 99.9% | 99.7% | 99.4%
2.18 | 99.8% | 99.8% | 99.9% | 99.7% | 99.4%
2.19 | 99.8% | 99.8% | 99.9% | 99.7% | 99.4%
2.2 | 99.8% | 99.9% | 99.9% | 99.6% | 99.3%
2.21 | 99.8% | 99.9% | 99.9% | 99.6% | 99.3%
2.22 | 99.8% | 99.9% | 99.9% | 99.6% | 99.3%
2.23 | 99.8% | 99.9% | 99.9% | 99.6% | 99.3%
2.24 | 99.8% | 99.9% | 99.9% | 99.6% | 99.3%
2.25 | 99.8% | 99.9% | 99.9% | 99.6% | 99.3%
2.26 | 99.9% | 99.9% | 99.9% | 99.5% | 99.2%
2.27 | 99.9% | 99.8% | 99.9% | 99.5% | 99.2%
2.28 | 99.9% | 99.8% | 99.9% | 99.5% | 99.2%
2.29 | 99.9% | 99.8% | 99.9% | 99.5% | 99.2%
2.3 | 99.9% | 99.8% | 100% | 99.4% | 99.2%
2.31 | 99.9% | 99.7% | 100% | 99.4% | 99.2%
2.32 | 99.9% | 99.7% | 100% | 99.4% | 99.1%
2.33 | 99.9% | 99.7% | 100% | 99.4% | 99.1%
2.34 | 99.9% | 99.7% | 100% | 99.3% | 99.1%
2.35 | 99.9% | 99.6% | 100% | 99.3% | 99.1%
2.36 | 99.9% | 99.6% | 100% | 99.3% | 99%
2.37 | 99.9% | 99.6% | 100% | 99.3% | 99%
2.38 | 100% | 99.6% | 99.9% | 99.2% | 98.9%
2.39 | 100% | 99.6% | 99.9% | 99.2% | 98.9%
2.4 | 100% | 99.6% | 99.8% | 99.2% | 98.8%
2.41 | 99.9% | 99.6% | 99.8% | 99.2% | 98.8%
2.42 | 99.9% | 99.6% | 99.7% | 99.2% | 98.8%
2.43 | 99.9% | 99.6% | 99.7% | 99.3% | 98.8%
2.44 | 99.9% | 99.6% | 99.6% | 99.3% | 98.7%
2.45 | 99.9% | 99.6% | 99.6% | 99.3% | 98.7%
2.46 | 99.9% | 99.6% | 99.6% | 99.3% | 98.7%
2.47 | 99.9% | 99.6% | 99.5% | 99.4% | 98.7%
2.48 | 99.9% | 99.6% | 99.5% | 99.4% | 98.7%
2.49 | 99.9% | 99.6% | 99.4% | 99.4% | 98.7%
2.5 | 99.9% | 99.6% | 99.5% | 99.4% | 98.7%
2.51 | 99.9% | 99.6% | 99.5% | 99.5% | 98.7%
2.52 | 99.9% | 99.6% | 99.5% | 99.5% | 98.7%
2.53 | 99.9% | 99.5% | 99.6% | 99.5% | 98.7%
2.54 | 99.9% | 99.5% | 99.6% | 99.5% | 98.7%
2.55 | 99.9% | 99.5% | 99.6% | 99.5% | 98.6%
2.56 | 99.9% | 99.5% | 99.6% | 99.5% | 98.6%
2.57 | 99.9% | 99.5% | 99.7% | 99.5% | 98.6%
2.58 | 99.9% | 99.5% | 99.7% | 99.5% | 98.6%
2.59 | 99.9% | 99.5% | 99.7% | 99.5% | 98.6%
2.6 | 99.9% | 99.5% | 99.8% | 99.5% | 98.5%
2.61 | 99.9% | 99.5% | 99.8% | 99.5% | 98.5%
2.62 | 99.9% | 99.5% | 99.7% | 99.5% | 98.4%
2.63 | 99.9% | 99.5% | 99.7% | 99.5% | 98.4%
2.64 | 99.9% | 99.5% | 99.7% | 99.5% | 98.4%
2.65 | 99.9% | 99.6% | 99.7% | 99.5% | 98.3%
2.66 | 99.9% | 99.6% | 99.6% | 99.4% | 98.3%
2.67 | 99.9% | 99.6% | 99.6% | 99.4% | 98.2%
2.68 | 100% | 99.6% | 99.6% | 99.3% | 98.2%
2.69 | 100% | 99.7% | 99.5% | 99.3% | 98.1%
2.7 | 99.9% | 99.7% | 99.5% | 99.2% | 98.1%
2.71 | 99.9% | 99.7% | 99.5% | 99.2% | 98%
2.72 | 99.9% | 99.8% | 99.4% | 99.1% | 98%
2.73 | 99.9% | 99.8% | 99.4% | 99.1% | 98%
2.74 | 99.9% | 99.8% | 99.4% | 99% | 97.9%
2.75 | 99.9% | 99.7% | 99.4% | 99% | 97.9%
2.76 | 99.9% | 99.7% | 99.4% | 98.9% | 97.9%
2.77 | 99.9% | 99.7% | 99.4% | 98.9% | 97.9%
2.78 | 99.9% | 99.6% | 99.4% | 98.9% | 97.8%
2.79 | 99.9% | 99.6% | 99.4% | 98.9% | 97.8%
2.8 | 99.9% | 99.6% | 99.4% | 98.8% | 97.8%
2.81 | 99.9% | 99.6% | 99.4% | 98.8% | 97.7%
2.82 | 99.9% | 99.5% | 99.4% | 98.8% | 97.7%
2.83 | 99.9% | 99.5% | 99.4% | 98.8% | 97.7%
2.84 | 99.9% | 99.5% | 99.4% | 98.8% | 97.6%
2.85 | 99.9% | 99.4% | 99.4% | 98.8% | 97.6%
2.86 | 99.9% | 99.4% | 99.3% | 98.7% | 97.5%
2.87 | 99.9% | 99.4% | 99.3% | 98.7% | 97.5%
2.88 | 99.9% | 99.4% | 99.3% | 98.7% | 97.4%
2.89 | 99.9% | 99.4% | 99.3% | 98.7% | 97.4%
2.9 | 99.9% | 99.4% | 99.3% | 98.6% | 97.3%
2.91 | 99.9% | 99.4% | 99.3% | 98.6% | 97.3%
2.92 | 99.9% | 99.4% | 99.3% | 98.6% | 97.2%
2.93 | 99.9% | 99.4% | 99.3% | 98.6% | 97.2%
2.94 | 99.9% | 99.4% | 99.3% | 98.5% | 97.1%
2.95 | 99.9% | 99.4% | 99.3% | 98.5% | 97.1%
2.96 | 99.9% | 99.3% | 99.3% | 98.5% | 97%
2.97 | 99.8% | 99.3% | 99.3% | 98.4% | 97%
2.98 | 99.8% | 99.3% | 99.2% | 98.4% | 96.9%
2.99 | 99.8% | 99.3% | 99.2% | 98.4% | 96.9%
3 | 99.8% | 99.3% | 99.2% | 98.4% | 96.9%
3.01 | 99.8% | 99.3% | 99.2% | 98.3% | 96.8%
3.02 | 99.8% | 99.3% | 99.1% | 98.3% | 96.8%
3.03 | 99.8% | 99.2% | 99.1% | 98.2% | 96.8%
3.04 | 99.7% | 99.2% | 99.1% | 98.2% | 96.7%
3.05 | 99.7% | 99.2% | 99.1% | 98.2% | 96.7%
3.06 | 99.7% | 99.2% | 99% | 98.1% | 96.7%
3.07 | 99.7% | 99.2% | 99% | 98.1% | 96.6%
3.08 | 99.7% | 99.2% | 99% | 98.1% | 96.6%
3.09 | 99.7% | 99.1% | 99% | 98% | 96.6%
3.1 | 99.7% | 99.1% | 99% | 98% | 96.5%
3.11 | 99.7% | 99% | 99% | 98% | 96.5%
3.12 | 99.6% | 99% | 99% | 97.9% | 96.4%
3.13 | 99.6% | 98.9% | 99% | 97.9% | 96.4%
3.14 | 99.6% | 98.9% | 99% | 97.9% | 96.3%
3.15 | 99.6% | 98.8% | 99% | 97.8% | 96.3%
3.16 | 99.6% | 98.8% | 99% | 97.8% | 96.2%
3.17 | 99.6% | 98.7% | 98.9% | 97.8% | 96.2%
3.18 | 99.6% | 98.7% | 98.9% | 97.8% | 96.1%
3.19 | 99.5% | 98.6% | 98.9% | 97.7% | 96.1%
3.2 | 99.5% | 98.6% | 98.9% | 97.7% | 96%
3.21 | 99.5% | 98.5% | 98.9% | 97.7% | 96%
3.22 | 99.5% | 98.5% | 98.9% | 97.6% | 95.9%
3.23 | 99.5% | 98.5% | 98.9% | 97.6% | 95.8%
3.24 | 99.5% | 98.6% | 98.9% | 97.6% | 95.8%
3.25 | 99.5% | 98.6% | 98.9% | 97.6% | 95.7%
3.26 | 99.5% | 98.6% | 98.9% | 97.6% | 95.7%
3.27 | 99.5% | 98.6% | 98.9% | 97.6% | 95.6%
3.28 | 99.5% | 98.6% | 98.9% | 97.6% | 95.6%
3.29 | 99.4% | 98.7% | 98.9% | 97.5% | 95.5%
3.3 | 99.4% | 98.7% | 98.9% | 97.5% | 95.4%
3.31 | 99.4% | 98.7% | 98.9% | 97.5% | 95.4%
3.32 | 99.4% | 98.7% | 98.9% | 97.5% | 95.3%
3.33 | 99.4% | 98.7% | 98.9% | 97.5% | 95.2%
3.34 | 99.4% | 98.7% | 98.9% | 97.5% | 95.2%
3.35 | 99.4% | 98.7% | 98.8% | 97.5% | 95.1%
3.36 | 99.4% | 98.7% | 98.8% | 97.5% | 95%
3.37 | 99.4% | 98.7% | 98.8% | 97.4% | 94.9%
3.38 | 99.4% | 98.6% | 98.7% | 97.4% | 94.9%
3.39 | 99.4% | 98.6% | 98.7% | 97.3% | 94.8%
3.4 | 99.4% | 98.6% | 98.6% | 97.3% | 94.8%
3.41 | 99.4% | 98.6% | 98.6% | 97.2% | 94.7%
3.42 | 99.4% | 98.5% | 98.5% | 97.1% | 94.6%
3.43 | 99.4% | 98.5% | 98.5% | 97.1% | 94.6%
3.44 | 99.4% | 98.5% | 98.5% | 97% | 94.5%
3.45 | 99.3% | 98.5% | 98.5% | 97% | 94.5%
3.46 | 99.3% | 98.5% | 98.5% | 96.9% | 94.4%
3.47 | 99.3% | 98.4% | 98.5% | 96.9% | 94.3%
3.48 | 99.3% | 98.4% | 98.5% | 96.8% | 94.3%
3.49 | 99.3% | 98.4% | 98.5% | 96.8% | 94.2%
3.5 | 99.3% | 98.4% | 98.4% | 96.7% | 94.1%
3.51 | 99.3% | 98.3% | 98.4% | 96.7% | 94%
3.52 | 99.3% | 98.3% | 98.4% | 96.7% | 94%
3.53 | 99.3% | 98.3% | 98.4% | 96.7% | 93.9%
3.54 | 99.3% | 98.3% | 98.4% | 96.6% | 93.8%
3.55 | 99.2% | 98.2% | 98.4% | 96.6% | 93.7%
3.56 | 99.2% | 98.2% | 98.4% | 96.6% | 93.6%
3.57 | 99.2% | 98.2% | 98.4% | 96.5% | 93.5%
3.58 | 99.2% | 98.2% | 98.3% | 96.5% | 93.5%
3.59 | 99.2% | 98.2% | 98.3% | 96.5% | 93.4%
3.6 | 99.2% | 98.2% | 98.2% | 96.4% | 93.3%
3.61 | 99.1% | 98.2% | 98.2% | 96.4% | 93.2%
3.62 | 99.1% | 98.2% | 98.1% | 96.3% | 93.2%
3.63 | 99.1% | 98.2% | 98.1% | 96.2% | 93.1%
3.64 | 99.1% | 98.1% | 98.1% | 96.2% | 93%
3.65 | 99% | 98.1% | 98% | 96.1% | 93%
3.66 | 99% | 98.1% | 98% | 96% | 92.9%
3.67 | 99% | 98.1% | 97.9% | 96% | 92.8%
3.68 | 99% | 98.1% | 97.9% | 95.9% | 92.7%
3.69 | 99% | 98.1% | 97.8% | 95.8% | 92.7%
3.7 | 98.9% | 98.1% | 97.8% | 95.8% | 92.6%
3.71 | 98.9% | 98% | 97.8% | 95.7% | 92.5%
3.72 | 98.9% | 98% | 97.8% | 95.7% | 92.4%
3.73 | 98.9% | 97.9% | 97.8% | 95.6% | 92.3%
3.74 | 98.9% | 97.9% | 97.8% | 95.6% | 92.2%
3.75 | 98.9% | 97.9% | 97.8% | 95.5% | 92.1%
3.76 | 98.9% | 97.8% | 97.8% | 95.5% | 91.9%
3.77 | 98.9% | 97.8% | 97.8% | 95.4% | 91.8%
3.78 | 98.8% | 97.7% | 97.8% | 95.4% | 91.7%
3.79 | 98.8% | 97.7% | 97.8% | 95.4% | 91.6%
3.8 | 98.8% | 97.7% | 97.8% | 95.3% | 91.5%
3.81 | 98.8% | 97.6% | 97.8% | 95.3% | 91.4%
3.82 | 98.8% | 97.6% | 97.7% | 95.2% | 91.3%
3.83 | 98.8% | 97.6% | 97.7% | 95.2% | 91.2%
3.84 | 98.8% | 97.6% | 97.7% | 95.2% | 91.1%
3.85 | 98.7% | 97.6% | 97.7% | 95.1% | 91%
3.86 | 98.7% | 97.6% | 97.6% | 95% | 90.9%
3.87 | 98.7% | 97.6% | 97.6% | 95% | 90.8%
3.88 | 98.7% | 97.6% | 97.6% | 94.9% | 90.7%
3.89 | 98.7% | 97.6% | 97.6% | 94.9% | 90.6%
3.9 | 98.6% | 97.5% | 97.5% | 94.8% | 90.5%
3.91 | 98.6% | 97.5% | 97.5% | 94.7% | 90.5%
3.92 | 98.6% | 97.5% | 97.5% | 94.7% | 90.4%
3.93 | 98.6% | 97.5% | 97.4% | 94.6% | 90.3%
3.94 | 98.5% | 97.5% | 97.4% | 94.5% | 90.2%
3.95 | 98.5% | 97.4% | 97.3% | 94.5% | 90%
3.96 | 98.5% | 97.4% | 97.2% | 94.4% | 89.9%
3.97 | 98.5% | 97.4% | 97.2% | 94.4% | 89.8%
3.98 | 98.4% | 97.3% | 97.1% | 94.3% | 89.7%
3.99 | 98.4% | 97.3% | 97% | 94.2% | 89.6%
4 | 98.4% | 97.3% | 97% | 94.2% | 89.5%
4.01 | 98.4% | 97.2% | 96.9% | 94.1% | 89.3%
4.02 | 98.3% | 97.2% | 96.8% | 94% | 89.2%
4.03 | 98.3% | 97.2% | 96.8% | 94% | 89.1%
4.04 | 98.3% | 97.1% | 96.7% | 93.9% | 89%
4.05 | 98.3% | 97.1% | 96.7% | 93.9% | 88.9%
4.06 | 98.3% | 97.1% | 96.6% | 93.8% | 88.8%
4.07 | 98.3% | 97% | 96.6% | 93.7% | 88.7%
4.08 | 98.2% | 97% | 96.6% | 93.7% | 88.6%
4.09 | 98.2% | 97% | 96.6% | 93.6% | 88.5%
4.1 | 98.2% | 96.9% | 96.6% | 93.5% | 88.3%
4.11 | 98.2% | 96.9% | 96.6% | 93.5% | 88.2%
4.12 | 98.2% | 96.9% | 96.6% | 93.4% | 88.1%
4.13 | 98.1% | 96.9% | 96.6% | 93.3% | 88%
4.14 | 98.1% | 96.8% | 96.6% | 93.2% | 87.9%
4.15 | 98.1% | 96.8% | 96.6% | 93.2% | 87.8%
4.16 | 98.1% | 96.8% | 96.6% | 93.1% | 87.7%
4.17 | 98.1% | 96.7% | 96.6% | 93% | 87.5%
4.18 | 98% | 96.7% | 96.5% | 92.9% | 87.4%
4.19 | 98% | 96.7% | 96.4% | 92.9% | 87.3%
4.2 | 98% | 96.6% | 96.4% | 92.8% | 87.1%
4.21 | 98% | 96.6% | 96.3% | 92.7% | 87%
4.22 | 97.9% | 96.5% | 96.2% | 92.7% | 86.9%
4.23 | 97.9% | 96.5% | 96.2% | 92.6% | 86.7%
4.24 | 97.9% | 96.4% | 96.1% | 92.5% | 86.6%
4.25 | 97.9% | 96.4% | 96% | 92.5% | 86.5%
4.26 | 97.8% | 96.4% | 96% | 92.4% | 86.4%
4.27 | 97.8% | 96.3% | 95.9% | 92.3% | 86.3%
4.28 | 97.8% | 96.3% | 95.8% | 92.3% | 86.1%
4.29 | 97.8% | 96.2% | 95.8% | 92.2% | 86%
4.3 | 97.7% | 96.2% | 95.8% | 92.2% | 85.9%
4.31 | 97.7% | 96.2% | 95.7% | 92.1% | 85.7%
4.32 | 97.7% | 96.1% | 95.7% | 92% | 85.6%
4.33 | 97.6% | 96.1% | 95.7% | 91.9% | 85.5%
4.34 | 97.6% | 96% | 95.6% | 91.8% | 85.3%
4.35 | 97.6% | 96% | 95.6% | 91.8% | 85.2%
4.36 | 97.5% | 96% | 95.6% | 91.7% | 85.1%
4.37 | 97.5% | 95.9% | 95.5% | 91.6% | 84.9%
4.38 | 97.5% | 95.9% | 95.5% | 91.5% | 84.8%
4.39 | 97.5% | 95.9% | 95.5% | 91.4% | 84.6%
4.4 | 97.5% | 95.8% | 95.4% | 91.3% | 84.5%
4.41 | 97.5% | 95.8% | 95.4% | 91.2% | 84.4%
4.42 | 97.4% | 95.7% | 95.4% | 91.1% | 84.2%
4.43 | 97.4% | 95.7% | 95.3% | 91% | 84.1%
4.44 | 97.4% | 95.7% | 95.3% | 90.9% | 84%
4.45 | 97.4% | 95.6% | 95.3% | 90.8% | 83.8%
4.46 | 97.4% | 95.6% | 95.2% | 90.8% | 83.7%
4.47 | 97.4% | 95.5% | 95.2% | 90.7% | 83.6%
4.48 | 97.4% | 95.5% | 95.1% | 90.6% | 83.4%
4.49 | 97.3% | 95.5% | 95.1% | 90.6% | 83.3%
4.5 | 97.3% | 95.4% | 95.1% | 90.5% | 83.2%
4.51 | 97.3% | 95.4% | 95% | 90.4% | 83.1%
4.52 | 97.3% | 95.3% | 95% | 90.4% | 83%
4.53 | 97.2% | 95.3% | 94.9% | 90.3% | 82.8%
4.54 | 97.2% | 95.2% | 94.9% | 90.2% | 82.7%
4.55 | 97.2% | 95.2% | 94.8% | 90.2% | 82.5%
4.56 | 97.2% | 95.2% | 94.8% | 90.1% | 82.4%
4.57 | 97.1% | 95.1% | 94.7% | 90% | 82.3%
4.58 | 97.1% | 95.1% | 94.7% | 89.9% | 82.1%
4.59 | 97.1% | 95% | 94.7% | 89.8% | 82%
4.6 | 97% | 95% | 94.6% | 89.6% | 81.8%
4.61 | 97% | 94.9% | 94.6% | 89.5% | 81.7%
4.62 | 97% | 94.9% | 94.5% | 89.4% | 81.5%
4.63 | 96.9% | 94.8% | 94.5% | 89.3% | 81.3%
4.64 | 96.9% | 94.8% | 94.4% | 89.1% | 81.2%
4.65 | 96.8% | 94.7% | 94.4% | 89% | 81%
4.66 | 96.8% | 94.7% | 94.3% | 88.9% | 80.8%
4.67 | 96.7% | 94.7% | 94.2% | 88.8% | 80.7%
4.68 | 96.7% | 94.6% | 94.2% | 88.6% | 80.5%
4.69 | 96.7% | 94.6% | 94.1% | 88.5% | 80.3%
4.7 | 96.6% | 94.6% | 94.1% | 88.4% | 80.2%
4.71 | 96.6% | 94.5% | 94% | 88.3% | 80%
4.72 | 96.5% | 94.5% | 93.9% | 88.2% | 79.8%
4.73 | 96.5% | 94.5% | 93.9% | 88.1% | 79.7%
4.74 | 96.5% | 94.4% | 93.8% | 88% | 79.5%
4.75 | 96.5% | 94.4% | 93.8% | 87.8% | 79.3%
4.76 | 96.4% | 94.4% | 93.7% | 87.7% | 79.2%
4.77 | 96.4% | 94.3% | 93.7% | 87.6% | 79%
4.78 | 96.4% | 94.3% | 93.6% | 87.5% | 78.9%
4.79 | 96.3% | 94.2% | 93.5% | 87.4% | 78.7%
4.8 | 96.3% | 94.2% | 93.5% | 87.3% | 78.5%
4.81 | 96.2% | 94.1% | 93.4% | 87.2% | 78.4%
4.82 | 96.2% | 94.1% | 93.4% | 87% | 78.2%
4.83 | 96.2% | 94% | 93.3% | 86.9% | 78.1%
4.84 | 96.1% | 94% | 93.3% | 86.7% | 77.9%
4.85 | 96.1% | 93.9% | 93.2% | 86.6% | 77.8%
4.86 | 96% | 93.9% | 93.2% | 86.5% | 77.6%
4.87 | 96% | 93.8% | 93.1% | 86.3% | 77.5%
4.88 | 95.9% | 93.8% | 93.1% | 86.2% | 77.3%
4.89 | 95.9% | 93.7% | 93% | 86.1% | 77.1%
4.9 | 95.9% | 93.7% | 92.9% | 85.9% | 77%
4.91 | 95.8% | 93.6% | 92.8% | 85.8% | 76.8%
4.92 | 95.8% | 93.5% | 92.8% | 85.6% | 76.7%
4.93 | 95.8% | 93.4% | 92.7% | 85.5% | 76.5%
4.94 | 95.7% | 93.4% | 92.6% | 85.4% | 76.3%
4.95 | 95.7% | 93.3% | 92.5% | 85.3% | 76.2%
4.96 | 95.6% | 93.2% | 92.4% | 85.2% | 76%
4.97 | 95.6% | 93.2% | 92.4% | 85.1% | 75.8%
4.98 | 95.6% | 93.1% | 92.3% | 85% | 75.6%
4.99 | 95.5% | 93% | 92.2% | 84.9% | 75.4%
5 | 95.5% | 92.9% | 92.1% | 84.8% | 75.3%
5.01 | 95.4% | 92.9% | 92% | 84.7% | 75.1%
5.02 | 95.4% | 92.8% | 92% | 84.6% | 74.9%
5.03 | 95.3% | 92.8% | 91.9% | 84.6% | 74.8%
5.04 | 95.3% | 92.7% | 91.8% | 84.5% | 74.6%
5.05 | 95.2% | 92.6% | 91.8% | 84.4% | 74.4%
5.06 | 95.2% | 92.6% | 91.7% | 84.2% | 74.3%
5.07 | 95.1% | 92.5% | 91.6% | 84.1% | 74.1%
5.08 | 95.1% | 92.5% | 91.6% | 84% | 73.9%
5.09 | 94.99% | 92.4% | 91.5% | 83.8% | 73.8%
5.1 | 94.99% | 92.4% | 91.5% | 83.7% | 73.6%
5.11 | 94.89% | 92.3% | 91.4% | 83.6% | 73.4%
5.12 | 94.89% | 92.3% | 91.3% | 83.5% | 73.3%
5.13 | 94.89% | 92.2% | 91.3% | 83.3% | 73.1%
5.14 | 94.79% | 92.1% | 91.2% | 83.2% | 72.9%
5.15 | 94.79% | 92.1% | 91.1% | 83.1% | 72.7%
5.16 | 94.69% | 92% | 91% | 82.9% | 72.5%
5.17 | 94.69% | 91.9% | 91% | 82.8% | 72.3%
5.18 | 94.59% | 91.8% | 90.9% | 82.6% | 72.1%
5.19 | 94.59% | 91.8% | 90.8% | 82.5% | 72%
5.2 | 94.49% | 91.7% | 90.7% | 82.4% | 71.8%
5.21 | 94.49% | 91.6% | 90.7% | 82.2% | 71.6%
5.22 | 94.39% | 91.5% | 90.6% | 82.1% | 71.4%
5.23 | 94.39% | 91.5% | 90.5% | 81.9% | 71.3%
5.24 | 94.29% | 91.4% | 90.4% | 81.8% | 71.1%
5.25 | 94.29% | 91.3% | 90.4% | 81.6% | 71%
5.26 | 94.19% | 91.3% | 90.3% | 81.5% | 70.8%
5.27 | 94.19% | 91.2% | 90.2% | 81.3% | 70.6%
5.28 | 94.09% | 91.2% | 90.1% | 81.2% | 70.5%
5.29 | 93.99% | 91.2% | 90% | 81% | 70.3%
5.3 | 93.99% | 91.1% | 89.9% | 80.9% | 70.1%
5.31 | 93.89% | 91.1% | 89.9% | 80.8% | 70%
5.32 | 93.89% | 91.1% | 89.8% | 80.6% | 69.8%
5.33 | 93.79% | 91% | 89.7% | 80.5% | 69.6%
5.34 | 93.79% | 91% | 89.6% | 80.4% | 69.4%
5.35 | 93.69% | 91% | 89.5% | 80.2% | 69.2%
5.36 | 93.59% | 90.9% | 89.5% | 80.1% | 69%
5.37 | 93.59% | 90.9% | 89.4% | 80% | 68.8%
5.38 | 93.49% | 90.8% | 89.3% | 79.9% | 68.6%
5.39 | 93.39% | 90.7% | 89.2% | 79.7% | 68.4%
5.4 | 93.39% | 90.6% | 89.1% | 79.6% | 68.2%
5.41 | 93.29% | 90.5% | 89% | 79.4% | 68%
5.42 | 93.29% | 90.4% | 88.9% | 79.3% | 67.8%
5.43 | 93.19% | 90.3% | 88.8% | 79.1% | 67.6%
5.44 | 93.09% | 90.1% | 88.7% | 78.9% | 67.4%
5.45 | 93.09% | 90% | 88.6% | 78.8% | 67.2%
5.46 | 92.99% | 89.9% | 88.6% | 78.6% | 67%
5.47 | 92.99% | 89.8% | 88.5% | 78.4% | 66.8%
5.48 | 92.89% | 89.7% | 88.4% | 78.3% | 66.6%
5.49 | 92.89% | 89.6% | 88.3% | 78.1% | 66.4%
5.5 | 92.79% | 89.5% | 88.2% | 78% | 66.2%
5.51 | 92.79% | 89.4% | 88.1% | 77.8% | 66%
5.52 | 92.79% | 89.3% | 88% | 77.6% | 65.8%
5.53 | 92.69% | 89.3% | 87.8% | 77.5% | 65.6%
5.54 | 92.69% | 89.2% | 87.7% | 77.3% | 65.5%
5.55 | 92.59% | 89.1% | 87.6% | 77.2% | 65.3%
5.56 | 92.59% | 89% | 87.5% | 77% | 65.1%
5.57 | 92.49% | 88.9% | 87.4% | 76.9% | 64.9%
5.58 | 92.39% | 88.9% | 87.3% | 76.7% | 64.8%
5.59 | 92.39% | 88.8% | 87.2% | 76.5% | 64.6%
5.6 | 92.29% | 88.7% | 87% | 76.4% | 64.4%
5.61 | 92.19% | 88.6% | 86.9% | 76.2% | 64.3%
5.62 | 92.09% | 88.5% | 86.8% | 76.1% | 64.1%
5.63 | 92.09% | 88.4% | 86.7% | 75.9% | 64%
5.64 | 91.99% | 88.3% | 86.6% | 75.8% | 63.8%
5.65 | 91.89% | 88.2% | 86.5% | 75.6% | 63.7%
5.66 | 91.79% | 88.1% | 86.4% | 75.5% | 63.5%
5.67 | 91.79% | 88% | 86.3% | 75.3% | 63.3%
5.68 | 91.69% | 87.9% | 86.2% | 75.2% | 63.2%
5.69 | 91.59% | 87.9% | 86.1% | 75% | 63%
5.7 | 91.49% | 87.8% | 86.1% | 74.8% | 62.8%
5.71 | 91.39% | 87.7% | 86% | 74.7% | 62.6%
5.72 | 91.39% | 87.6% | 85.9% | 74.5% | 62.4%
5.73 | 91.29% | 87.5% | 85.8% | 74.4% | 62.2%
5.74 | 91.19% | 87.4% | 85.7% | 74.2% | 62.1%
5.75 | 91.09% | 87.3% | 85.5% | 74.1% | 61.9%
5.76 | 90.99% | 87.2% | 85.4% | 73.9% | 61.7%
5.77 | 90.89% | 87.1% | 85.3% | 73.7% | 61.5%
5.78 | 90.79% | 86.9% | 85.2% | 73.6% | 61.3%
5.79 | 90.69% | 86.8% | 85.1% | 73.4% | 61.1%
5.8 | 90.69% | 86.7% | 85% | 73.2% | 60.9%
5.81 | 90.59% | 86.6% | 84.9% | 73.1% | 60.7%
5.82 | 90.49% | 86.5% | 84.8% | 72.9% | 60.5%
5.83 | 90.39% | 86.4% | 84.6% | 72.7% | 60.4%
5.84 | 90.29% | 86.3% | 84.5% | 72.6% | 60.2%
5.85 | 90.19% | 86.2% | 84.4% | 72.4% | 60%
5.86 | 90.19% | 86.1% | 84.3% | 72.2% | 59.8%
5.87 | 90.09% | 86% | 84.2% | 72.1% | 59.7%
5.88 | 89.99% | 85.9% | 84% | 71.9% | 59.5%
5.89 | 89.89% | 85.8% | 83.9% | 71.7% | 59.3%
5.9 | 89.79% | 85.7% | 83.8% | 71.6% | 59.1%
5.91 | 89.79% | 85.6% | 83.7% | 71.4% | 58.9%
5.92 | 89.69% | 85.5% | 83.5% | 71.3% | 58.8%
5.93 | 89.59% | 85.4% | 83.4% | 71.1% | 58.6%
5.94 | 89.49% | 85.3% | 83.3% | 71% | 58.4%
5.95 | 89.39% | 85.2% | 83.2% | 70.8% | 58.2%
5.96 | 89.29% | 85.1% | 83% | 70.6% | 58%
5.97 | 89.29% | 85% | 82.9% | 70.5% | 57.8%
5.98 | 89.19% | 84.9% | 82.8% | 70.3% | 57.6%
5.99 | 89.09% | 84.8% | 82.7% | 70.2% | 57.4%
6 | 88.99% | 84.7% | 82.5% | 70% | 57.2%
6.01 | 88.89% | 84.6% | 82.4% | 69.8% | 57%
6.02 | 88.79% | 84.4% | 82.3% | 69.7% | 56.8%
6.03 | 88.69% | 84.3% | 82.2% | 69.5% | 56.7%
6.04 | 88.59% | 84.2% | 82.1% | 69.3% | 56.5%
6.05 | 88.49% | 84% | 81.9% | 69.1% | 56.3%
6.06 | 88.39% | 83.9% | 81.8% | 68.9% | 56.1%
6.07 | 88.29% | 83.8% | 81.7% | 68.7% | 55.9%
6.08 | 88.19% | 83.7% | 81.6% | 68.5% | 55.8%
6.09 | 88.09% | 83.5% | 81.5% | 68.3% | 55.6%
6.1 | 87.99% | 83.4% | 81.3% | 68.1% | 55.4%
6.11 | 87.89% | 83.3% | 81.2% | 68% | 55.3%
6.12 | 87.79% | 83.2% | 81.1% | 67.8% | 55.1%
6.13 | 87.69% | 83% | 81% | 67.6% | 54.9%
6.14 | 87.59% | 82.9% | 80.8% | 67.4% | 54.8%
6.15 | 87.49% | 82.8% | 80.7% | 67.3% | 54.6%
6.16 | 87.39% | 82.6% | 80.6% | 67.1% | 54.4%
6.17 | 87.29% | 82.5% | 80.4% | 66.9% | 54.3%
6.18 | 87.19% | 82.4% | 80.3% | 66.8% | 54.1%
6.19 | 87.09% | 82.3% | 80.2% | 66.6% | 53.9%
6.2 | 86.99% | 82.1% | 80% | 66.4% | 53.8%
6.21 | 86.89% | 82% | 79.9% | 66.2% | 53.6%
6.22 | 86.69% | 81.9% | 79.8% | 66.1% | 53.4%
6.23 | 86.59% | 81.8% | 79.6% | 65.9% | 53.2%
6.24 | 86.49% | 81.6% | 79.5% | 65.7% | 53%
6.25 | 86.39% | 81.5% | 79.3% | 65.6% | 52.9%
6.26 | 86.29% | 81.4% | 79.1% | 65.4% | 52.7%
6.27 | 86.19% | 81.3% | 79% | 65.2% | 52.5%
6.28 | 86.09% | 81.1% | 78.8% | 65% | 52.3%
6.29 | 85.99% | 81% | 78.6% | 64.8% | 52.1%
6.3 | 85.89% | 80.9% | 78.5% | 64.6% | 52%
6.31 | 85.79% | 80.8% | 78.3% | 64.5% | 51.8%
6.32 | 85.69% | 80.6% | 78.1% | 64.3% | 51.6%
6.33 | 85.49% | 80.5% | 78% | 64.1% | 51.4%
6.34 | 85.39% | 80.4% | 77.8% | 63.9% | 51.3%
6.35 | 85.29% | 80.2% | 77.7% | 63.7% | 51.1%
6.36 | 85.19% | 80.1% | 77.5% | 63.5% | 50.9%
6.37 | 85.09% | 79.9% | 77.4% | 63.4% | 50.8%
6.38 | 84.98% | 79.8% | 77.2% | 63.2% | 50.6%
6.39 | 84.78% | 79.6% | 77.1% | 63% | 50.4%
6.4 | 84.68% | 79.4% | 76.9% | 62.9% | 50.3%
6.41 | 84.58% | 79.3% | 76.8% | 62.7% | 50.1%
6.42 | 84.48% | 79.1% | 76.6% | 62.5% | 49.9%
6.43 | 84.38% | 79% | 76.5% | 62.4% | 49.8%
6.44 | 84.18% | 78.8% | 76.3% | 62.2% | 49.6%
6.45 | 84.08% | 78.7% | 76.2% | 62% | 49.4%
6.46 | 83.98% | 78.6% | 76% | 61.9% | 49.3%
6.47 | 83.78% | 78.4% | 75.9% | 61.7% | 49.1%
6.48 | 83.68% | 78.3% | 75.7% | 61.5% | 48.9%
6.49 | 83.58% | 78.2% | 75.5% | 61.3% | 48.8%
6.5 | 83.38% | 78% | 75.4% | 61.1% | 48.6%
6.51 | 83.28% | 77.9% | 75.2% | 61% | 48.4%
6.52 | 83.18% | 77.8% | 75.1% | 60.8% | 48.3%
6.53 | 82.98% | 77.6% | 74.9% | 60.6% | 48.1%
6.54 | 82.88% | 77.5% | 74.8% | 60.4% | 47.9%
6.55 | 82.78% | 77.4% | 74.6% | 60.2% | 47.8%
6.56 | 82.58% | 77.2% | 74.4% | 60% | 47.6%
6.57 | 82.48% | 77.1% | 74.3% | 59.8% | 47.4%
6.58 | 82.38% | 76.9% | 74.1% | 59.6% | 47.3%
6.59 | 82.18% | 76.8% | 74% | 59.4% | 47.1%
6.6 | 82.08% | 76.6% | 73.8% | 59.2% | 46.9%
6.61 | 81.98% | 76.5% | 73.7% | 59% | 46.8%
6.62 | 81.78% | 76.3% | 73.5% | 58.8% | 46.6%
6.63 | 81.68% | 76.2% | 73.3% | 58.7% | 46.4%
6.64 | 81.58% | 76% | 73.2% | 58.5% | 46.3%
6.65 | 81.38% | 75.8% | 73% | 58.3% | 46.1%
6.66 | 81.28% | 75.7% | 72.9% | 58.1% | 45.9%
6.67 | 81.08% | 75.5% | 72.7% | 57.9% | 45.8%
6.68 | 80.98% | 75.4% | 72.6% | 57.7% | 45.6%
6.69 | 80.88% | 75.2% | 72.4% | 57.5% | 45.4%
6.7 | 80.68% | 75.1% | 72.3% | 57.3% | 45.3%
6.71 | 80.58% | 74.9% | 72.1% | 57.1% | 45.1%
6.72 | 80.38% | 74.8% | 71.9% | 56.9% | 45%
6.73 | 80.28% | 74.6% | 71.8% | 56.8% | 44.8%
6.74 | 80.08% | 74.5% | 71.6% | 56.6% | 44.7%
6.75 | 79.98% | 74.3% | 71.4% | 56.5% | 44.5%
6.76 | 79.78% | 74.2% | 71.3% | 56.3% | 44.3%
6.77 | 79.68% | 74% | 71.1% | 56.2% | 44.2%
6.78 | 79.48% | 73.9% | 71% | 56% | 44%
6.79 | 79.38% | 73.7% | 70.8% | 55.9% | 43.9%
6.8 | 79.18% | 73.6% | 70.6% | 55.7% | 43.7%
6.81 | 79.08% | 73.4% | 70.5% | 55.6% | 43.6%
6.82 | 78.88% | 73.3% | 70.3% | 55.4% | 43.4%
6.83 | 78.78% | 73.1% | 70.1% | 55.3% | 43.2%
6.84 | 78.58% | 72.9% | 70% | 55.1% | 43%
6.85 | 78.38% | 72.7% | 69.8% | 55% | 42.9%
6.86 | 78.28% | 72.5% | 69.6% | 54.8% | 42.7%
6.87 | 78.08% | 72.4% | 69.5% | 54.6% | 42.5%
6.88 | 77.98% | 72.2% | 69.3% | 54.4% | 42.3%
6.89 | 77.78% | 72% | 69.1% | 54.3% | 42.2%
6.9 | 77.58% | 71.8% | 69% | 54.1% | 42%
6.91 | 77.48% | 71.7% | 68.8% | 53.9% | 41.8%
6.92 | 77.28% | 71.5% | 68.6% | 53.7% | 41.7%
6.93 | 77.08% | 71.3% | 68.5% | 53.5% | 41.5%
6.94 | 76.98% | 71.1% | 68.3% | 53.4% | 41.4%
6.95 | 76.78% | 71% | 68.1% | 53.2% | 41.3%
6.96 | 76.58% | 70.8% | 67.9% | 53% | 41.1%
6.97 | 76.48% | 70.6% | 67.8% | 52.8% | 41%
6.98 | 76.28% | 70.4% | 67.6% | 52.7% | 40.9%
6.99 | 76.18% | 70.2% | 67.4% | 52.5% | 40.7%
7 | 75.98% | 70.1% | 67.2% | 52.3% | 40.6%
7.01 | 75.78% | 69.9% | 67.1% | 52.1% | 40.5%
7.02 | 75.58% | 69.7% | 66.9% | 52% | 40.3%
7.03 | 75.48% | 69.5% | 66.7% | 51.8% | 40.2%
7.04 | 75.28% | 69.3% | 66.5% | 51.6% | 40.1%
7.05 | 75.08% | 69.2% | 66.4% | 51.4% | 39.9%
7.06 | 74.87% | 69% | 66.2% | 51.3% | 39.8%
7.07 | 74.77% | 68.8% | 66% | 51.1% | 39.6%
7.08 | 74.57% | 68.7% | 65.8% | 50.9% | 39.4%
7.09 | 74.37% | 68.5% | 65.6% | 50.7% | 39.3%
7.1 | 74.27% | 68.4% | 65.5% | 50.6% | 39.1%
7.11 | 74.07% | 68.2% | 65.3% | 50.4% | 39%
7.12 | 73.87% | 68.1% | 65.1% | 50.2% | 38.8%
7.13 | 73.67% | 67.9% | 64.9% | 50% | 38.6%
7.14 | 73.57% | 67.7% | 64.7% | 49.8% | 38.5%
7.15 | 73.37% | 67.6% | 64.5% | 49.7% | 38.3%
7.16 | 73.17% | 67.4% | 64.4% | 49.5% | 38.1%
7.17 | 72.97% | 67.3% | 64.2% | 49.3% | 38%
7.18 | 72.87% | 67.1% | 64% | 49.1% | 37.9%
7.19 | 72.67% | 66.9% | 63.8% | 48.9% | 37.7%
7.2 | 72.47% | 66.7% | 63.6% | 48.7% | 37.6%
7.21 | 72.27% | 66.5% | 63.4% | 48.6% | 37.4%
7.22 | 72.07% | 66.3% | 63.2% | 48.4% | 37.3%
7.23 | 71.97% | 66.1% | 63% | 48.2% | 37.2%
7.24 | 71.77% | 65.8% | 62.7% | 48.1% | 37%
7.25 | 71.57% | 65.6% | 62.5% | 47.9% | 36.9%
7.26 | 71.37% | 65.4% | 62.3% | 47.7% | 36.8%
7.27 | 71.17% | 65.2% | 62.1% | 47.5% | 36.6%
7.28 | 70.97% | 65% | 61.9% | 47.4% | 36.5%
7.29 | 70.77% | 64.8% | 61.7% | 47.2% | 36.3%
7.3 | 70.67% | 64.6% | 61.5% | 47% | 36.2%
7.31 | 70.47% | 64.4% | 61.3% | 46.8% | 36%
7.32 | 70.27% | 64.2% | 61.1% | 46.7% | 35.9%
7.33 | 70.07% | 63.9% | 60.9% | 46.5% | 35.7%
7.34 | 69.87% | 63.7% | 60.8% | 46.3% | 35.6%
7.35 | 69.67% | 63.5% | 60.6% | 46.2% | 35.4%
7.36 | 69.47% | 63.2% | 60.4% | 46% | 35.3%
7.37 | 69.27% | 63% | 60.2% | 45.8% | 35.1%
7.38 | 69.07% | 62.8% | 60.1% | 45.7% | 35%
7.39 | 68.77% | 62.6% | 59.9% | 45.5% | 34.9%
7.4 | 68.57% | 62.3% | 59.7% | 45.4% | 34.7%
7.41 | 68.37% | 62.1% | 59.5% | 45.2% | 34.6%
7.42 | 68.17% | 61.9% | 59.4% | 45% | 34.5%
7.43 | 67.97% | 61.7% | 59.1% | 44.9% | 34.4%
7.44 | 67.67% | 61.5% | 58.9% | 44.7% | 34.2%
7.45 | 67.47% | 61.4% | 58.7% | 44.5% | 34.1%
7.46 | 67.27% | 61.2% | 58.5% | 44.4% | 34%
7.47 | 67.07% | 61% | 58.2% | 44.2% | 33.8%
7.48 | 66.87% | 60.8% | 58% | 44% | 33.7%
7.49 | 66.67% | 60.6% | 57.8% | 43.8% | 33.6%
7.5 | 66.47% | 60.4% | 57.6% | 43.6% | 33.4%
7.51 | 66.27% | 60.2% | 57.3% | 43.5% | 33.2%
7.52 | 66.07% | 60.1% | 57.1% | 43.3% | 33.1%
7.53 | 65.87% | 59.9% | 56.9% | 43.1% | 32.9%
7.54 | 65.67% | 59.7% | 56.7% | 42.9% | 32.8%
7.55 | 65.47% | 59.4% | 56.5% | 42.8% | 32.6%
7.56 | 65.17% | 59.2% | 56.3% | 42.6% | 32.5%
7.57 | 64.96% | 59% | 56.1% | 42.4% | 32.3%
7.58 | 64.76% | 58.7% | 55.9% | 42.2% | 32.2%
7.59 | 64.56% | 58.5% | 55.7% | 42.1% | 32.1%
7.6 | 64.36% | 58.3% | 55.5% | 41.9% | 31.9%
7.61 | 64.16% | 58% | 55.3% | 41.7% | 31.8%
7.62 | 63.86% | 57.8% | 55.1% | 41.5% | 31.6%
7.63 | 63.66% | 57.6% | 54.9% | 41.4% | 31.5%
7.64 | 63.46% | 57.3% | 54.7% | 41.2% | 31.4%
7.65 | 63.16% | 57.1% | 54.5% | 41% | 31.2%
7.66 | 62.96% | 56.9% | 54.3% | 40.8% | 31.1%
7.67 | 62.66% | 56.7% | 54% | 40.7% | 31%
7.68 | 62.46% | 56.6% | 53.8% | 40.5% | 30.8%
7.69 | 62.26% | 56.4% | 53.6% | 40.3% | 30.7%
7.7 | 61.96% | 56.2% | 53.4% | 40.1% | 30.6%
7.71 | 61.76% | 56.1% | 53.2% | 40% | 30.4%
7.72 | 61.46% | 55.9% | 53% | 39.8% | 30.3%
7.73 | 61.26% | 55.8% | 52.8% | 39.6% | 30.2%
7.74 | 61.06% | 55.6% | 52.5% | 39.4% | 30%
7.75 | 60.86% | 55.4% | 52.3% | 39.2% | 29.9%
7.76 | 60.56% | 55.3% | 52.1% | 39.1% | 29.7%
7.77 | 60.36% | 55.1% | 51.9% | 38.9% | 29.6%
7.78 | 60.16% | 54.9% | 51.7% | 38.7% | 29.5%
7.79 | 59.96% | 54.6% | 51.5% | 38.5% | 29.3%
7.8 | 59.66% | 54.4% | 51.2% | 38.3% | 29.2%
7.81 | 59.46% | 54.1% | 51% | 38.2% | 29.1%
7.82 | 59.26% | 53.9% | 50.8% | 38% | 28.9%
7.83 | 58.96% | 53.6% | 50.6% | 37.9% | 28.8%
7.84 | 58.76% | 53.4% | 50.3% | 37.7% | 28.7%
7.85 | 58.56% | 53.1% | 50.1% | 37.6% | 28.5%
7.86 | 58.26% | 52.8% | 49.9% | 37.4% | 28.4%
7.87 | 58.06% | 52.6% | 49.7% | 37.2% | 28.3%
7.88 | 57.76% | 52.3% | 49.4% | 37.1% | 28.2%
7.89 | 57.56% | 52.1% | 49.2% | 36.9% | 28.1%
7.9 | 57.36% | 51.9% | 49% | 36.8% | 28%
7.91 | 57.06% | 51.6% | 48.8% | 36.6% | 27.8%
7.92 | 56.86% | 51.4% | 48.6% | 36.5% | 27.7%
7.93 | 56.56% | 51.2% | 48.4% | 36.3% | 27.6%
7.94 | 56.36% | 50.9% | 48.2% | 36.2% | 27.5%
7.95 | 56.06% | 50.7% | 48% | 36% | 27.4%
7.96 | 55.86% | 50.5% | 47.8% | 35.9% | 27.2%
7.97 | 55.56% | 50.2% | 47.6% | 35.7% | 27.1%
7.98 | 55.36% | 50% | 47.4% | 35.6% | 27%
7.99 | 55.16% | 49.8% | 47.2% | 35.4% | 26.8%
8 | 54.85% | 49.6% | 46.9% | 35.3% | 26.7%
8.01 | 54.65% | 49.3% | 46.7% | 35.1% | 26.6%
8.02 | 54.45% | 49.1% | 46.5% | 34.9% | 26.4%
8.03 | 54.15% | 48.8% | 46.3% | 34.8% | 26.3%
8.04 | 53.95% | 48.6% | 46% | 34.6% | 26.2%
8.05 | 53.75% | 48.3% | 45.8% | 34.5% | 26.1%
8.06 | 53.45% | 48.1% | 45.6% | 34.3% | 25.9%
8.07 | 53.25% | 47.8% | 45.3% | 34.1% | 25.8%
8.08 | 52.95% | 47.6% | 45.1% | 33.9% | 25.7%
8.09 | 52.75% | 47.3% | 44.8% | 33.8% | 25.6%
8.1 | 52.55% | 47% | 44.6% | 33.6% | 25.5%
8.11 | 52.25% | 46.8% | 44.4% | 33.4% | 25.4%
8.12 | 51.95% | 46.5% | 44.1% | 33.2% | 25.2%
8.13 | 51.75% | 46.3% | 43.9% | 33.1% | 25.1%
8.14 | 51.45% | 46% | 43.6% | 32.9% | 25%
8.15 | 51.25% | 45.8% | 43.4% | 32.7% | 24.9%
8.16 | 50.95% | 45.6% | 43.2% | 32.5% | 24.7%
8.17 | 50.65% | 45.4% | 43% | 32.4% | 24.6%
8.18 | 50.45% | 45.1% | 42.8% | 32.2% | 24.5%
8.19 | 50.15% | 44.9% | 42.6% | 32% | 24.3%
8.2 | 49.95% | 44.7% | 42.4% | 31.9% | 24.2%
8.21 | 49.65% | 44.5% | 42.2% | 31.7% | 24.1%
8.22 | 49.45% | 44.2% | 42% | 31.5% | 24%
8.23 | 49.15% | 44% | 41.8% | 31.4% | 23.8%
8.24 | 48.85% | 43.8% | 41.6% | 31.2% | 23.7%
8.25 | 48.65% | 43.5% | 41.4% | 31% | 23.6%
8.26 | 48.35% | 43.3% | 41.2% | 30.9% | 23.5%
8.27 | 48.15% | 43.1% | 40.9% | 30.7% | 23.3%
8.28 | 47.85% | 42.8% | 40.7% | 30.5% | 23.2%
8.29 | 47.55% | 42.6% | 40.4% | 30.4% | 23.1%
8.3 | 47.35% | 42.3% | 40.2% | 30.2% | 23%
8.31 | 47.05% | 42.1% | 39.9% | 30.1% | 22.8%
8.32 | 46.75% | 41.8% | 39.7% | 29.9% | 22.7%
8.33 | 46.55% | 41.6% | 39.4% | 29.8% | 22.6%
8.34 | 46.25% | 41.4% | 39.2% | 29.6% | 22.5%
8.35 | 45.95% | 41.1% | 38.9% | 29.5% | 22.4%
8.36 | 45.75% | 40.9% | 38.7% | 29.3% | 22.2%
8.37 | 45.45% | 40.6% | 38.4% | 29.2% | 22.1%
8.38 | 45.15% | 40.4% | 38.2% | 29% | 22%
8.39 | 44.94% | 40.1% | 37.9% | 28.9% | 21.9%
8.4 | 44.64% | 39.9% | 37.7% | 28.7% | 21.8%
8.41 | 44.34% | 39.6% | 37.5% | 28.6% | 21.7%
8.42 | 44.14% | 39.4% | 37.3% | 28.4% | 21.6%
8.43 | 43.84% | 39.1% | 37.1% | 28.2% | 21.5%
8.44 | 43.64% | 38.9% | 36.9% | 28.1% | 21.3%
8.45 | 43.34% | 38.6% | 36.6% | 27.9% | 21.2%
8.46 | 43.04% | 38.4% | 36.4% | 27.7% | 21.1%
8.47 | 42.84% | 38.1% | 36.2% | 27.5% | 21%
8.48 | 42.54% | 37.9% | 36% | 27.4% | 20.9%
8.49 | 42.34% | 37.6% | 35.8% | 27.2% | 20.7%
8.5 | 42.04% | 37.4% | 35.6% | 27% | 20.6%
8.51 | 41.74% | 37.1% | 35.3% | 26.9% | 20.5%
8.52 | 41.54% | 36.9% | 35.1% | 26.7% | 20.4%
8.53 | 41.24% | 36.6% | 34.9% | 26.5% | 20.3%
8.54 | 40.94% | 36.4% | 34.7% | 26.3% | 20.2%
8.55 | 40.74% | 36.1% | 34.4% | 26.2% | 20.1%
8.56 | 40.44% | 35.9% | 34.2% | 26% | 20%
8.57 | 40.14% | 35.7% | 34% | 25.8% | 19.8%
8.58 | 39.94% | 35.4% | 33.7% | 25.7% | 19.7%
8.59 | 39.64% | 35.2% | 33.5% | 25.5% | 19.6%
8.6 | 39.34% | 34.9% | 33.3% | 25.4% | 19.5%
8.61 | 39.14% | 34.7% | 33.1% | 25.2% | 19.4%
8.62 | 38.84% | 34.4% | 32.8% | 25% | 19.3%
8.63 | 38.64% | 34.2% | 32.6% | 24.9% | 19.2%
8.64 | 38.34% | 34% | 32.4% | 24.7% | 19%
8.65 | 38.04% | 33.8% | 32.2% | 24.5% | 18.9%
8.66 | 37.84% | 33.5% | 31.9% | 24.4% | 18.8%
8.67 | 37.54% | 33.3% | 31.7% | 24.3% | 18.7%
8.68 | 37.34% | 33.1% | 31.5% | 24.1% | 18.6%
8.69 | 37.04% | 32.8% | 31.3% | 24% | 18.4%
8.7 | 36.84% | 32.6% | 31% | 23.8% | 18.3%
8.71 | 36.54% | 32.4% | 30.8% | 23.7% | 18.2%
8.72 | 36.24% | 32.2% | 30.6% | 23.6% | 18.1%
8.73 | 36.04% | 31.9% | 30.4% | 23.4% | 18%
8.74 | 35.74% | 31.7% | 30.1% | 23.3% | 17.9%
8.75 | 35.44% | 31.5% | 29.9% | 23.1% | 17.8%
8.76 | 35.24% | 31.2% | 29.7% | 23% | 17.7%
8.77 | 34.93% | 31% | 29.5% | 22.9% | 17.6%
8.78 | 34.63% | 30.7% | 29.3% | 22.7% | 17.5%
8.79 | 34.43% | 30.5% | 29.1% | 22.6% | 17.4%
8.8 | 34.13% | 30.3% | 28.9% | 22.4% | 17.3%
8.81 | 33.83% | 30% | 28.7% | 22.3% | 17.2%
8.82 | 33.63% | 29.8% | 28.5% | 22.1% | 17.1%
8.83 | 33.33% | 29.5% | 28.4% | 22% | 17%
8.84 | 33.03% | 29.3% | 28.2% | 21.8% | 16.9%
8.85 | 32.83% | 29.1% | 28% | 21.6% | 16.8%
8.86 | 32.53% | 28.8% | 27.7% | 21.5% | 16.6%
8.87 | 32.33% | 28.6% | 27.5% | 21.3% | 16.5%
8.88 | 32.03% | 28.4% | 27.3% | 21.2% | 16.4%
8.89 | 31.83% | 28.1% | 27% | 21% | 16.3%
8.9 | 31.53% | 27.9% | 26.8% | 20.9% | 16.1%
8.91 | 31.33% | 27.7% | 26.5% | 20.7% | 16%
8.92 | 31.13% | 27.5% | 26.3% | 20.6% | 15.9%
8.93 | 30.83% | 27.2% | 26% | 20.4% | 15.8%
8.94 | 30.63% | 27% | 25.8% | 20.3% | 15.7%
8.95 | 30.33% | 26.8% | 25.6% | 20.1% | 15.6%
8.96 | 30.13% | 26.5% | 25.3% | 20% | 15.5%
8.97 | 29.83% | 26.3% | 25.1% | 19.8% | 15.4%
8.98 | 29.63% | 26.1% | 24.9% | 19.7% | 15.3%
8.99 | 29.33% | 25.9% | 24.7% | 19.5% | 15.2%
9 | 29.13% | 25.7% | 24.5% | 19.4% | 15.1%
9.01 | 28.93% | 25.5% | 24.3% | 19.2% | 15%
9.02 | 28.63% | 25.2% | 24.1% | 19.1% | 14.9%
9.03 | 28.43% | 25% | 23.9% | 19% | 14.8%
9.04 | 28.13% | 24.8% | 23.7% | 18.8% | 14.7%
9.05 | 27.93% | 24.6% | 23.6% | 18.7% | 14.6%
9.06 | 27.63% | 24.4% | 23.4% | 18.5% | 14.5%
9.07 | 27.43% | 24.2% | 23.2% | 18.4% | 14.4%
9.08 | 27.23% | 24% | 23% | 18.3% | 14.3%
9.09 | 26.93% | 23.8% | 22.8% | 18.1% | 14.2%
9.1 | 26.73% | 23.6% | 22.6% | 18% | 14.1%
9.11 | 26.53% | 23.3% | 22.5% | 17.8% | 14%
9.12 | 26.23% | 23.1% | 22.3% | 17.7% | 14%
9.13 | 26.03% | 22.9% | 22.1% | 17.6% | 13.9%
9.14 | 25.83% | 22.7% | 21.9% | 17.5% | 13.8%
9.15 | 25.63% | 22.5% | 21.7% | 17.4% | 13.7%
9.16 | 25.33% | 22.2% | 21.5% | 17.3% | 13.6%
9.17 | 25.13% | 22% | 21.4% | 17.1% | 13.5%
9.18 | 24.92% | 21.8% | 21.2% | 17% | 13.4%
9.19 | 24.72% | 21.6% | 21% | 16.9% | 13.3%
9.2 | 24.42% | 21.4% | 20.8% | 16.8% | 13.2%
9.21 | 24.22% | 21.2% | 20.6% | 16.7% | 13.1%
9.22 | 24.02% | 21% | 20.4% | 16.6% | 13.1%
9.23 | 23.82% | 20.8% | 20.2% | 16.5% | 13%
9.24 | 23.62% | 20.6% | 20% | 16.4% | 12.9%
9.25 | 23.32% | 20.5% | 19.8% | 16.3% | 12.8%
9.26 | 23.12% | 20.3% | 19.6% | 16.1% | 12.7%
9.27 | 22.92% | 20.1% | 19.4% | 16% | 12.6%
9.28 | 22.72% | 20% | 19.3% | 15.8% | 12.5%
9.29 | 22.52% | 19.8% | 19.1% | 15.7% | 12.5%
9.3 | 22.22% | 19.6% | 18.9% | 15.5% | 12.4%
9.31 | 22.02% | 19.5% | 18.7% | 15.4% | 12.3%
9.32 | 21.82% | 19.3% | 18.5% | 15.3% | 12.2%
9.33 | 21.62% | 19.1% | 18.3% | 15.1% | 12.1%
9.34 | 21.42% | 18.9% | 18.1% | 15% | 12%
9.35 | 21.22% | 18.7% | 17.9% | 14.8% | 11.9%
9.36 | 21.02% | 18.5% | 17.8% | 14.7% | 11.9%
9.37 | 20.82% | 18.4% | 17.6% | 14.5% | 11.8%
9.38 | 20.62% | 18.2% | 17.4% | 14.4% | 11.7%
9.39 | 20.42% | 18% | 17.3% | 14.3% | 11.6%
9.4 | 20.22% | 17.8% | 17.1% | 14.2% | 11.5%
9.41 | 20.02% | 17.6% | 16.9% | 14.1% | 11.4%
9.42 | 19.82% | 17.4% | 16.8% | 14% | 11.4%
9.43 | 19.62% | 17.2% | 16.6% | 13.9% | 11.3%
9.44 | 19.42% | 17% | 16.4% | 13.7% | 11.2%
9.45 | 19.22% | 16.8% | 16.3% | 13.6% | 11.1%
9.46 | 19.02% | 16.7% | 16.1% | 13.5% | 11%
9.47 | 18.82% | 16.5% | 15.9% | 13.4% | 10.9%
9.48 | 18.72% | 16.3% | 15.8% | 13.3% | 10.8%
9.49 | 18.52% | 16.2% | 15.6% | 13.2% | 10.7%
9.5 | 18.32% | 16% | 15.5% | 13.1% | 10.6%
9.51 | 18.12% | 15.9% | 15.3% | 13% | 10.6%
9.52 | 17.92% | 15.7% | 15.1% | 12.9% | 10.5%
9.53 | 17.72% | 15.6% | 15% | 12.8% | 10.4%
9.54 | 17.52% | 15.4% | 14.8% | 12.7% | 10.3%
9.55 | 17.32% | 15.3% | 14.7% | 12.6% | 10.3%
9.56 | 17.22% | 15.1% | 14.5% | 12.4% | 10.2%
9.57 | 17.02% | 14.9% | 14.3% | 12.3% | 10.1%
9.58 | 16.82% | 14.8% | 14.2% | 12.2% | 10%
9.59 | 16.72% | 14.7% | 14.1% | 12.1% | 10%
9.6 | 16.52% | 14.5% | 13.9% | 12% | 9.9%
9.61 | 16.42% | 14.4% | 13.8% | 11.9% | 9.8%
9.62 | 16.22% | 14.3% | 13.7% | 11.8% | 9.8%
9.63 | 16.02% | 14.1% | 13.5% | 11.7% | 9.7%
9.64 | 15.92% | 14% | 13.4% | 11.6% | 9.6%
9.65 | 15.72% | 13.9% | 13.3% | 11.5% | 9.6%
9.66 | 15.62% | 13.7% | 13.1% | 11.4% | 9.5%
9.67 | 15.42% | 13.6% | 13% | 11.3% | 9.4%
9.68 | 15.32% | 13.5% | 12.9% | 11.2% | 9.4%
9.69 | 15.12% | 13.3% | 12.7% | 11.1% | 9.3%
9.7 | 14.91% | 13.2% | 12.6% | 11% | 9.2%
9.71 | 14.81% | 13.1% | 12.5% | 10.9% | 9.2%
9.72 | 14.61% | 12.9% | 12.4% | 10.8% | 9.1%
9.73 | 14.51% | 12.8% | 12.2% | 10.7% | 9%
9.74 | 14.31% | 12.6% | 12.1% | 10.6% | 9%
9.75 | 14.21% | 12.5% | 12% | 10.5% | 8.9%
9.76 | 14.01% | 12.4% | 11.9% | 10.5% | 8.8%
9.77 | 13.91% | 12.2% | 11.7% | 10.4% | 8.8%
9.78 | 13.71% | 12.1% | 11.6% | 10.3% | 8.7%
9.79 | 13.61% | 12% | 11.5% | 10.2% | 8.6%
9.8 | 13.41% | 11.8% | 11.3% | 10.2% | 8.6%
9.81 | 13.31% | 11.7% | 11.2% | 10.1% | 8.5%
9.82 | 13.21% | 11.5% | 11.1% | 10% | 8.5%
9.83 | 13.01% | 11.4% | 11% | 9.9% | 8.4%
9.84 | 12.91% | 11.3% | 10.9% | 9.9% | 8.3%
9.85 | 12.81% | 11.2% | 10.8% | 9.8% | 8.3%
9.86 | 12.71% | 11.1% | 10.7% | 9.7% | 8.2%
9.87 | 12.51% | 11% | 10.6% | 9.6% | 8.2%
9.88 | 12.41% | 10.9% | 10.5% | 9.5% | 8.1%
9.89 | 12.31% | 10.8% | 10.3% | 9.4% | 8%
9.9 | 12.21% | 10.7% | 10.2% | 9.4% | 8%
9.91 | 12.11% | 10.6% | 10.1% | 9.3% | 7.9%
9.92 | 11.91% | 10.5% | 10% | 9.2% | 7.9%
9.93 | 11.81% | 10.4% | 9.9% | 9.1% | 7.8%
9.94 | 11.71% | 10.3% | 9.8% | 9% | 7.7%
9.95 | 11.61% | 10.2% | 9.7% | 8.9% | 7.7%
9.96 | 11.41% | 10.1% | 9.6% | 8.8% | 7.6%
9.97 | 11.31% | 10% | 9.6% | 8.7% | 7.6%
9.98 | 11.21% | 9.9% | 9.5% | 8.7% | 7.5%
9.99 | 11.11% | 9.8% | 9.4% | 8.6% | 7.5%
10 | 11.01% | 9.7% | 9.3% | 8.5% | 7.5%
10.01 | 10.91% | 9.6% | 9.2% | 8.5% | 7.4%
10.02 | 10.71% | 9.5% | 9.2% | 8.4% | 7.4%
10.03 | 10.61% | 9.4% | 9.1% | 8.3% | 7.3%
10.04 | 10.51% | 9.3% | 9% | 8.3% | 7.3%
10.05 | 10.41% | 9.2% | 8.9% | 8.2% | 7.2%
10.06 | 10.31% | 9.1% | 8.8% | 8.2% | 7.2%
10.07 | 10.31% | 9% | 8.8% | 8.1% | 7.1%
10.08 | 10.21% | 8.9% | 8.7% | 8% | 7.1%
10.09 | 10.11% | 8.9% | 8.6% | 8% | 7%
10.1 | 10.01% | 8.8% | 8.5% | 7.9% | 7%
10.11 | 9.91% | 8.7% | 8.5% | 7.8% | 7%
10.12 | 9.81% | 8.7% | 8.4% | 7.8% | 6.9%
10.13 | 9.71% | 8.6% | 8.3% | 7.7% | 6.9%
10.14 | 9.71% | 8.5% | 8.2% | 7.7% | 6.8%
10.15 | 9.61% | 8.5% | 8.2% | 7.6% | 6.8%
10.16 | 9.51% | 8.4% | 8.1% | 7.6% | 6.8%
10.17 | 9.409% | 8.3% | 8% | 7.5% | 6.7%
10.18 | 9.309% | 8.3% | 7.9% | 7.5% | 6.7%
10.19 | 9.209% | 8.2% | 7.9% | 7.4% | 6.6%
10.2 | 9.109% | 8.1% | 7.8% | 7.3% | 6.6%
10.21 | 9.109% | 8.1% | 7.7% | 7.3% | 6.5%
10.22 | 9.009% | 8% | 7.7% | 7.2% | 6.5%
10.23 | 8.909% | 8% | 7.6% | 7.2% | 6.5%
10.24 | 8.809% | 7.9% | 7.5% | 7.1% | 6.4%
10.25 | 8.709% | 7.8% | 7.5% | 7.1% | 6.4%
10.26 | 8.709% | 7.8% | 7.4% | 7% | 6.4%
10.27 | 8.609% | 7.7% | 7.3% | 7% | 6.3%
10.28 | 8.509% | 7.6% | 7.3% | 6.9% | 6.3%
10.29 | 8.408% | 7.6% | 7.2% | 6.9% | 6.2%
10.3 | 8.408% | 7.5% | 7.1% | 6.8% | 6.2%
10.31 | 8.308% | 7.5% | 7.1% | 6.8% | 6.2%
10.32 | 8.208% | 7.4% | 7% | 6.7% | 6.2%
10.33 | 8.208% | 7.4% | 7% | 6.7% | 6.1%
10.34 | 8.108% | 7.3% | 6.9% | 6.6% | 6.1%
10.35 | 8.008% | 7.3% | 6.9% | 6.6% | 6.1%
10.36 | 8.008% | 7.2% | 6.9% | 6.6% | 6%
10.37 | 7.908% | 7.2% | 6.8% | 6.5% | 6%
10.38 | 7.908% | 7.1% | 6.8% | 6.5% | 6%
10.39 | 7.808% | 7.1% | 6.7% | 6.5% | 6%
10.4 | 7.808% | 7% | 6.7% | 6.4% | 5.9%
10.41 | 7.708% | 7% | 6.6% | 6.4% | 5.9%
10.42 | 7.708% | 6.9% | 6.6% | 6.3% | 5.9%
10.43 | 7.608% | 6.9% | 6.6% | 6.3% | 5.9%
10.44 | 7.608% | 6.8% | 6.5% | 6.3% | 5.8%
10.45 | 7.508% | 6.8% | 6.5% | 6.2% | 5.8%
10.46 | 7.508% | 6.8% | 6.4% | 6.2% | 5.8%
10.47 | 7.407% | 6.7% | 6.4% | 6.2% | 5.8%
10.48 | 7.407% | 6.7% | 6.3% | 6.1% | 5.7%
10.49 | 7.307% | 6.6% | 6.3% | 6.1% | 5.7%
10.5 | 7.307% | 6.6% | 6.3% | 6.1% | 5.7%
10.51 | 7.307% | 6.6% | 6.2% | 6% | 5.7%
10.52 | 7.207% | 6.5% | 6.2% | 6% | 5.7%
10.53 | 7.207% | 6.5% | 6.1% | 6% | 5.6%
10.54 | 7.107% | 6.5% | 6.1% | 6% | 5.6%
10.55 | 7.107% | 6.4% | 6.1% | 5.9% | 5.6%
10.56 | 7.107% | 6.4% | 6% | 5.9% | 5.6%
10.57 | 7.007% | 6.4% | 6% | 5.9% | 5.6%
10.58 | 7.007% | 6.4% | 6% | 5.8% | 5.5%
10.59 | 6.907% | 6.3% | 6% | 5.8% | 5.5%
10.6 | 6.907% | 6.3% | 6% | 5.8% | 5.5%
10.61 | 6.907% | 6.3% | 5.9% | 5.8% | 5.5%
10.62 | 6.807% | 6.3% | 5.9% | 5.7% | 5.5%
10.63 | 6.807% | 6.2% | 5.9% | 5.7% | 5.5%
10.64 | 6.807% | 6.2% | 5.9% | 5.7% | 5.4%
10.65 | 6.707% | 6.2% | 5.8% | 5.7% | 5.4%
10.66 | 6.707% | 6.2% | 5.8% | 5.7% | 5.4%
10.67 | 6.707% | 6.2% | 5.8% | 5.6% | 5.4%
10.68 | 6.607% | 6.1% | 5.8% | 5.6% | 5.4%
10.69 | 6.607% | 6.1% | 5.7% | 5.6% | 5.4%
10.7 | 6.607% | 6.1% | 5.7% | 5.6% | 5.4%
10.71 | 6.607% | 6.1% | 5.7% | 5.5% | 5.3%
10.72 | 6.507% | 6% | 5.7% | 5.5% | 5.3%
10.73 | 6.507% | 6% | 5.6% | 5.5% | 5.3%
10.74 | 6.507% | 6% | 5.6% | 5.5% | 5.3%
10.75 | 6.507% | 6% | 5.6% | 5.5% | 5.3%
10.76 | 6.406% | 5.9% | 5.6% | 5.5% | 5.3%
10.77 | 6.406% | 5.9% | 5.5% | 5.4% | 5.3%
10.78 | 6.406% | 5.9% | 5.5% | 5.4% | 5.2%
10.79 | 6.406% | 5.9% | 5.5% | 5.4% | 5.2%
10.8 | 6.306% | 5.9% | 5.5% | 5.4% | 5.2%
10.81 | 6.306% | 5.8% | 5.5% | 5.4% | 5.2%
10.82 | 6.306% | 5.8% | 5.4% | 5.4% | 5.2%
10.83 | 6.306% | 5.8% | 5.4% | 5.3% | 5.2%
10.84 | 6.306% | 5.8% | 5.4% | 5.3% | 5.2%
10.85 | 6.206% | 5.8% | 5.4% | 5.3% | 5.2%
10.86 | 6.206% | 5.7% | 5.4% | 5.3% | 5.2%
10.87 | 6.206% | 5.7% | 5.4% | 5.3% | 5.1%
10.88 | 6.206% | 5.7% | 5.3% | 5.3% | 5.1%
10.89 | 6.206% | 5.7% | 5.3% | 5.3% | 5.1%
10.9 | 6.206% | 5.7% | 5.3% | 5.2% | 5.1%
10.91 | 6.106% | 5.7% | 5.3% | 5.2% | 5.1%
10.92 | 6.106% | 5.6% | 5.3% | 5.2% | 5.1%
10.93 | 6.106% | 5.6% | 5.3% | 5.2% | 5.1%
10.94 | 6.106% | 5.6% | 5.3% | 5.2% | 5%
10.95 | 6.106% | 5.6% | 5.2% | 5.2% | 0.05
10.96 | 6.106% | 5.6% | 5.2% | 5.2% | 0.05
10.97 | 6.106% | 5.6% | 5.2% | 5.2% | 0.05
10.98 | 6.006% | 5.6% | 5.2% | 5.1% | 0.05
10.99 | 6.006% | 5.6% | 5.2% | 5.1% | 0.05
11 | 6.006% | 5.6% | 5.2% | 5.1% | 0.05
11.01 | 6.006% | 5.5% | 5.2% | 5.1% | 0.05
11.02 | 6.006% | 5.5% | 5.2% | 5.1% | 0.05
11.03 | 6.006% | 5.5% | 5.2% | 5.1% | 0.05
11.04 | 6.006% | 5.5% | 5.2% | 5.1% | 0.049
11.05 | 6.006% | 5.5% | 5.1% | 5.1% | 0.049
11.06 | 5.906% | 5.5% | 5.1% | 5.1% | 0.049
11.07 | 5.906% | 5.5% | 5.1% | 5.1% | 0.049
11.08 | 5.906% | 5.5% | 5.1% | 5.1% | 0.049
11.09 | 5.906% | 5.5% | 5.1% | 5% | 0.049
11.1 | 5.906% | 5.5% | 5.1% | 5% | 0.049
11.11 | 5.906% | 5.5% | 5.1% | 5% | 0.049
11.12 | 5.906% | 5.5% | 5.1% | 5% | 0.049
11.13 | 5.906% | 5.4% | 5.1% | 5% | 0.049
11.14 | 5.906% | 5.4% | 5.1% | 5% | 0.049
11.15 | 5.906% | 5.4% | 5.1% | 5% | 0.049
11.16 | 5.806% | 5.4% | 5.1% | 5% | 0.049
11.17 | 5.806% | 5.4% | 5.1% | 4.9% | 0.049
11.18 | 5.806% | 5.4% | 5.1% | 4.9% | 0.049
11.19 | 5.806% | 5.4% | 5.1% | 4.9% | 0.049
11.2 | 5.806% | 5.4% | 5.1% | 4.9% | 0.049
11.21 | 5.806% | 5.4% | 5% | 4.9% | 0.048
11.22 | 5.806% | 5.4% | 5% | 4.9% | 0.048
11.23 | 5.806% | 5.4% | 5% | 4.9% | 0.048
11.24 | 5.806% | 5.4% | 5% | 4.9% | 0.048
11.25 | 5.806% | 5.4% | 5% | 4.9% | 0.048
11.26 | 5.806% | 5.4% | 5% | 4.9% | 0.048
11.27 | 5.806% | 5.3% | 5% | 4.9% | 0.048
11.28 | 5.806% | 5.3% | 5% | 4.9% | 0.048
11.29 | 5.806% | 5.3% | 5% | 4.9% | 0.048
11.3 | 5.806% | 5.3% | 5% | 4.9% | 0.048
11.31 | 5.806% | 5.3% | 5% | 4.9% | 0.048
11.32 | 5.706% | 5.3% | 5% | 4.9% | 0.048
11.33 | 5.706% | 5.3% | 5% | 4.9% | 0.048
11.34 | 5.706% | 5.3% | 5% | 4.9% | 0.048
11.35 | 5.706% | 5.3% | 4.9% | 4.9% | 0.048
11.36 | 5.706% | 5.3% | 4.9% | 4.9% | 0.048
11.37 | 5.706% | 5.3% | 4.9% | 4.9% | 0.048
11.38 | 5.706% | 5.3% | 4.9% | 4.9% | 0.048
11.39 | 5.706% | 5.3% | 4.9% | 4.9% | 0.048
11.4 | 5.706% | 5.3% | 4.9% | 4.9% | 0.048
11.41 | 5.706% | 5.3% | 4.9% | 4.8% | 0.048
11.42 | 5.706% | 5.3% | 4.9% | 4.8% | 0.048
11.43 | 5.706% | 5.3% | 4.9% | 4.8% | 0.048
11.44 | 5.706% | 5.3% | 4.9% | 4.8% | 0.047
11.45 | 5.706% | 5.3% | 4.9% | 4.8% | 0.047
11.46 | 5.706% | 5.3% | 4.9% | 4.8% | 0.047
11.47 | 5.706% | 5.2% | 4.9% | 4.8% | 0.047
11.48 | 5.706% | 5.2% | 4.9% | 4.8% | 0.047
11.49 | 5.706% | 5.2% | 4.9% | 4.8% | 0.047
11.5 | 5.706% | 5.2% | 4.9% | 4.8% | 0.047
11.51 | 5.706% | 5.2% | 4.9% | 4.8% | 0.047
11.52 | 5.606% | 5.2% | 4.9% | 4.8% | 0.047
11.53 | 5.606% | 5.2% | 4.9% | 4.8% | 0.047
11.54 | 5.606% | 5.2% | 4.9% | 4.8% | 0.047
11.55 | 5.606% | 5.2% | 4.9% | 4.8% | 0.047
11.56 | 5.606% | 5.2% | 4.9% | 4.8% | 0.047
11.57 | 5.606% | 5.2% | 4.9% | 4.8% | 0.047
11.58 | 5.606% | 5.2% | 4.9% | 4.8% | 0.047
11.59 | 5.606% | 5.2% | 4.9% | 4.8% | 0.047
11.6 | 5.606% | 5.2% | 4.8% | 4.8% | 0.047
11.61 | 5.606% | 5.2% | 4.8% | 4.8% | 0.047
11.62 | 5.606% | 5.2% | 4.8% | 4.8% | 0.047
11.63 | 0.056056 | 5.2% | 4.8% | 4.8% | 0.047
11.64 | 0.056056 | 5.2% | 4.8% | 4.8% | 0.047
11.65 | 0.056056 | 5.2% | 4.8% | 4.8% | 0.047
11.66 | 0.056056 | 5.2% | 4.8% | 4.7% | 0.047
11.67 | 0.056056 | 5.2% | 4.8% | 4.7% | 0.047
11.68 | 0.056056 | 5.2% | 4.8% | 4.7% | 0.047
11.69 | 0.056056 | 5.2% | 4.8% | 4.7% | 0.047
11.7 | 0.056056 | 5.2% | 4.8% | 4.7% | 0.047
11.71 | 0.056056 | 5.2% | 4.8% | 4.7% | 0.047
11.72 | 0.056056 | 5.2% | 4.8% | 4.7% | 0.047
11.73 | 0.055055 | 5.2% | 4.8% | 4.7% | 0.047
11.74 | 0.055055 | 5.2% | 4.8% | 4.7% | 0.047
11.75 | 0.055055 | 5.2% | 4.8% | 4.7% | 0.047
11.76 | 0.055055 | 5.2% | 4.8% | 4.7% | 0.046
11.77 | 0.055055 | 5.2% | 4.8% | 4.7% | 0.046
11.78 | 0.055055 | 5.2% | 4.8% | 4.7% | 0.046
11.79 | 0.055055 | 5.2% | 4.8% | 4.7% | 0.046
11.8 | 0.055055 | 5.2% | 4.8% | 4.7% | 0.046
11.81 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.82 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.83 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.84 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.85 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.86 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.87 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.88 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.89 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.9 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.91 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.92 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.93 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.94 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.95 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.96 | 0.055055 | 5.1% | 4.8% | 4.7% | 0.046
11.97 | 0.055055 | 5.1% | 4.8% | 4.6% | 0.046
11.98 | 0.055055 | 5.1% | 4.8% | 4.6% | 0.046
11.99 | 0.055055 | 5.1% | 4.8% | 4.6% | 0.046
12 | 0.055055 | 5.1% | 4.7% | 4.6% | 0.046
12.01 | 0.055055 | 5.1% | 4.7% | 4.6% | 0.046
12.02 | 0.055055 | 5.1% | 4.7% | 4.6% | 0.046
12.03 | 0.055055 | 5.1% | 4.7% | 4.6% | 0.046
12.04 | 0.055055 | 5.1% | 4.7% | 4.6% | 0.046
12.05 | 0.055055 | 5.1% | 4.7% | 4.6% | 0.046
12.06 | 0.055055 | 5.1% | 4.7% | 4.6% | 0.046
12.07 | 0.055055 | 5.1% | 4.7% | 4.6% | 0.045
12.08 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.09 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.1 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.11 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.12 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.13 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.14 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.15 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.16 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.17 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.18 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.19 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.2 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.21 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.22 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.23 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.24 | 0.054054 | 5% | 4.7% | 4.6% | 0.045
12.25 | 0.054054 | 5% | 4.7% | 4.5% | 0.045
12.26 | 0.054054 | 5% | 4.7% | 4.5% | 0.045
12.27 | 0.054054 | 5% | 4.7% | 4.5% | 0.045
12.28 | 0.054054 | 5% | 4.7% | 4.5% | 0.045
12.29 | 0.054054 | 5% | 4.6% | 4.5% | 0.045
12.3 | 0.054054 | 5% | 4.6% | 4.5% | 0.045
12.31 | 0.054054 | 5% | 4.6% | 4.5% | 0.045
12.32 | 0.054054 | 5% | 4.6% | 4.5% | 0.045
12.33 | 0.054054 | 5% | 4.6% | 4.5% | 0.045
12.34 | 0.054054 | 5% | 4.6% | 4.5% | 0.045
12.35 | 0.054054 | 5% | 4.6% | 4.5% | 0.045
12.36 | 0.054054 | 5% | 4.6% | 4.5% | 0.045
12.37 | 0.054054 | 5% | 4.6% | 4.5% | 0.045
12.38 | 0.054054 | 5% | 4.6% | 4.5% | 0.045
12.39 | 0.054054 | 5% | 4.6% | 4.5% | 0.045
12.4 | 0.053053 | 5% | 4.6% | 4.5% | 0.045
12.41 | 0.053053 | 5% | 4.6% | 4.5% | 0.045
12.42 | 0.053053 | 5% | 4.6% | 4.5% | 0.045
12.43 | 0.053053 | 5% | 4.6% | 4.5% | 0.045
12.44 | 0.053053 | 5% | 4.6% | 4.5% | 0.044
12.45 | 0.053053 | 4.9% | 4.6% | 4.5% | 0.044
12.46 | 0.053053 | 4.9% | 4.6% | 4.5% | 0.044
12.47 | 0.053053 | 4.9% | 4.6% | 4.5% | 0.044
12.48 | 0.053053 | 4.9% | 4.6% | 4.5% | 0.044
12.49 | 0.053053 | 4.9% | 4.6% | 4.5% | 0.044
12.5 | 0.053053 | 4.9% | 4.6% | 4.5% | 0.044
12.51 | 0.053053 | 4.9% | 4.6% | 4.5% | 0.044
12.52 | 0.053053 | 4.9% | 4.6% | 4.5% | 0.044
12.53 | 0.053053 | 4.9% | 4.6% | 4.5% | 0.044
12.54 | 0.053053 | 4.9% | 4.6% | 4.5% | 0.044
12.55 | 0.053053 | 4.9% | 4.6% | 4.5% | 0.044
12.56 | 0.053053 | 4.9% | 4.6% | 4.5% | 0.044
12.57 | 0.053053 | 4.9% | 4.6% | 4.5% | 0.044
12.58 | 0.053053 | 4.9% | 4.6% | 4.4% | 0.044
12.59 | 0.053053 | 4.9% | 4.6% | 4.4% | 0.044
12.6 | 0.053053 | 4.9% | 4.6% | 4.4% | 0.044
12.61 | 0.053053 | 4.9% | 4.6% | 4.4% | 0.044
12.62 | 0.053053 | 4.9% | 4.6% | 4.4% | 0.044
12.63 | 0.053053 | 4.9% | 4.6% | 4.4% | 0.044
12.64 | 0.053053 | 4.9% | 4.6% | 4.4% | 0.044
12.65 | 0.053053 | 4.9% | 4.6% | 4.4% | 0.044
12.66 | 0.053053 | 4.9% | 4.6% | 4.4% | 0.044
12.67 | 0.053053 | 4.9% | 4.6% | 4.4% | 0.044
12.68 | 0.053053 | 4.9% | 4.5% | 4.4% | 0.044
12.69 | 0.053053 | 4.9% | 4.5% | 4.4% | 0.044
12.7 | 0.053053 | 4.8% | 4.5% | 4.4% | 0.044
12.71 | 0.053053 | 4.8% | 4.5% | 4.4% | 0.044
12.72 | 0.053053 | 4.8% | 4.5% | 4.4% | 0.044
12.73 | 0.053053 | 4.8% | 4.5% | 4.4% | 0.044
12.74 | 0.053053 | 4.8% | 4.5% | 4.4% | 0.044
12.75 | 0.053053 | 4.8% | 4.5% | 4.4% | 0.044
12.76 | 0.053053 | 4.8% | 4.5% | 4.4% | 0.044
12.77 | 0.053053 | 4.8% | 4.5% | 4.4% | 0.044
12.78 | 0.053053 | 4.8% | 4.5% | 4.4% | 0.044
12.79 | 0.053053 | 4.8% | 4.5% | 4.4% | 0.043
12.8 | 0.053053 | 4.8% | 4.5% | 4.4% | 0.043
12.81 | 0.053053 | 4.8% | 4.5% | 4.4% | 0.044
12.82 | 0.052052 | 4.8% | 4.5% | 4.4% | 0.044
12.83 | 0.052052 | 4.8% | 4.5% | 4.4% | 0.044
12.84 | 0.052052 | 4.8% | 4.5% | 4.4% | 0.044
12.85 | 0.052052 | 4.8% | 4.5% | 4.4% | 0.044
12.86 | 0.052052 | 4.8% | 4.5% | 4.4% | 0.044
12.87 | 0.052052 | 4.8% | 4.5% | 4.4% | 0.044
12.88 | 0.052052 | 4.8% | 4.5% | 4.4% | 0.044
12.89 | 0.052052 | 4.8% | 4.5% | 4.4% | 0.044
12.9 | 0.052052 | 4.8% | 4.5% | 4.4% | 0.044
12.91 | 0.052052 | 4.8% | 4.5% | 4.4% | 0.044
12.92 | 0.052052 | 4.8% | 4.5% | 4.4% | 0.044
12.93 | 0.052052 | 4.8% | 4.5% | 4.4% | 0.044
12.94 | 0.052052 | 4.8% | 4.5% | 4.4% | 0.044
12.95 | 0.052052 | 4.7% | 4.5% | 4.4% | 0.044
12.96 | 0.052052 | 4.7% | 4.5% | 4.3% | 0.043
12.97 | 0.052052 | 4.7% | 4.5% | 4.3% | 0.043
12.98 | 0.052052 | 4.7% | 4.5% | 4.3% | 0.043
12.99 | 0.052052 | 4.7% | 4.5% | 4.3% | 0.043
13 | 0.052052 | 4.7% | 4.4% | 4.3% | 0.043
13.01 | 0.052052 | 4.7% | 4.4% | 4.3% | 0.043
13.02 | 0.052052 | 4.7% | 4.4% | 4.3% | 0.043
13.03 | 0.052052 | 4.7% | 4.4% | 4.3% | 0.043
13.04 | 0.052052 | 4.7% | 4.4% | 4.3% | 0.043
13.05 | 0.052052 | 4.7% | 4.4% | 4.3% | 0.043
13.06 | 0.052052 | 4.7% | 4.4% | 4.3% | 0.043
13.07 | 0.052052 | 4.7% | 4.4% | 4.3% | 0.043
13.08 | 0.052052 | 4.7% | 4.4% | 4.3% | 0.043
13.09 | 0.052052 | 4.7% | 4.4% | 4.3% | 0.043
13.1 | 0.052052 | 4.7% | 4.4% | 4.3% | 0.043
13.11 | 0.052052 | 4.7% | 4.4% | 4.3% | 0.043
13.12 | 0.051051 | 4.7% | 4.4% | 4.3% | 0.043
13.13 | 0.051051 | 4.7% | 4.4% | 4.3% | 0.043
13.14 | 0.051051 | 4.7% | 4.4% | 4.3% | 0.043
13.15 | 0.051051 | 4.7% | 4.4% | 4.3% | 0.043
13.16 | 0.051051 | 4.7% | 4.4% | 4.3% | 0.043
13.17 | 0.051051 | 4.7% | 4.4% | 4.3% | 0.043
13.18 | 0.051051 | 4.7% | 4.4% | 4.3% | 0.043
13.19 | 0.051051 | 4.7% | 4.4% | 4.3% | 0.043
13.2 | 0.051051 | 4.7% | 4.4% | 4.3% | 0.043
13.21 | 0.051051 | 4.7% | 4.4% | 4.3% | 0.043
13.22 | 0.051051 | 4.7% | 4.4% | 4.3% | 0.043
13.23 | 0.051051 | 4.7% | 4.4% | 4.3% | 0.043
13.24 | 0.051051 | 4.7% | 4.4% | 4.3% | 0.043
13.25 | 0.051051 | 4.7% | 4.4% | 4.2% | 0.043
13.26 | 0.051051 | 4.7% | 4.4% | 4.2% | 0.043
13.27 | 0.051051 | 4.7% | 4.4% | 4.2% | 0.043
13.28 | 0.051051 | 4.7% | 4.4% | 4.2% | 0.043
13.29 | 0.051051 | 4.7% | 4.4% | 4.2% | 0.042
13.3 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.31 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.32 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.33 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.34 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.35 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.36 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.37 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.38 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.39 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.4 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.41 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.42 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.43 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.44 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.45 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.46 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.47 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.48 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.49 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.5 | 0.051051 | 4.7% | 4.3% | 4.2% | 0.042
13.51 | 0.051051 | 4.6% | 4.3% | 4.2% | 0.042
13.52 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.042
13.53 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.042
13.54 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.042
13.55 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.042
13.56 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.042
13.57 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.042
13.58 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.041
13.59 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.041
13.6 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.041
13.61 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.041
13.62 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.041
13.63 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.041
13.64 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.041
13.65 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.041
13.66 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.041
13.67 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.041
13.68 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.041
13.69 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.041
13.7 | 0.05005 | 4.6% | 4.3% | 4.2% | 0.041
13.71 | 0.05005 | 4.6% | 4.3% | 4.1% | 0.041
13.72 | 0.05005 | 4.6% | 4.3% | 4.1% | 0.041
13.73 | 0.05005 | 4.6% | 4.3% | 4.1% | 0.041
13.74 | 0.05005 | 4.6% | 4.3% | 4.1% | 0.041
13.75 | 0.05005 | 4.6% | 4.3% | 4.1% | 0.041
13.76 | 0.05005 | 4.6% | 4.3% | 4.1% | 0.041
13.77 | 0.05005 | 4.6% | 4.3% | 4.1% | 0.041
13.78 | 0.05005 | 4.6% | 4.3% | 4.1% | 0.041
13.79 | 0.05005 | 4.6% | 4.3% | 4.1% | 0.041
13.8 | 0.05005 | 4.6% | 4.3% | 4.1% | 0.041
13.81 | 0.05005 | 4.6% | 4.3% | 4.1% | 0.041
13.82 | 0.05005 | 4.6% | 4.3% | 4.1% | 0.041
13.83 | 0.05005 | 4.6% | 4.3% | 4.1% | 0.041
13.84 | 0.05005 | 4.6% | 4.2% | 4.1% | 0.041
13.85 | 0.05005 | 4.6% | 4.2% | 4.1% | 0.041
13.86 | 0.05005 | 4.6% | 4.2% | 4.1% | 0.041
13.87 | 0.05005 | 4.6% | 4.2% | 4.1% | 0.041
13.88 | 0.05005 | 4.5% | 4.2% | 4.1% | 0.041
13.89 | 0.05005 | 4.5% | 4.2% | 4.1% | 0.041
13.9 | 0.05005 | 4.5% | 4.2% | 4.1% | 0.041
13.91 | 0.049049 | 4.5% | 4.2% | 4.1% | 0.041
13.92 | 0.049049 | 4.5% | 4.2% | 4.1% | 0.041
13.93 | 0.049049 | 4.5% | 4.2% | 4.1% | 0.041
13.94 | 0.049049 | 4.5% | 4.2% | 4.1% | 0.041
13.95 | 0.049049 | 4.5% | 4.2% | 4.1% | 0.041
13.96 | 0.049049 | 4.5% | 4.2% | 4% | 0.041
13.97 | 0.049049 | 4.5% | 4.2% | 4% | 0.041
13.98 | 0.049049 | 4.5% | 4.2% | 4% | 0.041
13.99 | 0.049049 | 4.5% | 4.2% | 4% | 0.041
14 | 0.049049 | 4.5% | 4.2% | 4% | 0.041
14.01 | 0.049049 | 4.5% | 4.2% | 4% | 0.041
14.02 | 0.049049 | 4.5% | 4.2% | 4% | 0.041
14.03 | 0.049049 | 4.5% | 4.2% | 4% | 0.041
14.04 | 0.049049 | 4.5% | 4.2% | 4% | 0.041
14.05 | 0.049049 | 4.5% | 4.2% | 4% | 0.041
14.06 | 0.049049 | 4.5% | 4.2% | 4% | 0.041
14.07 | 0.049049 | 4.5% | 4.2% | 4% | 0.04
14.08 | 0.049049 | 4.5% | 4.2% | 4% | 0.04
14.09 | 0.049049 | 4.5% | 4.2% | 4% | 0.04
14.1 | 0.049049 | 4.5% | 4.2% | 4% | 0.04
14.11 | 0.049049 | 4.5% | 4.2% | 4% | 0.04
14.12 | 0.049049 | 4.5% | 4.2% | 4% | 0.04
14.13 | 0.049049 | 4.5% | 4.2% | 4% | 0.04
14.14 | 0.049049 | 4.4% | 4.2% | 4% | 0.04
14.15 | 0.049049 | 4.4% | 4.2% | 4% | 0.04
14.16 | 0.049049 | 4.4% | 4.2% | 4% | 0.04
14.17 | 0.049049 | 4.4% | 4.1% | 4% | 0.04
14.18 | 0.049049 | 4.4% | 4.1% | 4% | 0.04
14.19 | 0.049049 | 4.4% | 4.1% | 4% | 0.04
14.2 | 0.049049 | 4.4% | 4.1% | 4% | 0.04
14.21 | 0.049049 | 4.4% | 4.1% | 4% | 0.04
14.22 | 0.049049 | 4.4% | 4.1% | 4% | 0.04
14.23 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.24 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.25 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.26 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.27 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.28 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.29 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.3 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.31 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.32 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.33 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.34 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.35 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.36 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.37 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.38 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.39 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.4 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.41 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.42 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.43 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.44 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.45 | 0.048048 | 4.4% | 4.1% | 3.9% | 0.04
14.46 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.47 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.48 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.49 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.5 | 0.048048 | 4.4% | 4.1% | 4% | 0.04
14.51 | 0.048048 | 4.3% | 4.1% | 4% | 0.04
14.52 | 0.048048 | 4.3% | 4% | 4% | 0.039
14.53 | 0.048048 | 4.3% | 4% | 4% | 0.039
14.54 | 0.048048 | 4.3% | 4% | 4% | 0.039
14.55 | 0.048048 | 4.3% | 4% | 4% | 0.039
14.56 | 0.048048 | 4.3% | 4% | 4% | 0.039
14.57 | 0.048048 | 4.3% | 4% | 4% | 0.039
14.58 | 0.048048 | 4.3% | 4% | 4% | 0.039
14.59 | 0.048048 | 4.3% | 4% | 4% | 0.039
14.6 | 0.048048 | 4.3% | 4% | 4% | 0.039
14.61 | 0.048048 | 4.3% | 4% | 4% | 0.039
14.62 | 0.048048 | 4.3% | 4% | 4% | 0.039
14.63 | 0.047047 | 4.3% | 4% | 3.9% | 0.039
14.64 | 0.047047 | 4.3% | 4% | 3.9% | 0.039
14.65 | 0.047047 | 4.3% | 4% | 3.9% | 0.039
14.66 | 0.047047 | 4.3% | 4% | 3.9% | 0.039
14.67 | 0.047047 | 4.3% | 4% | 3.9% | 0.039
14.68 | 0.047047 | 4.3% | 4% | 3.9% | 0.039
14.69 | 0.047047 | 4.3% | 4% | 3.9% | 0.039
14.7 | 0.047047 | 4.3% | 4% | 3.9% | 0.039
14.71 | 0.047047 | 4.3% | 4% | 3.9% | 0.039
14.72 | 0.047047 | 4.3% | 4% | 3.9% | 0.039
14.73 | 0.047047 | 4.3% | 4% | 3.9% | 0.039
14.74 | 0.047047 | 4.3% | 4% | 3.9% | 0.039
14.75 | 0.047047 | 4.3% | 4% | 3.9% | 0.039
14.76 | 0.047047 | 4.2% | 4% | 3.9% | 0.039
14.77 | 0.047047 | 4.2% | 4% | 3.9% | 0.039
14.78 | 0.047047 | 4.2% | 4% | 3.9% | 0.039
14.79 | 0.047047 | 4.2% | 4% | 3.9% | 0.039
14.8 | 0.047047 | 4.2% | 4% | 3.9% | 0.039
14.81 | 0.047047 | 4.2% | 4% | 3.9% | 0.038
14.82 | 0.047047 | 4.2% | 4% | 3.9% | 0.038
14.83 | 0.047047 | 4.2% | 4% | 3.9% | 0.038
14.84 | 0.047047 | 4.2% | 4% | 3.9% | 0.038
14.85 | 0.047047 | 4.2% | 4% | 3.9% | 0.038
14.86 | 0.047047 | 4.2% | 4% | 3.9% | 0.038
14.87 | 0.047047 | 4.2% | 4% | 3.9% | 0.038
14.88 | 0.047047 | 4.2% | 4% | 3.9% | 0.038
14.89 | 0.047047 | 4.2% | 4% | 3.9% | 0.038
14.9 | 0.047047 | 4.2% | 4% | 3.9% | 0.038
14.91 | 0.047047 | 4.2% | 4% | 3.9% | 0.038
14.92 | 0.047047 | 4.2% | 3.9% | 3.8% | 0.038
14.93 | 0.047047 | 4.2% | 3.9% | 3.8% | 0.038
14.94 | 0.047047 | 4.2% | 3.9% | 3.9% | 0.038
14.95 | 0.047047 | 4.2% | 3.9% | 3.9% | 0.038
14.96 | 0.047047 | 4.2% | 3.9% | 3.9% | 0.038
14.97 | 0.047047 | 4.2% | 3.9% | 3.9% | 0.038
14.98 | 0.047047 | 4.2% | 3.9% | 3.9% | 0.038
14.99 | 0.046046 | 4.2% | 3.9% | 3.9% | 0.038
15 | 0.046046 | 4.2% | 3.9% | 3.9% | 0.038
15.01 | 0.046046 | 4.2% | 3.9% | 3.9% | 0.038
15.02 | 0.046046 | 4.2% | 3.9% | 3.9% | 0.038
15.03 | 0.046046 | 4.2% | 3.9% | 3.9% | 0.038
15.04 | 0.046046 | 4.2% | 3.9% | 3.9% | 0.038
15.05 | 0.046046 | 4.2% | 3.9% | 3.9% | 0.038
15.06 | 0.046046 | 4.2% | 3.9% | 3.9% | 0.038
15.07 | 0.046046 | 4.2% | 3.9% | 3.9% | 0.038
15.08 | 0.046046 | 4.2% | 3.9% | 3.9% | 0.038
15.09 | 0.046046 | 4.2% | 3.9% | 3.9% | 0.038
15.1 | 0.046046 | 4.2% | 3.9% | 3.9% | 0.038
15.11 | 0.046046 | 4.2% | 3.9% | 3.9% | 0.038
15.12 | 0.046046 | 4.2% | 3.9% | 3.9% | 0.038
15.13 | 0.046046 | 4.2% | 3.9% | 3.8% | 0.038
15.14 | 0.046046 | 4.2% | 3.9% | 3.8% | 0.038
15.15 | 0.046046 | 4.2% | 3.9% | 3.8% | 0.038
15.16 | 0.046046 | 4.2% | 3.9% | 3.8% | 0.038
15.17 | 0.046046 | 4.2% | 3.9% | 3.8% | 0.038
15.18 | 0.046046 | 4.2% | 3.9% | 3.8% | 0.038
15.19 | 0.046046 | 4.2% | 3.9% | 3.8% | 0.038
15.2 | 0.046046 | 4.2% | 3.9% | 3.8% | 0.038
15.21 | 0.046046 | 4.2% | 3.9% | 3.8% | 0.038
15.22 | 0.046046 | 4.2% | 3.9% | 3.8% | 0.038
15.23 | 0.046046 | 4.2% | 3.9% | 3.8% | 0.038
15.24 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.038
15.25 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.038
15.26 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.038
15.27 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.038
15.28 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.038
15.29 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.038
15.3 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.038
15.31 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.038
15.32 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.038
15.33 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.038
15.34 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.038
15.35 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.038
15.36 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.037
15.37 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.037
15.38 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.037
15.39 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.037
15.4 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.037
15.41 | 0.046046 | 4.1% | 3.9% | 3.8% | 0.037
15.42 | 0.046046 | 4.1% | 3.8% | 3.8% | 0.037
15.43 | 0.046046 | 4.1% | 3.8% | 3.8% | 0.037
15.44 | 0.046046 | 4.1% | 3.8% | 3.8% | 0.037
15.45 | 0.046046 | 4.1% | 3.8% | 3.8% | 0.037
15.46 | 0.046046 | 4.1% | 3.8% | 3.8% | 0.037
15.47 | 0.045045 | 4.1% | 3.8% | 3.8% | 0.037
15.48 | 0.045045 | 4.1% | 3.8% | 3.8% | 0.037
15.49 | 0.045045 | 4.1% | 3.8% | 3.8% | 0.037
15.5 | 0.045045 | 4.1% | 3.8% | 3.8% | 0.037
15.51 | 0.045045 | 4.1% | 3.8% | 3.8% | 0.037
15.52 | 0.045045 | 4.1% | 3.8% | 3.8% | 0.037
15.53 | 0.045045 | 4.1% | 3.8% | 3.8% | 0.037
15.54 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.55 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.56 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.57 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.58 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.59 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.6 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.61 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.62 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.63 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.64 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.65 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.66 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.67 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.68 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.69 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.7 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.71 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.72 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.73 | 0.045045 | 4.1% | 3.8% | 3.7% | 0.037
15.74 | 0.045045 | 4% | 3.8% | 3.7% | 0.037
15.75 | 0.045045 | 4% | 3.8% | 3.7% | 0.037

## PDD Corrections
Energy | SSD | Field Size | R100 | R50 | R95 | R90 | R80 | Rp | Ds | d_ref | PDD(1.38) | Field size inline | PDD(1.3) | Field size inline | PDD Correction
6 MeV | 100.00 cm | 3x3 | 1.26 cm | 2.41 cm | 1.64 cm | 1.78 cm | 1.98 cm | 3.03 cm | 81% |  | 99.5% | 3 | 99.8% | 3
6 MeV | 100.00 cm | 4x4+ | 1.38 cm | 2.41 cm | 1.69 cm | 1.82 cm | 2.01 cm | 2.99 cm | 77.8% | 1.3 | 100% | 4 | 99.8% | 4 |  |  | PDD Correction
Energy
Energy | SSD | Field Size | R100 | R50 | R95 | R90 | R80 | Rp | Ds |  | PDD(2.13) | Field Size | PDD(2.1) | Field Size | PDD(2.1) |  | Field Size | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV
9 MeV | 100.00 cm | 3x3 | 1.62 cm | 3.49 cm | 2.25 cm | 2.48 cm | 2.80 cm | 4.46 cm | 87.7% |  | 96.8% | 3 | 97.3% | 3 |  |  | 3 | 100% | 97.3% | 92.9% | 96% | 99.5%
9 MeV | 100.00 cm | 4x4 | 1.97 cm | 3.61 cm | 2.52 cm | 2.73 cm | 3.02 cm | 4.43 cm | 83.8% |  | 99.7% | 4 | 99.8% | 4 |  |  | 4 | 100% | 99.8% | 98.1% | 98.8% | 99.8%
9 MeV | 100.00 cm | 6x6+ | 2.13 cm | 3.63 cm | 2.64 cm | 2.83 cm | 3.10 cm | 4.41 cm | 81.9% | 2.1 | 100% | 6 | 100% | 6 |  |  | 6 |  | 100% | 100% |  | 100.1%
[17 empty] | 6.5 |  |  |  | 100%
Energy | SSD | Field Size | R100 | R50 | R95 | R90 | R80 | Rp | Ds |  | PDD(2.78) | Field size | PDD(2.9) | Field size | PDD(2.9) |  | 7 |  |  |  |  | 100.1%
12 MeV | 100.00 cm | 3x3 | 1.96 cm | 4.56 cm | 2.76 cm | 3.08 cm | 3.54 cm | 6.08 cm | 91.2% |  | 94.8% | 3 | 92.9% | 3 |  |  | 10 |  |  |  | 100% | 100%
12 MeV | 100.00 cm | 4x4 | 2.17 cm | 4.86 cm | 3.21 cm | 3.54 cm | 3.97 cm | 6.11 cm | 89.4% |  | 99% | 4 | 98.1% | 4
12 MeV | 100.00 cm | 6x6+ | 2.78 cm | 5.02 cm | 3.66 cm | 3.93 cm | 4.30 cm | 6.05 cm | 87% | 2.9 | 100% | 6 | 100% | 6
[16 empty] | slope | 3 |  | 0.024996 | 0.051981 | 0.028028 | 0.0019324
Energy | SSD | Field Size | R100 | R50 | R95 | R90 | R80 | Rp | Ds |  | PDD(3.22) | Field size inline | PDD(3.0) | Field size inline | PDD(3.0) | intercept | 4 |  | 0.89801 | 0.77306 | 0.87588 | 0.98962
16 MeV | 100.00 cm | 3x3 | 1.68 cm | 5.55 cm | 3.11 cm | 3.57 cm | 4.17 cm | 7.75 cm | 93.1% |  | 94.09% | 3 | 96% | 3 |  | slope | 6 |  | 0.001002 | 0.0095095 | 0.0048048
16 MeV | 100.00 cm | 4x4 | 2.02 cm | 6.08 cm | 3.64 cm | 4.13 cm | 4.74 cm | 8.00 cm | 92.8% |  | 97.9% | 4 | 98.8% | 4 |  | intercept | 6.5 |  | 0.99399 | 0.94294 | 0.96877
16 MeV | 100.00 cm | 6x6 & 7x7 | 2.89 cm | 6.50 cm | 4.35 cm | 4.84 cm | 5.42 cm | 7.96 cm | 91.9% |  | 99.7% | 6.5 | 100% | 6.5 |  | slope
16 MeV | 100.00 cm | 10x10+ | 3.22 cm | 6.63 cm | 4.64 cm | 5.09 cm | 5.63 cm | 8.00 cm |  | 3 | 100% | 10 | 100% | 10
PDD Correction
Energy | SSD | Field Size | R100 | R50 | R95 | R90 | R80 | Rp | Ds |  | PDD(2.37) | Field size crossline | PDD(2.2) | Field size crossline | PDD(2.2) |  | Field Size | 6 MeV | 9 MeV | 12 MeV | 16 MeV | 20 MeV
20 MeV | 100.00 cm | 3x3 | 1.61 cm | 6.42 cm | 3.36 cm | 3.95 cm | 4.71 cm | 9.23 cm | 0.942 |  | 99% | 3 | 99.3% | 3 |  |  | 2.8 | 1 | 0.968 | 0.9186 | 0.95435 | 0.99503
20 MeV | 100.00 cm | 4x4 | 1.78 cm | 7.13 cm | 3.85 cm | 4.56 cm | 5.36 cm | 9.80 cm | 0.937 |  | 99.3% | 4 | 99.6% | 4 |  |  | 2.9 | 1 | 0.9705 | 0.9238 | 0.95716 | 0.99523
20 MeV | 100.00 cm | 6x6 | 1.85 cm | 7.85 cm | 4.49 cm | 5.28 cm | 6.19 cm | 9.98 cm | 0.938 |  | 100% | 6 | 99.9% | 6 |  |  | 3 | 1 | 0.973 | 0.929 | 0.95996 | 0.99542
20 MeV | 100.00 cm | 7x7 | 1.86 cm | 7.98 cm | 4.58 cm | 5.45 cm | 6.36 cm | 9.97 cm | 0.93 |  | 99.6% | 7 | 99.9% | 7 |  |  | 3.1 | 1 | 0.9755 | 0.9342 | 0.96276 | 0.99561
20 MeV | 100.00 cm | 10x10+ | 2.37 cm | 8.20 cm | 5.10 cm | 5.88 cm | 6.75 cm | 9.99 cm | 0.923 | 2.2 | 99.9% | 10 | 99.8% | 10 |  |  | 3.2 | 1 | 0.978 | 0.9394 | 0.96557 | 0.99581
[17 empty] | 3.3 | 1 | 0.9805 | 0.94459 | 0.96837 | 0.996
[17 empty] | 3.4 | 1 | 0.983 | 0.94979 | 0.97117 | 0.99619
[17 empty] | 3.5 | 1 | 0.9855 | 0.95499 | 0.97397 | 0.99639
[17 empty] | 3.6 | 1 | 0.988 | 0.96019 | 0.97678 | 0.99658
[17 empty] | 3.7 | 1 | 0.9905 | 0.96539 | 0.97958 | 0.99677
[17 empty] | 3.8 | 1 | 0.993 | 0.97058 | 0.98238 | 0.99697
[17 empty] | 3.9 | 1 | 0.9955 | 0.97578 | 0.98519 | 0.99716
[17 empty] | 4 | 1 | 0.998 | 0.98098 | 0.98799 | 0.99735
[17 empty] | 4.1 | 1 | 0.9981 | 0.98193 | 0.98847 | 0.99755
[17 empty] | 4.2 | 1 | 0.9982 | 0.98288 | 0.98895 | 0.99774
[17 empty] | 4.3 | 1 | 0.9983 | 0.98383 | 0.98943 | 0.99793
[17 empty] | 4.4 | 1 | 0.9984 | 0.98478 | 0.98991 | 0.99813
[17 empty] | 4.5 | 1 | 0.9985 | 0.98574 | 0.99039 | 0.99832
[17 empty] | 4.6 | 1 | 0.9986 | 0.98669 | 0.99087 | 0.99851
[17 empty] | 4.7 | 1 | 0.9987 | 0.98764 | 0.99135 | 0.99871
[17 empty] | 4.8 | 1 | 0.9988 | 0.98859 | 0.99183 | 0.9989
[17 empty] | 4.9 | 1 | 0.9989 | 0.98954 | 0.99231 | 0.99909
[17 empty] | 5 | 1 | 0.999 | 0.99049 | 0.99279 | 0.99929
[17 empty] | 5.1 | 1 | 0.9991 | 0.99144 | 0.99327 | 0.99948
[17 empty] | 5.2 | 1 | 0.9992 | 0.99239 | 0.99375 | 0.99967
[17 empty] | 5.3 | 1 | 0.9993 | 0.99334 | 0.99423 | 0.99987
[17 empty] | 5.4 | 1 | 0.9994 | 0.99429 | 0.99471 | 1.0001
[17 empty] | 5.5 | 1 | 0.9995 | 0.99525 | 0.9952 | 1.0003
[17 empty] | 5.6 | 1 | 0.9996 | 0.9962 | 0.99568 | 1.0004
[17 empty] | 5.7 | 1 | 0.9997 | 0.99715 | 0.99616 | 1.0006
[17 empty] | 5.8 | 1 | 0.9998 | 0.9981 | 0.99664 | 1.0008
[17 empty] | 5.9 | 1 | 0.9999 | 0.99905 | 0.99712 | 1.001
[17 empty] | 6 | 1 | 1 | 1 | 0.9976 | 1
[17 empty] | 6.1 | 1 | 1 | 1 | 0.99808 | 1
[17 empty] | 6.2 | 1 | 1 | 1 | 0.99856 | 1
[17 empty] | 6.3 | 1 | 1 | 1 | 0.99904 | 1
[17 empty] | 6.4 | 1 | 1 | 1 | 0.99952 | 1
[17 empty] | 6.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 6.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 6.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 6.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 6.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 7.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 7.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 7.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 7.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 7.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 7.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 7.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 7.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 7.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 8.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 8.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 8.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 8.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 8.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 8.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 8.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 8.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 8.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 9.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 9.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 9.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 9.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 9.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 9.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 9.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 9.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 9.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 10 | 1 | 1 | 1 | 1 | 1
[17 empty] | 10.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 10.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 10.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 10.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 10.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 10.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 10.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 10.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 10.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 11 | 1 | 1 | 1 | 1 | 1
[17 empty] | 11.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 11.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 11.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 11.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 11.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 11.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 11.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 11.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 11.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 12 | 1 | 1 | 1 | 1 | 1
[17 empty] | 12.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 12.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 12.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 12.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 12.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 12.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 12.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 12.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 12.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 13 | 1 | 1 | 1 | 1 | 1
[17 empty] | 13.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 13.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 13.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 13.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 13.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 13.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 13.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 13.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 13.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 14 | 1 | 1 | 1 | 1 | 1
[17 empty] | 14.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 14.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 14.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 14.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 14.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 14.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 14.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 14.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 14.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 15 | 1 | 1 | 1 | 1 | 1
[17 empty] | 15.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 15.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 15.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 15.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 15.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 15.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 15.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 15.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 15.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 16 | 1 | 1 | 1 | 1 | 1
[17 empty] | 16.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 16.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 16.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 16.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 16.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 16.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 16.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 16.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 16.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 17 | 1 | 1 | 1 | 1 | 1
[17 empty] | 17.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 17.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 17.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 17.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 17.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 17.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 17.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 17.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 17.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 18 | 1 | 1 | 1 | 1 | 1
[17 empty] | 18.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 18.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 18.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 18.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 18.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 18.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 18.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 18.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 18.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 19 | 1 | 1 | 1 | 1 | 1
[17 empty] | 19.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 19.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 19.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 19.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 19.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 19.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 19.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 19.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 19.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 20 | 1 | 1 | 1 | 1 | 1
[17 empty] | 20.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 20.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 20.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 20.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 20.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 20.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 20.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 20.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 20.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 21 | 1 | 1 | 1 | 1 | 1
[17 empty] | 21.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 21.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 21.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 21.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 21.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 21.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 21.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 21.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 21.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 22 | 1 | 1 | 1 | 1 | 1
[17 empty] | 22.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 22.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 22.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 22.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 22.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 22.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 22.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 22.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 22.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 23 | 1 | 1 | 1 | 1 | 1
[17 empty] | 23.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 23.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 23.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 23.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 23.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 23.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 23.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 23.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 23.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 24 | 1 | 1 | 1 | 1 | 1
[17 empty] | 24.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 24.2 | 1 | 1 | 1 | 1 | 1
[17 empty] | 24.3 | 1 | 1 | 1 | 1 | 1
[17 empty] | 24.4 | 1 | 1 | 1 | 1 | 1
[17 empty] | 24.5 | 1 | 1 | 1 | 1 | 1
[17 empty] | 24.6 | 1 | 1 | 1 | 1 | 1
[17 empty] | 24.7 | 1 | 1 | 1 | 1 | 1
[17 empty] | 24.8 | 1 | 1 | 1 | 1 | 1
[17 empty] | 24.9 | 1 | 1 | 1 | 1 | 1
[17 empty] | 25 | 1 | 1 | 1 | 1 | 1
[17 empty] | 25.1 | 1 | 1 | 1 | 1 | 1
[17 empty] | 25.2 | 1 | 1 | 1 | 1 | 1

## Dmax
Energy | SSD | Field Size | R100 | R50 | d_ref | PDD(1.38) | Field size inline | PDD(1.3)
6 MeV | 100.00 cm | 3x3 | 1.26 cm | 2.41 cm |  | 99.5% | 3 | 99.8%
6 MeV | 100.00 cm | 4x4+ | 1.38 cm | 2.41 cm | 1.3 | 100% | 4 | 99.8%
Energy | SSD | Field Size | R100 | R50 |  | PDD(2.13) | Field Size | PDD(2.1)
9 MeV | 100.00 cm | 3x3 | 1.62 cm | 3.49 cm |  | 96.8% | 3 | 97.3%
9 MeV | 100.00 cm | 4x4 | 1.97 cm | 3.61 cm |  | 99.7% | 4 | 99.8%
9 MeV | 100.00 cm | 6x6+ | 2.13 cm | 3.63 cm | 2.1 | 100% | 6 | 100%
Energy | SSD | Field Size | R100 | R50 |  | PDD(2.78) | Field size | PDD(2.9)
12 MeV | 100.00 cm | 3x3 | 1.96 cm | 4.56 cm |  | 94.8% | 3 | 92.9%
12 MeV | 100.00 cm | 4x4 | 2.17 cm | 4.86 cm |  | 99% | 4 | 98.1%
12 MeV | 100.00 cm | 6x6+ | 2.78 cm | 5.02 cm | 2.9 | 100% | 6 | 100%
Energy | SSD | Field Size | R100 | R50 |  | PDD(3.22) | Field size inline | PDD(3.0)
16 MeV | 100.00 cm | 3x3 | 1.68 cm | 5.55 cm |  | 94.09% | 3 | 96%
16 MeV | 100.00 cm | 4x4 | 2.02 cm | 6.08 cm |  | 97.9% | 4 | 98.8%
16 MeV | 100.00 cm | 6x6 & 7x7 | 2.89 cm | 6.50 cm |  | 99.7% | 6.5 | 100%
16 MeV | 100.00 cm | 10x10+ | 3.22 cm | 6.63 cm | 3 | 100% | 10 | 100%
Energy | SSD | Field Size | R100 | R50 |  | PDD(2.37) | Field size crossline | PDD(2.2)
20 MeV | 100.00 cm | 3x3 | 1.61 cm | 6.42 cm |  | 99% | 3 | 99.3%
20 MeV | 100.00 cm | 4x4 | 1.78 cm | 7.13 cm |  | 99.3% | 4 | 99.6%
20 MeV | 100.00 cm | 6x6 | 1.85 cm | 7.85 cm |  | 100% | 6 | 99.9%
20 MeV | 100.00 cm | 7x7 | 1.86 cm | 7.98 cm |  | 99.6% | 7 | 99.9%
20 MeV | 100.00 cm | 10x10+ | 2.37 cm | 8.20 cm | 2.2 | 99.9% | 10 | 99.8%