## CutOut Image
Expected Dimensions to Bottom of Insert
4.9131 | 4.5178
Expected Dimensions to Top of Insert
4.831 | 4.4423

## CutOut Parameters
Plan | NOSE | NOSE | Expected
Field | GA45 9MeV | GA45 6MEV
RadiationType | ELECTRON | ELECTRON | ELECTRON
SetupTechnique | FIXED_SSD | FIXED_SSD | FIXED_SSD
ToleranceTable | ELECTRON PLANNED | ELECTRON PLANNED | ELECTRON PLANNED | ?
Linac | TR2 | TR2 | TR2, TR3
Energy | 9 | 6
GantryAngle | 45 | 45
ApplicatorID | A06 | A06 | A06, A10, A15, A20, A25
AccessoryCode | 4000 | 4000 | 4000
BlockTrayID | FFDA(A06) | FFDA(A06) | FFDA(A06)
InsertCode | 126 | 126 | 126
BlockType | APERTURE | APERTURE | APERTURE
MaterialID | Cut-out | Cut-out | Cut-out
BlockDivergence | PRESENT | PRESENT | PRESENT
BlockName | CUSTOM | CUSTOM
SourceToBlockTrayDistance | 95 | 95 | 95
Cutout Area | 19.1391
Cutout Perimeter | 15.9688
Cutout Eq. Sq. | 4.7941 |  | 6
Cutout Extent | 2.66 |  | 3

## CutOut Coordinates
Block Points |  | Scale Factor | 94.5472 |  |  | Distances
X | Y | X Scaled | Y Scaled | X Range | Y Range | Source to top of electron insert | 93 | cm
2.55 | 0.44 | 2.411 | 0.41601 | 2.56 | 2.5 | Source to bottom of electron insert | 95 | cm
2.56 | 0.19 | 2.4204 | 0.17964 | -2.66 | -2.3 | Source to bottom of electron applicator | 96.8 | cm
2.55 | 0.18 | 2.411 | 0.17019 | 5.22 | 4.8
2.53 | 0.1 | 2.392 | 0.094547 |  |  | SSD | 100.4528
2.5 | -0.17 | 2.3637 | -0.16073 |  |  | Insert Size | 6
2.51 | -0.18 | 2.3731 | -0.17019
2.52 | -0.43 | 2.3826 | -0.40655
2.51 | -0.44 | 2.3731 | -0.41601 |  |  | Scale Factors
2.49 | -0.62 | 2.3542 | -0.58619 |  |  | bottom of electron insert | 94.5472
2.45 | -0.78 | 2.3164 | -0.73747 |  |  | top of electron insert | 92.5472
2.37 | -0.94 | 2.2408 | -0.88874
2.31 | -1.18 | 2.184 | -1.1157
2.23 | -1.42 | 2.1084 | -1.3426
2.11 | -1.66 | 1.9949 | -1.5695
2.07 | -1.7 | 1.9571 | -1.6073
2.05 | -1.74 | 1.9382 | -1.6451
1.9 | -1.89 | 1.7964 | -1.7869
1.86 | -1.91 | 1.7586 | -1.8059
1.82 | -1.95 | 1.7208 | -1.8437
1.74 | -1.99 | 1.6451 | -1.8815
1.66 | -2.07 | 1.5695 | -1.9571
1.62 | -2.09 | 1.5317 | -1.976
1.58 | -2.13 | 1.4938 | -2.0139
1.38 | -2.23 | 1.3048 | -2.1084
1.22 | -2.27 | 1.1535 | -2.1462
1.06 | -2.29 | 1.0022 | -2.1651
0.65 | -2.3 | 0.61456 | -2.1746
0.64 | -2.29 | 0.6051 | -2.1651
0.52 | -2.27 | 0.49165 | -2.1462
0.44 | -2.25 | 0.41601 | -2.1273
0.32 | -2.21 | 0.30255 | -2.0895
0.16 | -2.13 | 0.15128 | -2.0139
0.14 | -2.11 | 0.13237 | -1.9949
0.1 | -2.09 | 0.094547 | -1.976
-0.06 | -2.05 | -0.056728 | -1.9382
-0.18 | -2.01 | -0.17019 | -1.9004
-0.3 | -1.95 | -0.28364 | -1.8437
-0.46 | -1.91 | -0.43492 | -1.8059
-0.64 | -1.85 | -0.6051 | -1.7491
-0.76 | -1.79 | -0.71856 | -1.6924
-0.78 | -1.77 | -0.73747 | -1.6735
-0.82 | -1.75 | -0.77529 | -1.6546
-0.84 | -1.73 | -0.7942 | -1.6357
-0.88 | -1.71 | -0.83202 | -1.6168
-1 | -1.59 | -0.94547 | -1.5033
-1.2 | -1.57 | -1.1346 | -1.4844
-1.46 | -1.55 | -1.3804 | -1.4655
-1.58 | -1.53 | -1.4938 | -1.4466
-1.66 | -1.51 | -1.5695 | -1.4277
-1.78 | -1.47 | -1.6829 | -1.3898
-1.9 | -1.41 | -1.7964 | -1.3331
-1.92 | -1.39 | -1.8153 | -1.3142
-1.96 | -1.37 | -1.8531 | -1.2953
-1.98 | -1.35 | -1.872 | -1.2764
-2.02 | -1.33 | -1.9099 | -1.2575
-2.19 | -1.16 | -2.0706 | -1.0967
-2.21 | -1.12 | -2.0895 | -1.0589
-2.23 | -1.1 | -2.1084 | -1.04
-2.31 | -0.94 | -2.184 | -0.88874
-2.35 | -0.82 | -2.2219 | -0.77529
-2.37 | -0.78 | -2.2408 | -0.73747
-2.39 | -0.76 | -2.2597 | -0.71856
-2.41 | -0.72 | -2.2786 | -0.68074
-2.49 | -0.64 | -2.3542 | -0.6051
-2.51 | -0.6 | -2.3731 | -0.56728
-2.53 | -0.58 | -2.392 | -0.54837
-2.59 | -0.46 | -2.4488 | -0.43492
-2.61 | -0.4 | -2.4677 | -0.37819
-2.63 | -0.32 | -2.4866 | -0.30255
-2.65 | -0.22 | -2.5055 | -0.208
-2.66 | 0.13 | -2.515 | 0.12291
-2.65 | 0.14 | -2.5055 | 0.13237
-2.63 | 0.22 | -2.4866 | 0.208
-2.61 | 0.28 | -2.4677 | 0.26473
-2.53 | 0.44 | -2.392 | 0.41601
-2.51 | 0.56 | -2.3731 | 0.52946
-2.41 | 0.96 | -2.2786 | 0.90765
-2.37 | 1.08 | -2.2408 | 1.0211
-2.27 | 1.28 | -2.1462 | 1.2102
-2.25 | 1.3 | -2.1273 | 1.2291
-2.23 | 1.34 | -2.1084 | 1.2669
-2.06 | 1.51 | -1.9477 | 1.4277
-2.03 | 1.52 | -1.9193 | 1.4371
-1.97 | 1.64 | -1.8626 | 1.5506
-1.95 | 1.66 | -1.8437 | 1.5695
-1.93 | 1.7 | -1.8248 | 1.6073
-1.91 | 1.72 | -1.8059 | 1.6262
-1.89 | 1.76 | -1.7869 | 1.664
-1.66 | 1.99 | -1.5695 | 1.8815
-1.62 | 2.01 | -1.5317 | 1.9004
-1.6 | 2.03 | -1.5128 | 1.9193
-1.56 | 2.05 | -1.4749 | 1.9382
-1.54 | 2.07 | -1.456 | 1.9571
-1.34 | 2.17 | -1.2669 | 2.0517
-1.32 | 2.19 | -1.248 | 2.0706
-1.28 | 2.21 | -1.2102 | 2.0895
-1.26 | 2.23 | -1.1913 | 2.1084
-1.22 | 2.25 | -1.1535 | 2.1273
-1.2 | 2.27 | -1.1346 | 2.1462
-1.08 | 2.33 | -1.0211 | 2.203
-0.96 | 2.37 | -0.90765 | 2.2408
-0.88 | 2.39 | -0.83202 | 2.2597
-0.78 | 2.41 | -0.73747 | 2.2786
-0.54 | 2.45 | -0.51056 | 2.3164
-0.4 | 2.47 | -0.37819 | 2.3353
-0.26 | 2.48 | -0.24582 | 2.3448
-0.26 | 2.5 | -0.24582 | 2.3637
0.28 | 2.49 | 0.26473 | 2.3542
0.44 | 2.45 | 0.41601 | 2.3164
0.56 | 2.39 | 0.52946 | 2.2597
0.62 | 2.37 | 0.58619 | 2.2408
0.72 | 2.35 | 0.68074 | 2.2219
0.88 | 2.31 | 0.83202 | 2.184
1.08 | 2.21 | 1.0211 | 2.0895
1.12 | 2.17 | 1.0589 | 2.0517
1.16 | 2.15 | 1.0967 | 2.0328
1.24 | 2.07 | 1.1724 | 1.9571
1.44 | 1.97 | 1.3615 | 1.8626
1.46 | 1.95 | 1.3804 | 1.8437
1.54 | 1.93 | 1.456 | 1.8248
1.74 | 1.83 | 1.6451 | 1.7302
1.78 | 1.79 | 1.6829 | 1.6924
1.82 | 1.77 | 1.7208 | 1.6735
1.94 | 1.65 | 1.8342 | 1.56
1.98 | 1.63 | 1.872 | 1.5411
2.02 | 1.59 | 1.9099 | 1.5033
2.06 | 1.57 | 1.9477 | 1.4844
2.15 | 1.48 | 2.0328 | 1.3993
2.17 | 1.44 | 2.0517 | 1.3615
2.21 | 1.4 | 2.0895 | 1.3237
2.31 | 1.2 | 2.184 | 1.1346
2.39 | 0.88 | 2.2597 | 0.83202
2.41 | 0.78 | 2.2786 | 0.73747
2.43 | 0.76 | 2.2975 | 0.71856
2.51 | 0.6 | 2.3731 | 0.56728
2.55 | 0.44 | 2.411 | 0.41601

## Block Coordinates
| NOSE | NOSE | NOSE | NOSE
| GA45 6MEV | GA45 6MEV | GA45 9MeV | GA45 9MeV
| X | Y | X | Y
0 | 2.55 | 0.44 | 2.55 | 0.44
1 | 2.56 | 0.19 | 2.56 | 0.19
2 | 2.55 | 0.18 | 2.55 | 0.18
3 | 2.53 | 0.1 | 2.53 | 0.1
4 | 2.5 | -0.17 | 2.5 | -0.17
5 | 2.51 | -0.18 | 2.51 | -0.18
6 | 2.52 | -0.43 | 2.52 | -0.43
7 | 2.51 | -0.44 | 2.51 | -0.44
8 | 2.49 | -0.62 | 2.49 | -0.62
9 | 2.45 | -0.78 | 2.45 | -0.78
10 | 2.37 | -0.94 | 2.37 | -0.94
11 | 2.31 | -1.18 | 2.31 | -1.18
12 | 2.23 | -1.42 | 2.23 | -1.42
13 | 2.11 | -1.66 | 2.11 | -1.66
14 | 2.07 | -1.7 | 2.07 | -1.7
15 | 2.05 | -1.74 | 2.05 | -1.74
16 | 1.9 | -1.89 | 1.9 | -1.89
17 | 1.86 | -1.91 | 1.86 | -1.91
18 | 1.82 | -1.95 | 1.82 | -1.95
19 | 1.74 | -1.99 | 1.74 | -1.99
20 | 1.66 | -2.07 | 1.66 | -2.07
21 | 1.62 | -2.09 | 1.62 | -2.09
22 | 1.58 | -2.13 | 1.58 | -2.13
23 | 1.38 | -2.23 | 1.38 | -2.23
24 | 1.22 | -2.27 | 1.22 | -2.27
25 | 1.06 | -2.29 | 1.06 | -2.29
26 | 0.65 | -2.3 | 0.65 | -2.3
27 | 0.64 | -2.29 | 0.64 | -2.29
28 | 0.52 | -2.27 | 0.52 | -2.27
29 | 0.44 | -2.25 | 0.44 | -2.25
30 | 0.32 | -2.21 | 0.32 | -2.21
31 | 0.16 | -2.13 | 0.16 | -2.13
32 | 0.14 | -2.11 | 0.14 | -2.11
33 | 0.1 | -2.09 | 0.1 | -2.09
34 | -0.06 | -2.05 | -0.06 | -2.05
35 | -0.18 | -2.01 | -0.18 | -2.01
36 | -0.3 | -1.95 | -0.3 | -1.95
37 | -0.46 | -1.91 | -0.46 | -1.91
38 | -0.64 | -1.85 | -0.64 | -1.85
39 | -0.76 | -1.79 | -0.76 | -1.79
40 | -0.78 | -1.77 | -0.78 | -1.77
41 | -0.82 | -1.75 | -0.82 | -1.75
42 | -0.84 | -1.73 | -0.84 | -1.73
43 | -0.88 | -1.71 | -0.88 | -1.71
44 | -1 | -1.59 | -1 | -1.59
45 | -1.2 | -1.57 | -1.2 | -1.57
46 | -1.46 | -1.55 | -1.46 | -1.55
47 | -1.58 | -1.53 | -1.58 | -1.53
48 | -1.66 | -1.51 | -1.66 | -1.51
49 | -1.78 | -1.47 | -1.78 | -1.47
50 | -1.9 | -1.41 | -1.9 | -1.41
51 | -1.92 | -1.39 | -1.92 | -1.39
52 | -1.96 | -1.37 | -1.96 | -1.37
53 | -1.98 | -1.35 | -1.98 | -1.35
54 | -2.02 | -1.33 | -2.02 | -1.33
55 | -2.19 | -1.16 | -2.19 | -1.16
56 | -2.21 | -1.12 | -2.21 | -1.12
57 | -2.23 | -1.1 | -2.23 | -1.1
58 | -2.31 | -0.94 | -2.31 | -0.94
59 | -2.35 | -0.82 | -2.35 | -0.82
60 | -2.37 | -0.78 | -2.37 | -0.78
61 | -2.39 | -0.76 | -2.39 | -0.76
62 | -2.41 | -0.72 | -2.41 | -0.72
63 | -2.49 | -0.64 | -2.49 | -0.64
64 | -2.51 | -0.6 | -2.51 | -0.6
65 | -2.53 | -0.58 | -2.53 | -0.58
66 | -2.59 | -0.46 | -2.59 | -0.46
67 | -2.61 | -0.4 | -2.61 | -0.4
68 | -2.63 | -0.32 | -2.63 | -0.32
69 | -2.65 | -0.22 | -2.65 | -0.22
70 | -2.66 | 0.13 | -2.66 | 0.13
71 | -2.65 | 0.14 | -2.65 | 0.14
72 | -2.63 | 0.22 | -2.63 | 0.22
73 | -2.61 | 0.28 | -2.61 | 0.28
74 | -2.53 | 0.44 | -2.53 | 0.44
75 | -2.51 | 0.56 | -2.51 | 0.56
76 | -2.41 | 0.96 | -2.41 | 0.96
77 | -2.37 | 1.08 | -2.37 | 1.08
78 | -2.27 | 1.28 | -2.27 | 1.28
79 | -2.25 | 1.3 | -2.25 | 1.3
80 | -2.23 | 1.34 | -2.23 | 1.34
81 | -2.06 | 1.51 | -2.06 | 1.51
82 | -2.03 | 1.52 | -2.03 | 1.52
83 | -1.97 | 1.64 | -1.97 | 1.64
84 | -1.95 | 1.66 | -1.95 | 1.66
85 | -1.93 | 1.7 | -1.93 | 1.7
86 | -1.91 | 1.72 | -1.91 | 1.72
87 | -1.89 | 1.76 | -1.89 | 1.76
88 | -1.66 | 1.99 | -1.66 | 1.99
89 | -1.62 | 2.01 | -1.62 | 2.01
90 | -1.6 | 2.03 | -1.6 | 2.03
91 | -1.56 | 2.05 | -1.56 | 2.05
92 | -1.54 | 2.07 | -1.54 | 2.07
93 | -1.34 | 2.17 | -1.34 | 2.17
94 | -1.32 | 2.19 | -1.32 | 2.19
95 | -1.28 | 2.21 | -1.28 | 2.21
96 | -1.26 | 2.23 | -1.26 | 2.23
97 | -1.22 | 2.25 | -1.22 | 2.25
98 | -1.2 | 2.27 | -1.2 | 2.27
99 | -1.08 | 2.33 | -1.08 | 2.33
100 | -0.96 | 2.37 | -0.96 | 2.37
101 | -0.88 | 2.39 | -0.88 | 2.39
102 | -0.78 | 2.41 | -0.78 | 2.41
103 | -0.54 | 2.45 | -0.54 | 2.45
104 | -0.4 | 2.47 | -0.4 | 2.47
105 | -0.26 | 2.48 | -0.26 | 2.48
106 | -0.26 | 2.5 | -0.26 | 2.5
107 | 0.28 | 2.49 | 0.28 | 2.49
108 | 0.44 | 2.45 | 0.44 | 2.45
109 | 0.56 | 2.39 | 0.56 | 2.39
110 | 0.62 | 2.37 | 0.62 | 2.37
111 | 0.72 | 2.35 | 0.72 | 2.35
112 | 0.88 | 2.31 | 0.88 | 2.31
113 | 1.08 | 2.21 | 1.08 | 2.21
114 | 1.12 | 2.17 | 1.12 | 2.17
115 | 1.16 | 2.15 | 1.16 | 2.15
116 | 1.24 | 2.07 | 1.24 | 2.07
117 | 1.44 | 1.97 | 1.44 | 1.97
118 | 1.46 | 1.95 | 1.46 | 1.95
119 | 1.54 | 1.93 | 1.54 | 1.93
120 | 1.74 | 1.83 | 1.74 | 1.83
121 | 1.78 | 1.79 | 1.78 | 1.79
122 | 1.82 | 1.77 | 1.82 | 1.77
123 | 1.94 | 1.65 | 1.94 | 1.65
124 | 1.98 | 1.63 | 1.98 | 1.63
125 | 2.02 | 1.59 | 2.02 | 1.59
126 | 2.06 | 1.57 | 2.06 | 1.57
127 | 2.15 | 1.48 | 2.15 | 1.48
128 | 2.17 | 1.44 | 2.17 | 1.44
129 | 2.21 | 1.4 | 2.21 | 1.4
130 | 2.31 | 1.2 | 2.31 | 1.2
131 | 2.39 | 0.88 | 2.39 | 0.88
132 | 2.41 | 0.78 | 2.41 | 0.78
133 | 2.43 | 0.76 | 2.43 | 0.76
134 | 2.51 | 0.6 | 2.51 | 0.6
135 | 2.55 | 0.44 | 2.55 | 0.44

## Plan Data
| NOSE | NOSE
| GA45 9MeV | GA45 6MEV
FieldType | STATIC | STATIC
RadiationType | ELECTRON | ELECTRON
Weight | 1 | 1
Linac | TR2 | TR2
SetupField | TREATMENT | TREATMENT
SAD | 100 | 100
NumberOfBlocks | 1 | 1
NumberOfBoli | 1 | 1
NumberOfControlPoints | 2 | 2
NumberOfWedges | 0 | 0
CollimatorAngle | 345 | 345
DoseRate | 600 | 600
GantryAngle | 45 | 45
Energy | 9 | 6
CouchAngle | 345 | 345
Actual SSD | 100.4528 | 100.4528
Isocentre
AccessoryCode | 4000 | 4000
ApplicatorID | A06 | A06
ApplicatorType | ELECTRON_SQUARE | ELECTRON_SQUARE
ApplicatorApertureShape | SYM_SQUARE | SYM_SQUARE
ApplicatorOpening | 6 | 6
BlockName | CUSTOM | CUSTOM
BlockTrayID | FFDA(A06) | FFDA(A06)
MaterialID | Cut-out | Cut-out
BlockType | APERTURE | APERTURE
InsertCode | 126 | 126
BlockDivergence | PRESENT | PRESENT
SourceToBlockTrayDistance | 95 | 95
BlockMountingPosition | PATIENT_SIDE | PATIENT_SIDE
BlockThickness | 2.5 | 2.5
SourceToBlockDistance | 92.5 | 92.5
MUs | 234.8439 | 160.1974
Beam Dose | 2.1 | 1.4
SetupTechnique | FIXED_SSD | FIXED_SSD
PatientOrientation | HFS | HFS
ToleranceTable | ELECTRON PLANNED | ELECTRON PLANNED
PatientID | ElectronCutout | ElectronCutout

## Applicators & Inserts
Applicator
ID | Name | Internal Code | Field Size X [cm] | Field Size Y [cm] | Distance from the source [cm] | Tray Id | Internal Code | Add-On Material | Block ID | Thickness [cm]
A25 | A25 | 4004 | 25 | 25 | 95 | FFDA(A10+) | 4094 | Electron Insert | Cut-out | 2.5
A20 | A20 | 4003 | 20 | 20 | 95 | FFDA(A10+) | 4094 | Electron Insert | Cut-out | 2.5
A15 | A15 | 4002 | 15 | 15 | 95 | FFDA(A10+) | 4094 | Electron Insert | Cut-out | 2.5
A10 | A10 | 4001 | 10 | 10 | 95 | FFDA(A10+) | 4094 | Electron Insert | Cut-out | 2.5
A06 | A06 | 4000 | 6 | 6 | 95 | FFDA(A06) | 126 | Electron Insert | Cut-out | 2.5